# Finger Stick Blood Sugar:  Glucometer (PXP) Overview

- Controls

  - Frequency (every 24 hours)

  - 2 Levels: Low & high

  - Opened bottle expiration date (90 days, NOT manufacturer’s printed date)

- Perform Patient Test

  - Verbally confirm patient’s name and date of birth before performing test

  - Sample collection

    - WIPE OFF FIRST DROP OF BLOOD! (interstitial fluid contamination will cause a falsely low result)

  - Test linearity (20-500 mg/dl): Any value below or above this range will be displayed as <20 or >500 mg/dl, respectively.

  - Critical values (<50 mg/dl and >400 mg/dl for adults, <25 mg/dl for neonates)

    - Readback policy (Provider must be notified and notification documented in patient’s record)

  - Manual result entry must be done immediately to record the result into the patient’s medical record.

- Cleaning the Instrument:  Use Saniwipes before and after every use

Please review the lab procedures in the BDHS organization-wide Policies and Procedures manual.  This manual may be found in SharePoint under the Policies and Procedures tab, Organization Wide Procedures, Laboratory.  You will be asked to locate these procedures by an inspector.

1/2014 DDe