Document Title: 	Factor Assays Intrinsic on the ACL TOP 500

| Author |  | Effective Date |  | Supersedes Procedure # |
| --- | --- | --- | --- | --- |
| Michael Delski |  |  |  |  |

| Revised By | Date Revised | Effective (adopted) Date | Version # | Reason for Revision |
| --- | --- | --- | --- | --- |
| RJM | 3/15/2017 |  | 2 | Update to calculations and result reporting |

| Version # | Approval Signature | Approval Date |
| --- | --- | --- |
| 2 | Majed Refaai, MD, Medical Director |  |

| Version # | Distributed To | Date | # of Copies | Date of Removal |
| --- | --- | --- | --- | --- |
| 2 | QC, Lab Bench, SP |  | 1,1 |  |

# TITLE: Factor Assays Intrinsic

- Purpose

This procedure provides instructions for the quantitative determination of Factor VIII, IX, XI or XII in citrated plasma based on Activated Partial Thromboplastin Time (APTT) assay using HemosIL SynthASil or HemosIL APTT-SP reagent in liquid form, on the ACL TOP® Family.1,6,7,8,9,11

- PRINCIPLE

Factor VIII, IX, XI or XII activity in a patient’s plasma determined by performing a modified APTT assay. Patient plasma is diluted and added to the corresponding Factor Deficient plasma reagent. Correction of the clotting time of the deficient plasma is proportional to the concentration (% activity) of that factor in the patient plasma interpreted from a calibration curve.6,7,8,9

- IV.	SCOPE

- This procedure will be used by the University of Rochester Medical Center, Strong Memorial Hospital, Hematology and Chemistry Laboratory staff.

- IV.	RESPONSIBILITIES

| - TITLE | - RESPONSIBILITY |
| --- | --- |
| - Laboratory Director | - Approval of Procedure |
| - Supervisory Staff | - Implementation of procedure |
| - Staff | - Adherence to procedure |

- V.	SPECIMENS

- Citrated blood 9:1 (blood to anticoagulant) 3.2 % sodium citrate.Follow NCCLS guideline H21-A2 and H3-A3.  No other anticoagulant is acceptable.

B.	Centrifugation:12 minutes at 4,000 rpm (RCF = 3756g). or  2  minutes at 16,000 RPM (Eppendorf High Speed Centrifuge.

C.	Plasma Storage: 4 hours at 20  ºC ,1 month at -70 ºC.  Frozen plasma should be thawed only once at 37 ºC for 5 minutes.

D.	Unacceptable Specimens:  Samples that are short draws, clotted or hemolyzed may yield incorrect results.

- Sample volume is CRITICAL to obtain accurate coagulation results. A 90% draw is the minimum volume acceptable for accurate testing. Do not run or report these samples. Place “.ND” in the result field with the canned text comment “Sample volume inadequate, unable to performed required testing” In Order Entry add the SPROB test by following procedure. See Documentation of Patient/Specimen SH.CP.SM.loe.0180.0005.in the comment field and call for redraw.  Refer to the Minimum volume indicator on the tube, see the BD Tube draw volume guide.

- Elevated hematocrit specimens (specimens determined to have a hematocrit >55%) may require the preparation of a special collection tube: Refer to General Operating Procedure for the ACL TOP 500, Section IV ( SH.CP.AU.coa.0036 or see SH.CP.AU.jad.0136).

- Clotted samples: Each specimen is checked visually for the presence of clots prior to analysis.  If a clot is suspected, the tube is uncapped, and checked with a pair of applicator sticks.  ANY clot present in the specimen makes it inadequate for ALL coag testing. Clotted specimens are rejected and a new specimen should be requested.

- Hemolyzed samples: Grossly hemolyzed samples should be rejected. If sample has been centrifuged and hemolysis is present, check for a clot. If a clot is present, follow procedure for reporting clotted samples.

- VI	QUALITY CONTROL

- Normal and abnormal controls are recommended for a complete quality control program. HemosIL Controls for Prothrombin Time are designed for this program. Each laboratory should establish its own mean and standard deviation and should establish a quality control program to monitor laboratory testing.  Controls should be analyzed at least once every 8 hour shift and with reagent change in accordance with good laboratory practice.  Refer to Westgard et al for identification and resolution of out-of-control situations.

QC data is transmitted to the LIS. Remedial action for out-of range values is documented in the LIS and on the QC log for that instrument. When a value is out of range, the operator should repeat the control using the same control material – if acceptable document in LIS.

- If still not acceptable reconstitute a new control and rerun – if acceptable document in LIS.

- If still not acceptable, run the Enhanced Probe Clean procedure and rerun the same well mixed control and well mixed reagent. If acceptable, document and continue processing samples.

- If still not acceptable, reconstitute/use new reagent and rerun. If acceptable, a lookback of the last five samples analyzed (or an appropriate number of samples) for that assay on that analyzer should be done. These samples should be from the time when the last reagent vial was in use and the last QC point was in range, in accordance with the Quality Control Lookback Policy SH.CP.AU.gen.0002. If the samples do not agree, continue to look back at samples analyzed on that analyzer since the last successful QC run. Notify supervisor or technical specialist.

- In the case of any if >4.0 SD QC failure (R1:XS in Soft QC), rerun the QC and evaluate according to Quality Control Policy SH.CP.AU.gen.0001. Perform a lookback as necessary according to the aforementioned policy if any patient samples were analyzed.

- If still not acceptable, document and notify Hematology Supervisor or a Tech Specialist. Use backup analyzer for STAT/routine testing. Call service for further guidance and troubleshooting help.

- VII.	SPECIAL SAFETY PRECAUTIONS

- All patient specimens should be considered potentially infectious and must be handled with precautions used for human blood, as described in CDC (Center for Disease Control) recommendations and in compliance with the Federal OSHA (Occupational Safety and Health Administration) Blood-borne Pathogen Standard, 29 CFR (Code of Federal Regulations) part 1910.1030.  Follow specimen handling as outlined by Laboratory Safety Policy, (SH.CP.AU.gen.0005.0001).

- VIII.	MATERIALS

- Equipment

ACL TOP 500 analyzer

centrifuge

- Supplies

Pipettes

Pipette tips

Plastic test tubes

TYPE –I Reagent Grade Water

- Reagents

HemosIL SynthASil (PN 0020006800)  contains reagent and calcium chloride in liquid form.1

HemosIL Factor VIII Deficient plasma (PN0020011800) is lyophilized human plasma, artificially depleted of FVIII, and containing buffer and stabilizers. Residual FVIII activity < 1% and other factors have normal levels.6

HemosIL Factor IX Deficient plasma (PN 0020011900) is lyophilized human plasma, artificially depleted of FIX, and containing buffer and stabilizers. Residual FIX activity < 1% and other factors have normal levels.7

HemosIL Factor XI Deficient plasma (PN 0020011300) is lyophilized human plasma, artificially depleted of FXI, and containing buffer and stabilizers. Residual FXI activity < 1% and other factors have normal levels.8

HemosIL Factor XII Deficient plasma (PN 0020011200) is lyophilized human plasma, artificially depleted of FXII, and containing buffer and stabilizers. Residual FXII activity < 1% and other factors have normal levels.9

Please refer to the package inserts for further detail.1,6,7,8,9,11

The following are not supplied with the kits and those required may be purchased separately.1,6,7,8,9

HemosIL Calibration Plasma 		00020003700

HemosIL Normal Control ASSAYED		00020003110

HemosIL Special Test Control Level 2		00020012000

HemosIL Factor Diluent		00009757600

IX.	PROCEDURE – (STEP/ACTION)

- Reagent Preparation

SynthASil (PN 0020006800): Each vial of APTT reagent must be equilibrated at 15-25˚C for at least 15 minutes and mixed thoroughly before use.1 Calcium Chloride (packaged with either APTT reagent): Ready for use.1,11 Factor Deficient plasmas: Dissolve the contents of each required vial with 1 mL of CLSI CLR water or equivalent.11 Replace the stopper and swirl gently. Ensure complete reconstitution of the product. Keep at 15-25˚C for 30 minutes and invert to mix before use. Do not shake. Avoid foam formation.6,7,8,9

- Reagent Storage and Stability

Unopened reagents are stable until the expiration date shown on the vial when stored at 2-8°C.

SynthASil: Opened reagent is stable 30 days at 2-8°C in the original vial and 10 days at 15°C on the ACL TOP® Family. Do not freeze.1

Calcium Chloride: Opened reagent is stable 30 days at 2-30°C.1

Calcium Chloride: Opened reagent is stable 30 days at 2-8°C.11

Factor Deficient plasmas: Stability after reconstitution:  Plasma is stable 24 hours at 2-8˚C in the original vial or 24 hours at 15˚C on the ACL TOP® Family.6,7,8,9

For optimal stability remove reagents from the system and store them at 2-8°C in the original vial.1,6,7,8,9,11

- Calibration

Calibration and storage of a valid specific Factor Assay calibration are required to obtain Factor results. Calibration is performed: - With a change of reagent lot numbers - With a change of major instrument components - To satisfy local regulatory requirements - At laboratory discretion Method for Calibration (if necessary): - Define Materials if necessary (Setup, Material List). Select the appropriate materials from the Material List. Add appropriate Factor assay reagents (SynthASil Reagent and APTT Calcium Chloride), Factor Deficient Plasma(s), Calibration Plasma and Factor Diluent to the Materials Programming Window if necessary (Setup, Display, Materials Programming Window). - Define Results Units and Rerun Rules in the Specific Factor Test Definition if necessary (Setup, Test List, Test Code, Result Units and Rerun Rules). - Choose Setup, Materials List. - Double-click on the appropriate calibrator to open the Materials Definition screen. - Choose the Lot Specific Information tab and enter the Calibrator lot number and Expiration Date. - Enable Lot Management from the Lot Specific Information tab. - Select the Save icon to store the lot number. Once the lot number is saved, the Assign Values icon becomes available. - Select the Assign Values icon. - Enter the calibration value from the package insert. Press OK. - Choose the Previous Screen icon to exit. - Load the appropriate Factor assay reagents (APTT Reagent and APTT Calcium Chloride, Factor Deficient Plasma(s),) Calibration Plasma and Factor Diluent onto the ACL TOP® Family instrument. - Select Calibration, Status List. - Double-click on the appropriate Factor Assay test code to open the Calibration Details screen. - Choose the Run icon. - Select OK at the “Do you confirm the operation?” prompt. - Choose the Previous Screen icon to exit. - Verify the Job Status for the Factor Assay test code says Active. - Once the calibration is complete, review calibration results. The instrument will fail the calibration if the r2 value is less than the value in the chart below: Intrinsic Factors Assay; Math Model High (r2 ≥); Math Model Low (r2 ≥). Factor VIII SS; 0.985; 0.985. Factor IX SS; 0.985; 0.985. Factor XI SS; 0.985; 0.985. Factor XII SS; 0.985; 0.985. - If there are no errors/failures and the calibration is acceptable, choose the Validate icon to validate the calibration curve.2

- Test Procedure

- Create/Edit QC files, if necessary (Setup, QC List, Test Code to access the QC Definition Screen). Add FVIII, FIX, FXI and/or FXII test codes to the Test/Profiles Programming Window, if necessary (Setup, Display, Test Programming Window).

- Add appropriate Factor assay reagents (APTT Reagent and APTT Calcium Chloride), Factor Deficient Plasma(s), and Factor Diluent to the Materials Programming Window if necessary (Setup, Display, Materials Programming Window).

- Define Parallelism parameters, if necessary (Setup, Test List, Test Code, Parallelism).

- Load the appropriate Factor assay reagents (APTT Reagent and APTT Calcium Chloride), Factor Deficient Plasma(s), and Factor Diluent on to the ACL TOP® Family instrument.

- Place QC materials with the barcodes facing out in a Diluent Rack and load onto an ACL TOP® Family instrument in a Diluent track. (If running the QC from the sample rack, refer to Quality Control, Performing a QC Test in the ACL TOP® On-Line Help Manual).

- Choose QC from the Main Menu and select Test Status List.

- Double-click on a test code show Test Materials Definition tree.

- Select the appropriate Factor Assay QC Control and choose Program QC icon. This will run all QC levels for that test.

- Place sample tubes in a sample rack with barcodes facing outwards.

- Select an available sample track and load the sample rack when the barcode reader is in position.

- Verify the samples have been identified and have a test ordered. If not, program the sample ID manually and/or order the test manually from the test and programming window.

- Choose the Run icon if the ACL TOP® Family instrument is not currently running.2

For instructions on loading samples without barcodes or LIS, please refer to Samples Analysis, Managing Patient Samples, Programming Bar Coded Samples and Programming Non-Bar Coded Samples in the ACL TOP®  Family On-Line Help Manual.2

X.	LIMITATIONS

Samples with excessive hemolysis, icterus, or lipemia, should not be used.

Also refer to the SynthASil package insert for relative interference claims.1,6,7,8,9,11

XI.	CALCULATIONS

Patient samples are run at multiple dilutions. The analyzer calculates the % activity for each dilution from a calibration curve. Additional checks are performed including the variance between dilutions and comparison between the slope of the calibration curve and the patient sample dilution curve (parallelism). If the checks are outside prescribed limits analyzer flags are generated and sent to the LIS. In the absence of flags the average of the 50% and 25% sample dilution results (Mean CR) is reported.

XII.	INTERPRETATION

Decreased levels of Factor VIII, IX, XI or XII may be found in congenital deficiency of each Factor or may be acquired secondary to other diseases such as liver disease, Disseminated Intravascular Coagulation (DIC) or due to a specific Factor Inhibitor.1,6,7,8,9

Current laboratory reference range for ACL TOP® Family:

| Factor | Normal Range (% Activity) |
| --- | --- |
| FVIII | 68 – 156 |
| FIX | 92 -161 |
| FXI | 69 - 154 |
| FXII | 57 - 163 |

XIII.	RESULT REPORTING

Analyzer results with no instrument or parallelism errors are auto-verified and released through the LIS. Results with instrument flags are stopped in the LIS interface from reporting. A parallelism report should be printed and reviewed for indications of interference from an inhibitor. Interpretation of the reports should be performed by coagulation faculty, resident, supervisor or technical specialist.

If determined to have an inhibitor present the results from the highest dilution (25% or higher) should be reported with the canned comment “Factor dilutional analysis indicates possible inhibitor”.

  - TRAINING

Staff are trained by a laboratory designated trainer and a training record is completed and signed by both trainer and staff (trainee).

- XV.	REFERENCES

- HemosIL SynthASil (PN 0020006800) package insert

- ACL TOP® Family On-Line Help Manual

- Clinical and Laboratory Standards Institute. Collection, Transport, and Processing of Blood Specimens for Testing Plasma-Based Coagulation and Molecular Hemostasis Assays; Approved Guideline - Fifth Edition, CLSI Document H21-A5; Vol. 28 No. 5

- Westgard JO, and Barry PL. Cost-Effective Quality Control; Managing the Quality and Productivity of Analytical Process, AACC Press, 1986

- Clinical and laboratory Standards Institute.  One Stage Prothrombin Time (PT) Test and Activated Partial Thromboplastin Time (APTT) Test; Approved Guideline – Second Edition, CLSI Document H47-A2; Vol. 28 No. 20

- HemosIL Factor VIII Deficient plasma (PN 0020011800) package insert

- HemosIL Factor IX Deficient plasma (PN 0020011900) package insert

- HemosIL Factor XI Deficient plasma (PN 0020011300) package insert

- HemosIL Factor XII Deficient plasma (PN 0020011200) package insert

- Clinical and Laboratory Standards Institute.  Preparation and Testing of Reagent Water in the Clinical Laboratory, Fourth Edition, CLSI Document C3-A4; Vol. 26 No. 22

Factor Assays, Intrinsic on the ACL TOP 500

Knowledge Check

In the event of a question answered incorrectly:  Single-line through the incorrect answer, initial & date, then select the correct answer.

ALWAYS HAVE CHANGES INITIALED BY YOUR TRAINER.

Circle True or False for each of the following statements.

| 1. | True or False | Factor Assay samples may be processed by centrifuging specimens for 12 minutes at 4000 r.p.m. |
| --- | --- | --- |
| 2. | True or False | All factor assay results will stop and hold on the LIS interface |
| 3. | True or False | All three patient samples dilution results are averaged and reported. |
| 4. | True or False | The slope of the factor assay calibration curve is compared to the patient sample curve. |
| 5. | True or False | Factor inhibitor reports can be interpreted by SMS staff. |

Any incorrect answers I may have initially written have been discussed and corrected.  I now understand the answers I may have gotten wrong.

PASSING GRADE IS 75% OR GREATER

________________________

Employee name (print)

________________________		______________________

Employee signature				(Date)

________________________

Supervisor/Manager name (print)

________________________		______________________

Supervisor/Manager signature			(Date)