TITLE:	 Automated Urinalysis: iRICELL™3000 Procedure

- I.	PURPOSE

  - This procedure provides instructions for the Iris: iRICELL 3000 Automated Urinalysis

II.	PRINCIPLE

The iRICELL™ is an in-vitro diagnostic system. Distinguished by throughput, the work cell includes the iRICELL 3000 (iQ 200 SPRINT™ plus the iChemVELOCITY). The work cell is composed of the Automated Urine Chemistry Analyzer module [The iChemVELOCITY], the iQ Series Automated Urine Microscopy Module [iQ Series], results/analysis processor, computer monitor, mouse and keyboard. The iRICELL provides a fully integrated, automated chemical and microscopic analysis of urine. This procedure provides instructions for performing a complete urinalysis using the iRICELL.

THEORY OF OPERATION:

The iChemVELOCITY performs the chemistry panel, determines the specific gravity, color and clarity of a urine specimen. The chemistry panel is performed using a test strip, which detects the presence of 9 elements – glucose, protein, bilirubin, urobilinogen, pH, blood, ketones, nitrite, and leukocytes by wavelength reflectance. Specific gravity is determined by measuring the refractive index. Color is measured by transmitted light and clarity is measured by scattered light. The iQ Series performs the microscopic portion of the urinalysis and provides a quantitative or qualitative count of formed elements such as cells, casts, crystals, and organisms. The iQ Series photographs particles as they are passed in front of a digital camera. The images are classified, counted and stored for verification by the user. The work cell consists of a computer that is interfaced with an approved chemistry analyzer and the iQ Series modules. At the work cell, results of the chemistry profile and the microscopic analysis are collated, compared to user-defined criteria for auto-release, and stored for auto-release or verified. The majority of specimens can be auto released if user-defined criteria are entered. The user can verify results including the images of the formed elements. As needed, the user may sub-classify or verify results. After verification the results may be sent to the host computer or printed.

PRINCIPLE: iChemVELOCITY URINE CHEMISTRY SYSTEM

The iChemVELOCITY is a urine chemistry analyzer that measures the chemical constituents of the urine using iChemVELOCITY strips, which are read by a dual wavelength reflectance system. The iChemVELOCITY strips consist of a plastic strip containing nine (9) pads impregnated with chemicals specific for the determination of a particular constituent. The nine analytes measured are: glucose, protein, bilirubin, urobilinogen, pH, blood, ketones, nitrite, leukocytes esterase, ascorbic acid and a color compensation pad. A color compensation pad is included on the strip to compensate for the natural color of urine and its effect on the color of the reaction pads. Test strips are placed onto a strip conveyor system by a mechanical extractor. The sample probe mixes the sample, aspirates an aliquot of urine and dispenses it onto each reagent pad. At defined wavelengths, the iChemVELOCITY analyzes the color changes and the intensity of reflected light from the reaction pads. These measurements are used to calculate clinically meaningful results.

PRINCIPLE: iQ200 MICROSCOPY SYSTEM

The microscopic portion of a routine urinalysis is performed on the iQ200 Analyzer. The iQ200 Analyzer auto-identifies and processes specimens by mixing, sampling and analyzing the data obtained from the sample. Approximately 1mL of the mixed specimen is aspirated and is sandwiched between enveloping layers of a suspending fluid. This fluid or “lamina” is positioned exactly within the depth of focus and field of view of the objective lens of a microscope that is coupled to a video camera. The iQ® Lamina is used to position the formed elements in the best orientation that presents the particles with their largest profile facing the direction of view. The camera captures five hundred pictures per sample. The flash of a strobe lamp illuminates each field. The pictures are digitized and sent to the instrument processor. Individual particle images are classified into one of 12 categories using size, shape, contrast and texture. The auto-classified categories are RBCs, WBCs, WBC clumps (WBCC), hyaline casts, unclassified casts (UNCC), squamous epithelial cells, non-squamous epithelial cells (NSE), bacteria, budding yeast, unclassified crystals (UNCX), mucus, and sperm. Any images that do not classify into any one of these 12 categories are placed in the UNCL category. The particle concentration is calculated using the number of images, normalization factor and the volume scanned. User defined criteria for the auto-release of results is checked and results are sent either directly to the host computer and/or printer or to the system monitor for verification. Only non-auto released specimens will appear on the system screen for verification. During the verification process, individual images may be displayed. The operator may manually reclassify digital images when necessary and in accordance with the laboratory’s policy. Once the verification process has been completed and “ACCEPT” has been chosen, the results will be sent to the LIS, screen or printer.

iChemVELOCITY DIPSTICK CHEMICAL REACTIONS:

| Bilirubin | This test is based on the coupling of bilirubin and diazonium salt in an acidic medium. A pinkish tan color proportional to bilirubin concentration is generated. |
| --- | --- |
| Urobilinogen | This test is based on the coupling reaction of urobilinogen with a stable diazonium salt in buffer. A pink to red color proportional to the urobilinogen concentration is generated. |
| Ketone | This test is based on Legal’s method in which the test pad contains sodium nitroprusside and glycine in an alkaline medium. A violet color proportional to methylketone is generated. |
| Ascorbic Acid | This test is based on Tillman’s reaction in which the presence of ascorbic acid leads to the decolorization of the pad from gray-blue to orange. |
| Glucose | This two-step enzymatic reaction uses glucose oxidase, peroxidase and a chromogen. Glucose oxidase catalyzes the formation of gluconic acid and hydrogen peroxide via the oxidation of glucose. Peroxidase then catalyzes the reaction of hydrogen peroxide with a chromogen via the oxidation of chromogen to colors ranging between green and gray-blue. |
| Protein | This test is based on the “protein error of pH indicators” on the green color developed from the presence of protein. This dye-binding is particularly strong with albumin. |
| Blood | This pseudo-enzymatic test contains organic peroxide and a chromogen. The peroxidase effect of hemoglobin and myoglobin causes a color change to green. |
| pH | This test contains a mixed indicator which assures a marked change in color between pH5 and pH9. Colors range from orange through yellow and green to cyan. |
| Nitrite | This test is based on modified Griess reaction in which nitrite in the urine reacts with amide to form a diazonium compound. The subsequent coupling reaction yields a pink color in the presence of nitrite. Some Gram positive and non nitrite-forming bacteria are not detected in this test. |
| Leukocytes | This enzymatic test pad contains an indoxyl ester and a diazonium salt. Granulocyte esterases react with indoxyl ester and diazonium salt to generate a violet color. |

- iii.	SCOPE

- This procedure will be used by the University of Rochester Medical    Center, Clinical and Pathology Laboratories.

- iv.	RESPONSIBILITIES

| Roles | Responsibilities |
| --- | --- |
| Quality | Supports the development of this procedure. Review and approval of this document. |
| Medical Director | Ensure that procedure is followed. Review and approval of this document. |
| Management | Ensure that procedure is followed. Review and approval of this document. |
| Employees | Follows the procedure. |

- V.Specimens

A.	Patient Preparation:  Collect urine in clean or sterile container.  The patient should be given instructions for collecting a “clean catch” urine specimen.  Off site, the specimen should be transferred to a BD preservative tube (Ref# 364992).  At SMH, specimen transfer should be to a 16x100mm round bottom tube.

B.	Specimen Type:  A freshly voided urine sample collected by the “clean catch” method is the specimen of choice.  “clean catch” urine is recommended to prevent the possibility of a positive leukocyte test caused by leukocytes external to the renal-urinary system.

C.	The specimen volume placed on the iRICELL must be at least 4 mL.

If testing on the iChemVELOCITY urine chemistry analyzer alone, the minimum volume is 2 ml.

If testing on the iQ200® urine microscopy analyzer alone, the minimum volume is 3 ml.

D.	Handling:  Specimens should be delivered to the Laboratory as soon after collection as possible.  All specimens should be handled using the principles of Universal Precautions, due to the potential presence of pathogenic material.

Do not add disinfectant or detergent to the specimen.  Do not centrifuge the specimen.

E.	Stability:  Non preservative urine specimens that are kept at room temperature are stable for two hours.  If the specimen is not processed within two hours after collection, cap the container tightly and store at 2 - 8° C.  Non preservative urines that are refrigerated are stable for 4 hours.  Preservative tubes are stable for 72 hours.  Boric acid tubes are stable for 24 hours.

F.	Tracking/ Storage: Patient specimens are tracked to racks labeled yellow, black or white and stored for 3 days in the urine refrigerator (SR015).

- vi.	Quality control

A.	Assay QC once every 24 hours for both chemistry and microscopy modules.

B.	QC material for microscopy module:

iQ® Positive Control

iQ® Negative Control

iQ® Focus reagent.

Upper limits, lower limits and target values are encoded within the -specific barcodes for the Positive, Negative and Focus iQ control material.

C.	QC material for chemistry module:

IRISpec CA/CB/CC controls

D.	Dipsticks:

Sticks are checked for reactivity and accuracy once every 24 hours by running the CA/CB/CC control material.

Controls will also need to be tested on the iChemVELOCITY when a new shipment of Velocity test strips is received.

- Criteria for Acceptable Control Results:

  - Quality control material results must fall within the ranges provided by the manufacturer of the material.

  - Refer to the section below for corrective action(s) if the values fall outside of the published limits.

Note: The system will “lock out” patient testing for microscopy when microscopy controls fail and patient testing for chemistry when chemistry controls fail.

| Step | Action |  |
| --- | --- | --- |
| 1 | - Remove an iChemVELOCITY control rack. - Pour 2 mL of the CA, CB and CC control into 3 separate 16x100mm glass tubes. - Return tightly capped controls to the refrigerator immediately. - Allow aliquots to warm to room temperature. - Use aliquots within one hour of pouring. |  |
| 2 | - Place the CA control in position 8. [color coded on bottle and tube position as RED] - Place the CB control in position 9. [color coded on bottle and tube position as BLUE] - Place the CC control in position 10. [color coded on bottle and tube position as YELLOW] |  |
| 3 | Place the rack on the right side of the iChemVELOCITY sampler touching the edge nearest the operator. Note: The rack will automatically advance forward [blue MEASURE light will appear]. |  |
| 4 | Control results will automatically print on the Quality Review screen as CA, CB and CC. |  |
| 5 | Look at QC results. |  |
|  | QC Failed | - Look at the result printout to see why the control failed. - Adjust positions, if necessary. - Re-pour and re-run all controls. - If results are still not acceptable – check strips loaded in instrument for discoloration. If discolored, discard and add new strips and repeat. - If results are still not acceptable, notify your local distributor, your Iris Product Specialist or Iris’ Clinical Support. - Refer to the Westguard rules chart in the QC log of the urine maintenance book for the correct action log |

| Step | Action |  |
| --- | --- | --- |
| 1 | - Place appropriate barcode labels on 16x100mm glass tubes [Positive, Negative, Focus]. - Mix iQ Positive control and iQ Focus by holding the bottle upside down and giving five hard sharp shakes followed by five gentle inversions. |  |
| 2 | Pour the following into 16x100 mm glass tubes and place in the iQ200 QC rack: - Position 1: 3ml of Iris Cleanser - Position 2: 3ml of Iris Diluent - Position 3: 3ml of Iris Diluent - Position 4: Leave Empty - Position 5: 6 ml of Focus reagent - Position 6: 3 ml of Positive Control - Position 7: 3 ml of Negative Control Control rack must be run immediately after pouring focus and controls. |  |
| 3 | - Place the rack on the iQ200 sampler. - Press the START button located at the upper left corner of the iQ200, if the instrument is in the standby mode (green light). - If in measure mode (blue light), the rack will be detected and processed automatically. |  |
| 4 | Review results under: - Quality Review screen - QC Statistics |  |
|  | QC Failed | - Look at message code to see why the control failed. Note: If control failed due to an identification error or QC out of order, resolve this error. - Pour fresh aliquots and re-run control. - If results are still not acceptable, notify Iris’ Clinical Support. - Refer to the Westguard rules chart in the QC log of the urine maintenance book for the correct action log. |

    - VELOCITY C

- CALIBRATION

  - CalChek Kit

- This material is used for the reflectance check of the strips and for Specific Gravity, color and clarity checks.  Refer to the iChem VELOCITY Maintenance quick reference Guide, Quarterly Maintenance section, for procedure details.

    - iQ200 Module Calibration

- The iQ200 module is calibrated on the first of the month as part of the monthly maintenance schedule.  Refer to the iQ200 Maintenance Quick Reference Guide for procedure details.

- VII.  	special safety precautions

- All patient specimens should be considered potentially infectious and must be handled with precautions used for human blood, as described in CBC (Center for Disease Control) recommendations and in compliance with the Federal OSHA (Occupational Safety and Health Administration) Blood-borne Pathogen Standard, 29 CFR (Code of Federal Regulations) part 1910.1030.  Follow specimen handling as outlined by Laboratory Safety Policy, (SH.CP.AU.0005.0001).

- VIII.	 MATERIALS

| Materials/ Equipment | Consumables and Part Number | Storage | Packaging and Use |
| --- | --- | --- | --- |
| iChemVELOCITY Consumables |  |  |  |
|  | iChemVELOCITY Strips [800-7212] | Room temperature. Stable unopened until expiration date on bottle. Open vial stability=5 days | Bottle of 100 strips; 1 each. Use daily. |
|  | iChem Wash Solution and Wash Filter [800-7704] | Room temperature. Stable unopened until expiration date on bottle. Open bottle stability=3 months | 4 bottles/case; 5.5ml/test. Use daily. |
|  | IRISpec™ CA/CB/CC [800-7702] | Refrigerate 2-8 °C. Unopened stability= date on bottle. Bring poured aliquot to room temperature before use. Open stability= 15 days.  Record open and expiration date on bottle. | 9 bottles (3 bottles of each control); 2 ml of each per day. Use daily. |
|  | iChemVELOCITY CalChek Kit: Strips [800-7703] | Room temperature Stable unopened until expiration date on bottle. Open stability=CalChek strips are single use only. | 2 vials/kit; 1 vial /quarter; single use. Use quarterly. |
|  | iChemVELOCITY CalChek Kit: Reagents [800-7703] | Room temperature Stable unopened until expiration date on bottle. Open tube stability=8 hours | 2 sets/kit; 1 set/quarter; single use. Use quarterly. |

| Materials/ Equipment | Consumables and Part Number | Storage | Packaging and Use |
| --- | --- | --- | --- |
| iQ200 Consumables |  |  |  |
|  | IQ® Lamina [800-3102] | Room temperature. Unopened stability= date on bottle. Change filter when replacing first bottle from a new case of 4. | 4 bottles/cs; 7000 mL/bottle; 14mL/test. Use daily. |
|  | Iris System Cleanser [800-3203] | Room temperature Unopened stability= date on bottle. | 4 bottles/cs 475 mL/bottle; 3ml/test. Use daily. |
|  | Iris Diluent [800-3202] | Room temperature Unopened stability= date on bottle. | 4 bottles/cs 475 mL/bottle; 3 ml/test. Use daily. |
|  | iQ® Calibrator [800-3103] |  Refrigerate 2-8 °C.  Unopened stability= date on bottle. Open bottle stability=24 hours | 4 bottles/cs 125 mL/bottle; 2ml/test. Use monthly. |
|  | iQ® Control/Focus Set [800-3104] |  Refrigerate 2-8 °C. Unopened stability= date on bottle. Open bottle stability=30 days | 1 bottle ea Pos/Neg; 2 bottles of Focus; 125 mL/bottle; Lot-specific Barcodes: Positive control, Negative control; Focus reagent; 3ml of Control/test; 6 ml of Focus/test. Use daily. |
|  | Dilution Code labels [800-3211] | N/A | 1 set; 1 each /test as needed. Use as needed. |
|  |  |  | . |

- IX.	PROCEDURE

- Logging on to the iRICELL:

1.	Access the Instrument Screen.

Click on the Logon button.

2.	Click on the down arrow.

Select your identifier (Do Not type in your identifier). Press tab.

    - Enter your password, click ok.

  - Sample Preparation

    - Place the patient’s bar code label on a sample tube.  Apply it to the tube so that the start of the bar code (not the label edge) is about ½” from the top of the tube.  This leaves room for the dilution label should it be required.

- 2.	 Pour at least 4.0-6.0ml of a well mixed urine specimen into the bar coded tube.

- Tubes 16x100 mm glass or plastic; B-D tiger tops with preservative, (Ref# 364992) B-D yellow with no preservative and gray tops can be run.  Grays must be aliquoted to a 16x100mm tube.

- Note:	Remember to add the comment regarding Boric Acid to samples received in gray top tubes.  Add the comment in results entry, after color, backslash then F5 to bring up the comment field.  Chose @BAT

- 3.	Put the sample tube in a rack so that the bar code is centered between the uprights and facing toward the instrument when the rack is placed correctly o the system

- 4.	Testing can be ordered using the LIS codes

- a.	UAR-urine with reflex to microscopic based on

- a positive leukocyte esterase, blood or protein

- b.	UARM	-urine will always have a microscopic performed

- c.	UARWC-urine with reflex to culture based on

- WBC’s >5/hpf

- Or Bacteria  >1+

- Or Nitrites pos

      - Urobilinogen and Bilirubin are not routinely reported as part of the urinalysis macroscopic profile.  They can be individually ordered by using UUROB or UBILI.

- C.	Velocity and iQ200 Sprint Operation

1.	Place rack(s) on right hand side of the iChemVELOCITY.

Ensure that the notch of the rack base is placed onto the Sampler track ridge.  The instrument will automatically advance the rack.

THE REMAINDER OF THE PROCESSING IS PERFORMED AUTOMATICALLY ON THE SYSTEM:

- After the rack is transferred via the bridge to the iQ Analyzer, the sample rack will be moved along the iQ Sampler to the barcode reader.

- The iQ Analyzer barcode reader reads the specimen barcode.

- If a microscopic examination is to be done (as determined by the user The sample rack will be moved along the sample transport tray to the barcode reader.

- After the barcode is read, the probe mixes the sample, aspirates an aliquot, analyzes the SG, color, clarity and dispenses the sample onto a test strip.

- When the sample processing is complete, the sample rack will be automatically transferred, via the bridge, to the iQ Analyzer.

- If the specimen does not need to be run on the iQ Analyzer, remove the rack from the instrument at this point.  If the rack is to be run on the iQAnalyzer, allow the rack to transfer across the bridge.

- If a microscopic examination is to be done (as determined by the lab defined criteria), the probe will mix the sample, aspirate an aliquot and perform the microscopic examination.  If a microscopic examination is not to be performed, the tube will be passed.

- After sample processing is complete, unload the sample racks from the left side of the Analyzer.

- Verify any pending, flagged results on the Work List as follows in next section of this document.

    - Reviewing Instrument Test Results

- If a specimen has an abnormal microscopic result it is not auto-released to the host computer. The results are to be reviewed at the workstation monitor.

- For a review of the specimens needing to be edited/reviewed, click on “Work List”.

- This brings up the Work List screen that contains all unreleased specimen results.

- At this screen, a specimen may be deleted or undeleted.

- The default list order is time order, oldest first, with any flagged specimens at the top.  The list may be sorted for any parameter by choosing Sort Specimen List or by clicking on the heading at the top of the row.  Clicking a second time will reverse the order; i.e., oldest to newest, newest to oldest or highest specimen number to lowest or vice versa.  The small triangle in the header indicates which header is being used to sort at any time.

- To review a specimen result, double click on it or highlight it then click on “Specimen” at the top of the screen

- The Results screen for that specimen will be displayed.  On the right side are the chemistry results and on the left are the microscopic.

- The microscopic screen from left to right lists the particles, the concentration of the particles and a graphic representation of the particle concentration.

- If the concentration is normal the green bar will display.  If the concentration is abnormal the red bar will display.

- If flags are displayed on the right side of the screen, they must be cleared before any particle type detail can be reviewed.

- Click on “Clear flags” to remove them.  Then click on “Accept”.

- If the only thing preventing auto-release was a flag, the result will be auto-released.  If the specimen result does not meet the lab criteria for auto-release the specimen results are moved to the work list in the time slot that it would have had without the flag and brings up the next specimen.

- It is most efficient to handle all flagged specimens then move on to the timed list.  The cleared flag specimens, if not auto-released, will be encountered again in their proper time order.

- In the specimen screen click on the button of the first particle to be reviewed.

- Images of the particles in that classification will be displayed.

- If the classification of particles is acceptable continue the review by clicking on the right arrow on the right side near the top of the screen.  This takes you to the next set of images.  Clicking on the left arrow takes you to the previous screen.

- Particles not categorized by APR are placed in UNCL (previously ART)

a) Users cannot classify particles into UNCL

b) Users can classify particles from UNCL into any category including ART

r.	The ART category still exists for true artifacts. ART starts with no images.

  - Neither UNCL nor ART are reported to the LIS.

t.	The UNCL category is checked for particles that may have been missed such as casts, if Protein is positive, bacteria, if Nitrates or Leukocyte Esterase is positive. Do not classify every cell back to a category already called such as WBC or Squamous epithelium unless it would make a clinical difference.  The only exception is that all casts must be accounted for.

  - Continue reviewing until you return to the specimen screen.

- If everything is acceptable click on the “accept” button at the bottom right of the screen.  The results will be transmitted to the host computer.

- The iQ will auto-release results if they fall with in the established criteria of: less than or equal to RBC’s - 2, Nucleated cells less than or equal to 5, the presence of squamous cells or mucous.

3.	Sub classifying or Reclassifying Particles:

- In the specimen screen click on the button of the first particle to be reviewed.

- Images of the particles in that category will be displayed.  Note: there may be multiple pages of the same particle type.

- If there are images that are incorrectly classified, they may be reclassified if the reclassification will change the clinical result.  Example 1: you look at WBC and find 24 images of WBC’s with 2 of artifact.  Do not take the time to reclassify the artifact as it will not change the clinical result.

Example 2: you click on BYST and see 6 images of amorphous or artifact.  Click on ART to reclassify the images.  This will remove the BYST category from the report.

- On the right side of the screen, all the categories are listed.  To reclassify, click on the particle type that an image is to be classified into and then click on the image(s) in question.

- This transfers the image to the chosen category.  Note: if you click on an image in error, re-click on the space to return the image to the screen.

- Continue to reclassify by clicking on the category and then clicking on all the images that should go into it.  Much of the time all the particles will be sub-classified or reclassified into the same category, in this case, leaving the screen moves all images on the screen into the category indicated by the lit button on the right.  This saves the time and effort of clicking on the individual images.

- When all images are not being moved to the same category, move the fewest images individually, then choose the category for the rest of the images and leave the screen.  Example: You are reviewing UNCC (unclassified casts).  A couple of the images are of Cellular casts (CELL) and the rest are Granular (GRAN).  Click on the CELL button and select the images that are cellular casts.   Then choose GRAN and leave the screen.

- If the remaining images are to remain in the original category, click on that category at the upper left corner.  This confirms all the images on the screen in their original category.   If you do not, all the images on the screen will go to whichever category is highlighted.

- When everything has been reviewed and edited, return to the specimen screen by clicking on the Results button at the bottom or by continuing through the screens using the right arrow.

Review the results.  After all classification is acceptable, click on the “Accept” button at the bottom right of the screen.

4.	Dilutions

1.	Cloudy and bloody* specimens will need to be diluted before performing testing on the iQ200 module.  Slightly cloudy specimens and specimens with high amorphous may also need to be diluted.  Urine samples that are sent with (less than) <3 ml can also be diluted so there is a sufficient amount of sample for testing.  Remember: You cannot run a diluted specimen on the VELOCITY.

2.	All dilutions are made with Iris Diluent.  If enough sample, perform a 1:2 dilution on specimens containing <3 ml.  Perform a 1:3 dilution on slightly cloudy samples and a 1:6 or 10 on extremely cloudy or bloody samples.  The diluent will dissolve amorphous so you will be able to tell immediately if the cloudy is due to amorphous or cells.  If a sample is viscous or contains large particles, do not run on the iQ but dip and run on the iCHEM and read manually.

| Dilutions must be prepared for the iQ for the following specimens: - Grossly bloody - Dense/Viscous - Short Samples Note: - Dilutions are run on the iQ urine microscopy analyzer only - Do Not perform dilutions for the iChemVELOCITY urine chemistry analyzer. - Identify specimens that require a dilution before placing specimen on the iQ. |  |  |  |
| --- | --- | --- | --- |
|  | Select dilution and corresponding dilution code (refer to Dilutions under Formed Particles or Fluid Type under Settings on the Instrument Screen). |  |  |
|  | You have barcodes | - Label the tube with the appropriate secondary barcode dilution label (fix label below the patient barcode). - Prepare the dilution, in this tube, using Iris Diluent. - Put rack on the iQ analyzer - Press START to run the sample. |  |

- LIMITATIONS

  - **Remember to add the comment “@SBAT (interpret with caution-sample collected with boric acid)” for the gray top tubes. *** Add this comment after color.

  - Specimens exhibiting gross hematuria need to be diluted for the microscopic.  Gross hematuria may cause incorrect results in subsequent samples. Run the sample on the iCHEM and see if it will result.  If not, report the color and clarity and ND the rest of the macro and report as “unable to result due to urine color” (@SDIP).  The urine can then be diluted and run on the iQ or can be spun down and read at the microscope.  A positive result for blood will trigger the microscopic. A “UMIC2” can also be added in Order Entry if necessary.

discarding.

- C.	If a request is made to add on a urine test to a non preserved container, the urine can be done up to 24hr from collection time with the comment “interpret with caution, sample greater than 8 hours old.” If greater than 24 hours, recollect the sample.

D.	Specimens exhibiting gross hematuria, 2+ cloudy or appearing “chunky” cannot be run on the Velocity.  They should be tested manually on the backup iChem analyzer. If the sample needs a microscopic, it should be diluted or spun down and read manually at the microscope.

E.	 INTERFERENCES:

| ANALYTE | CAUSES OF FALSE NEGATIVE RESULTS | CAUSES OF FALSE POSITIVE RESULTS |
| --- | --- | --- |
| Bilirubin | Elevated concentrations of nitrite may inhibit the reaction. Bilirubin is light sensitive and prolonged exposure of urine specimens to light may result in diminished or false negative values | Some urine specimens may contain impurities such as food dyes and therapeutic pigments to produce a yellowish or reddish discoloration of the test pad that may lead to the interference. Elevated Urobilinogen concentrations may slightly enhance the response to this test pad. |
| Urobilinogen | This test is inhibited by elevated concentrations of formaldehyde and nitrite ≥10 mg/dl. Prolonged exposure to light may lead to diminished or false negative values. | Food dyes and medications that have an intrinsic red color in acidic medium such as red beets, azo dyes, phenazopyridine and p-aminobenzoic acid may produce false positive results. |
| Ketones | Elevated concentrations of phenylpyruvic acid may interfere with the test pad and produce a variety of colors. Phthaleins and anthraquinone derivates exhibit a red color in alkaline medium and this may mask the response. Large amounts of levodopa and medications containing sulfhydrl groups may produce atypical color reactions. | N/A |
| Ascorbic Acid | No interferences reported | No interference reported. |

| Glucose | Ascorbic acid concentrations of up to 50 mg/dL did not interfere with glucose assay test results (no false negative results). Acetoacetic Acid concentrations of up to 200 mg/dL did not interfere with glucose assay test results (no false negative results). High specific gravity, acidic pH values and gentisic acid may inhibit color formation. | Cleaning agents such as hypochlorite and peroxide may lead to false positive results. |
| --- | --- | --- |
| Protein | Food dyes such as red beets and therapeutic pigments such as methylene blue and pyridium may mask the coloration of the test pad. Interference may occur with high specific gravity. Interference may also occur with disinfectants, wetting agents and blood substitutes (quaternary ammonium compounds, polyvinylpyrolidone, chlorohexidine). |  |
| ANALYTE | CAUSES OF FALSE NEGATIVE RESULTS | CAUSES OF FALSE POSITIVE RESULTS |
| Blood | Reducing agents such as ascorbic acid, uric acid, glutathione and gentisic acid may cause false negative results. Samples with a pH of 5 may interfere with this test. High concentrations of nitrite can delay the reaction. | Preservatives (formalin) and cleaning agents such as hypochlorite may result in false positives. |
| pH | No interferences reported. | No interferences reported. |
| Nitrite | A negative response in the presence of bacteriuria may be caused by the following: non-nitrite producing microorganisms, low nitrate diet, antibiotic therapy, strong diuresis, or insufficient urinary retention time in the bladder. | Food dyes and therapeutic pigments such as red beets and pyridium may cause false positive responses. |
| Leukocytes | High concentrations of protein, glucose, cephalexin and gentamicin may diminish the color response. The test can be negative in the presence of visible leukocytes if they have not lysed and/or are not granulocytes. | False positive results may occur in the presence of preservatives such as formaldehyde and formalin. Test results may be positive in the absence of observable cells if the granulocytes have lysed. |
| Specific Gravity | N/A-measured by Refractometer | N/A-measured by Refractometer |
| Color | N/A –measured by scattered light | N/A –measured by scattered light |

- CALCULATIONS

Not Applicable

- INTERPRETATION

CLINICAL SIGNIFICANCE OF URINE CHEMISTRY RESULTS:

| Glucose: | The presence of glucose in the urine called glucosuria is caused by hyperglycemia or renal condition. Diabetes mellitus is the most common disease resulting in hyperglycemia. Renal conditions causing dysfunction of tubular re-absorption of glucose occur in many conditions including pregnancy. |
| --- | --- |
| Protein: | The presence of protein in urine is often the first indicator of renal disease, but its appearance in the urine doesn’t always signify renal disease. Although proteinuria may indicate nephritic syndrome, multiple myeloma, glomerulonephritis, and pre-eclampsia, a transient mild proteinuria can be present after exposure to cold, strenuous exercise, high fever, dehydration, or an acute phase of a severe illness. The strip is primarily sensitive to albumin. |
| Bilirubin: | The appearance of urinary bilirubin can be a sign of liver disease or extra-or intra-hepatic biliary obstruction. |
| Urobilinogen: | The normal urine has a small amount of urobilinogen (less than or equal to 2.0 mg/dL). The strip is unable to detect a decreased amount, which may appear in infants, patients on antibiotic therapy, or patients with obstructive disease. Increased amounts appear in hemolytic anemias and liver dysfunction. |
| pH: | Along with the lungs, the kidneys are the major regulator of acid-base balance. Freshly voided urine has a pH of 5.0 – 6.0. The pH of urine can be controlled by dietary regulation and medication. |
| Blood: | A positive reaction for blood may indicate red cells, hemoglobin, or myoglobin present in the urine. Hematuria can be seen due to bleeding as a result of trauma or irritation (renal calculi, glomerulonephritis, tumors, toxic or chemical exposure). Hemoglobinuria occurs when there is lysis of red cells in the urinary tract, intravascular hemolysis or transfusion reactions. Very dilute or extremely alkaline urine can also lyse the cells. Myoglobinuria indicates muscular destruction that may appear in hypothermia, convulsions, and extensive exertions. |
| Ketones: | Ketonuria appears when there is an increased use of fat instead of carbohydrate as a source of metabolism. Conditions of ketonuria include diabetes mellitus, vomiting, inadequate intake of carbohydrates due to starvation, weight reduction, or pregnancy. |

| Nitrite: | Bacteria, specifically gram negative organisms, are detected by this nitrite reducing reaction. In order for the reaction to take place there must be adequate dietary nitrates, and the urine must be in the bladder at least four hours for the bacteria to react with nitrate for a positive reaction. Unusually colored urine due to medication or dyes can interfere with this reaction. |
| --- | --- |
| Leukocytes: | The presence of white blood cells in the urine is an indicator of inflammation. Lysed and intact WBCs are detected because both may have produced esterase. |
| Specific Gravity: | Specific gravity is a measure of the dissolved substances present in the urine. Specific gravity is one measure of the concentrating and diluting ability of the kidneys and hydration status of the patient. The specific gravity is obtained by measuring the refraction angles of light passing through a triangle prism. An LED emits a beam of light through a slit and a lens. The refractive index changes according to the specific gravity of the sample, the higher the specific gravity the greater is the angle of measurement. The change in the angle of the light is reported as the specific gravity. The result is automatically corrected for elevated protein or glucose concentrations as measured by the test strip. |
| Color: | Color variation can indicate the presence of a disease process, metabolic abnormality, or an ingested food or drug, or the variation simply due to excessive physical activity or stress. The color of the specimen is measured by transmitted light. The colors obtained are colorless, yellow, orange, brown, red, violet, blue, green and other, including light and dark of each. |

| Clarity: | Substances that cause urine turbidity may be pathologic or non-pathologic. The clarity or turbidity of a urine specimen is measured by passing a beam of light through the sample and measuring how the light is scattered. The amount of scattered light increases as the specimen becomes more turbid. The amount of clarity is reported as clear, turbid or extremely turbid. |
| --- | --- |

- RESULT REPORTING

A normal sample that does not need a microscopic examination will be transmitted to the host computer and will auto verify

.	A specimen with pathologic findings such as Trichomonas or cystine, leucine or tyrosine crystals, will not auto verify but remain on the iQ interface.  This allows the user to add a comment that these findings were reviewed manually.

If the sample is abnormal or meets the criteria requiring a microscopic examination, the results can be reviewed at the workstation along with the microscopic findings.

There are three comments that pertain to macro or sample problems.  These are found after color, backslash then F5 to bring up the comment screen

a. 	@SBAT-canned comment for gray top with Boric Acid preservative

b.	@SDIP-canned comment refers to urine color interference for the dipsticks.

      - SUCN-canned comment pertaining to fecal contamination of the sample

REPORTING OF RESULTS: - Glucose is reported as NORM,50,150,>=500 mg/dl; protein as NEG,30,100,>=500 mg/dl; ketones are graded as NEG, trace, 1+,2+ - Bilirubin is graded as NEG, 1+,2+. - Leukocytes are graded as NEG, trace, 1+,2+,3+. - Urobilinogen is graded as NORM, 1+,2+. - pH is reported in quantitative units. 5,6,7,8,9 - Blood is graded as NEG, 1+,2+,3+. - Nitrite is reported semi-quantitatively as NEG,POS. - Specific Gravity is reported by refractive index quantitatively with a value to 3 decimal places, ranging from 1.000 to >1.060. in 0.001 increments. - Clarity is reported as clear, cloudy 1+, cloudy 2+, cloudy 3+. - Color is reported as a color: colorless, straw, yellow, amber, brown, red, blue” by the iChem VELOCITY.

EXPECTED RESULTS:

| Bilirubin | In normal urine, no detectable level of bilirubin should be obtained. Positive results require further investigation. Performance Characteristics: The test is specifically developed for bilirubin. Biliverdin does not react with this test pad. |
| --- | --- |
| Urobilinogen | In normal urine, urobilinogen is usually present at concentration up to 1 mg/dl. A result of 2 mg/dl represents the transition from normal to abnormal state and the specimen should be further investigated for possible liver disease and hemolytic disorders. Performance characteristics: The diazonium-based test is more specific for urobilinogen than Ehrlich’s reagent based- test. Test strips cannot determine the absence of urobilinogen, which may be significant in biliary obstruction |
| Ketone | Detectable amounts of ketone do appear in the urine of normal specimens. Positive ketone values may result from the following conditions: starvation, dietary imbalance, diabetes mellitus, eclampsia, insulin dosage monitoring, vomiting and other metabolic disorders. Performance Characteristics: The test does not measure B-hydroxybutric acid and is only slightly sensitive to acetone. |
| Ascorbic Acid | Ascorbic acid is found in various food supplies and dietary supplements. Concentrations of ascorbic acid greater than 20 mg/dl can be expected to cause strong interference with glucose, blood, and nitrite. Performance Characteristics: The oxidized form, dehydroascorbic acid, does not react with this test pad. |
| Glucose | A small amount of glucose (up to 20 mg/dl) may be present in normal urine. The detectable sensitivity of this test has been adjusted to exclude the minute amount of glucose. Therefore, any positive response should be further evaluated. Performance Characteristics: The test pad doe not react with other reducing sugars such as fructose and galactose. |
| Blood | Normal urine contains no detectable hemoglobin or intact red blood cells. Any positive results should be further evaluated. Performance Characteristics: 0.015 mg/dl blood |
| pH | The normal pH of urine can vary between pH 4.5and pH 8.0. Performance Characteristics: pH values are determined to within 0.5 unit over the range from 5.0 to 9.0. |
| Nitrite | Normal urine contains no detectable nitrite. However, a negative result does not rule out a urinary tract infection. Performance Characteristics: This test is specific for nitrite. Results may depend on the ability of bacteria to reduce nitrate to nitrite, the number of bacteria and the retention time of the urine in the bladder. |
| Leukocytes | A normal urine specimen should not produce a positive result. The test can be negative in the presence of observable leukocytes if they are not lysed/and or are not granulocytes (example:lymphocytes). The test may be positive in the absence of observable cells if the granulocytes have lysed. Performance Characteristics: Leukocyte test detects the presence of esterase in the granulocytic white blood cells. The test result is most frequently accompanied by the presence of bacteria that may or may not produce a nitrite positive reaction. |

- TRAINING

| Role | Training Needed |
| --- | --- |
| Management | Read |
| Designated Employees | Knowledge Check |

- REFERENCES

    - Iris automated Urinalysis iRICELL Procedure Template, Rev A 300-5347 5/2012

    - Iris Operators manual, Rev B 11/2/2011

XVI.	REFERENCE RANGES:		UNITS

## Chemistry Results

Glucose			Neg

Protein, Qualitative		Neg

Bilirubin			Neg

Urobilinogen			Norm

PH				5.0 – 8.0

Blood				Neg

Ketones			Neg

Nitrite				Neg

Leukocyte Esterase		Neg

Clarity				Clear

Specific Gravity		1.002 – 1.030

Color				Yellow

## Microscopy Results

WBC				M/F -0-5	/HPF

RBC				M/F- 0-2	/HPF

Bacteria 			None

Epithelial Cells

Squamous 			0-3+

Transitional 			0-1+

Renal 				None

# Casts

Hyaline Casts		0 - 2	/LPF

Granular Casts		0	/LPF

Cellular Casts		0	/LPF

WBC Cast		0	/LPF

RBC Cast			0	/LPF

Waxy Cast			0	/LPF

Fatty Cast			0	/LPF

Epi Cell Cast			0	/LPF

# Crystals

Amorphous Crystals	None

Calcium Oxalate 		None

Leucine Crystals		None

Cystine Crystals		None

Tyrosine Crystals		None

Miscellaneous Particles

Yeast				None

WBC Clumps		None

Oval Fat Body		None

Trichomonas			None

Mucus				None