# 1.0	Synonym

# High Density Lipoprotein, Good Cholesterol.

# 2.0	Clinical Significance

# Many epidemiological investigations have demonstrated the strong and independent inverse association between HDL-Cholesterol and the risk of coronary artery disease.1, 2 It has been proposed that HDL particles, through the uptake and transport of Cholesterol from peripheral tissue to the liver (reverse Cholesterol transport), protects against the development of atheromatous plaques.

# The guidelines issued by The National Cholesterol Education Program Adult Treatment Panel 3 (NCEP ATP 3),4 recommends lipoprotein analysis (total cholesterol, LDL cholesterol, HDL cholesterol and triglycerides) as the preferred initial test, rather than screening for total cholesterol and HDL alone.

# In 2001, the NCEP increased the high-risk medical decision point to <40 mg/dL.

# The guidelines classify HDL- C levels as follows:

# 1.  < 40 mg/dL as indicative of a major risk factor for Coronary Heart Disease.

# 2.  ≥ 60 mg/dL as a negative risk factor for Coronary Heart Disease.

# 3.0	Principle

# 3.1	This HDL cholesterol test is a two reagent homogenous system for the selective measurement of serum or plasma HDL cholesterol in the presence of other lipoprotein particles. The assay is comprised of two distinct phases. In phase one, free cholesterol in non-HDL-lipoproteins is solubilized and consumed by cholesterol oxidase, peroxidase, and DSBmT to generate a colorless end product. In phase two, a unique detergent selectively solubilizes HDL. The HDL cholesterol is released for reaction with cholesterol esterase and cholesterol oxidase, in the presence of chromogens, to produce a colour product.

# 3.2	The HDL reagent measures the HDL cholesterol concentration by a timed-endpoint method. The system automatically proportions the appropriate HDL cholesterol sample and reagent volumes into a cuvette. The ratio used is one part sample to 93 parts reagent. The system monitors the change in absorbance at 560 nanometers. This change in absorbance is directly proportional to the concentration of cholesterol in the sample and is used by the system to calculate and express the HDL cholesterol concentration.

# 3.3	Chemical Reaction Scheme

# 4.0	Specimen

# 4.1	Type of Specimen

# Biological fluid samples should be collected in the same manner routinely used for any laboratory test.  Freshly drawn serum or plasma are the preferred specimens.

# 4.2	Specimen Storage and Stability

# 4.2.1	Tubes of blood are to be kept closed at all times and in a vertical position. It 			is recommended that the serum or plasma be physically separated from 				contact with cells within two hours from the time of collection.

# 4.2.2	Use fresh samples for analysis when possible. HDL cholesterol is stable in serum and plasma for 8 hours at +15ºC to +30ºC and 7 days at +2°C to +8°C.  If longer storage is required, samples can be frozen for up to 30 days at  ≤ -20ºC. Frozen samples should be thawed only once.

# 4.3	Sample Volume

# The optimum volume, when using a 0.5 mL sample cup, is 0.3 mL of sample.

# 4.4	Criteria for Unacceptable Specimens

# Refer to the “Type of Specimen”, “Limitation” and “Interference’ sections of this 			procedure for information on unacceptable specimens.

# 5.0	Reagents

# 5.1	Each kit contains the following items:

# Two HDL Reagent Cartridges (2 x 200 tests)

# 5.2	Volumes per Test

# Table 1

| # Sample Volume |  | # 3 mL |
| --- | --- | --- |
| # Total Reagent Volume |  | # 280 mL |
| # Cartridge Volumes |  |  |
|  | # A | # 210 mL |
|  | # B | # 70 mL |
|  | # C | # – – |

# 5.3	Reactive Ingredients

# Table 2

| # Reagent Costituents |  |
| --- | --- |
| # Cholesterol esterase (Pseudomonas) | # 375 U/L |
| # Cholesterol oxidase (E. coli) | # 750 U/L |
| # Peroxidase (Horseradish) | # 5000 U/L |
| # Ascorbate oxidase (Cicurbita sp.) | # 2250 U/L |
| # DSBmT | # 0.75 mmol/L |
| # 4-aminoantipyrine | # 0.25 mmol/L |
| # Detergent | # 0.375% |
| # Preservative | # 0.05% |
| # Also non-reactive chemicals necessary for optimal system performance. |  |

# 5.4	Caution

## GHS HAZARD CLASSIFICATION

| HDL Cholesterol Reagent (Compartment A) | DANGER |  |
| --- | --- | --- |
|  | H317 | May cause an allergic skin reaction. |
|  | H334 | May cause allergy or asthma symptoms or breathing difficulties if inhaled. |
|  | P261 | Avoid breathing vapours. |
|  | P280 | Wear protective gloves, protective clothing and eye/face protection. |
|  | P284 | In case of inadequate ventilation, wear respiratory protection. |
|  | P304+P340 | IF INHALED: Remove person to fresh air and keep at rest in a position comfortable for breathing. |
|  | P342+P311 | If experiencing respiratory symptoms: Call a POISON CENTER or doctor/physician. |
|  | P362+P364 | Take off contaminated clothing and wash it before use. |
|  |  | Peroxidase 0.1 - 1% |
|  |  | reaction mass of: 5-chloro-2-methyl-4-isothiazolin-3-one [EC# 247-500-7] and 2-methyl-4-isothiazolin-3-one [EC# 220-239-6](3:1) < 0.05% |

| HDL Cholesterol Reagent (Compartment B) | WARNING |  |
| --- | --- | --- |
|  | H317 | May cause an allergic skin reaction. |
|  | P280 | Wear protective gloves, protective clothing and eye/face protection. |
|  | P333+P313 | If skin irritation or rash occurs: Get medical advice/attention. |
|  | P362+P364 | Take off contaminated clothing and wash it before use. |
|  |  | reaction mass of: 5-chloro-2-methyl-4-isothiazolin-3-one [EC# 247-500-7] and 2-methyl-4-isothiazolin-3-one [EC# 220-239-6](3:1) < 0.05% |

|  | Safety Data Sheet is available at techdocs.beckmancoulter.com |
| --- | --- |

# 5.5	Materials Needed but not Supplied with Reagent Kit

# 5.5.1	UniCel DxC SYNCHRON Systems Calibrator

# 5.5.2	Deionized Water

# 5.5.3	At least two levels of control material

# 5.5.4   Saline

# 5.6	Reagent preparation

# No preparation is required.

# 5.7	Acceptable Reagent Performance

# The acceptability of a reagent is determined by successful calibration and by 			ensuring that quality control results are within acceptance criteria.

# 5.8	Reagent Storage and Stability

# HDL Cholesterol Reagent when stored unopened at +2°C to +8°C, will attain the 			shelf-life indicated on the cartridge label. Once opened, the reagent is stable for 60 		days at +2°C to +8°C unless the expiration date is exceeded. DO NOT FREEZE.

# 6.0	Instrument and Equipment

# Beckman Coulter DxC .

# 7.0	Procedure

# 7.1	Calibration

# 7.1.1	Calibrator Required

# UniCel DxC SYNCHRON Systems HDL Calibrator PN B23634

Deionized water ( Level 1 calibrator)

# 7.2	Calibrator Preparation

7.2.1 	Use deionized water as the Level 1 calibrator.

7.2.2	Remove the metal cap around the HDL calibrator bottle and gently tap the bottle on the table to remove powder at the top of the stopper. Open the HDL calibrator bottle carefully, avoiding loss of lyophilizate.

7.2.3	Add exactly 1.00 mL of deionized water to the bottle of calibrator. Replace the stopper and let stand for 5 minutes at room temperature.

7.2.4	Gently invert until the contents are dissolved avoiding the formation of foam. DO NOT SHAKE.

# 7.3	Calibrator Storage and Stability

The unopened, UniCel SYNCHRON Systems HDL Calibrator may be stored at +2ºC to +8ºC until the expiration date printed on the label. Reconstituted calibrators that are resealed are stable for 14 days at +2ºC to +8ºC or for 30 days at ≤ -20ºC unless the expiration date is exceeded.  Frozen calibrator should be thawed only once.

Visible signs of microbial growth, gross turbidity, or precipitate in the calibrator may indicate degradation and warrant discontinuation of use.

# 7.4	Caution

# Because this product is of human origin, it should be handled as though 				capable of transmitting infectious diseases.  Each serum or plasma donor

# unit used in the preparation of this material was tested by United States 				Food and Drug Administration (FDA) approved methods and found to be 				negative for antibodies to HIV and HCV and nonreactive for HbsAg.  				Because no test method can offer complete assurance that HIV, hepatitis B 			virus, and hepatitis C virus or other infectious agents are absent, this 				material should be handled as though capable of transmitting infectious 				diseases.  This product may also contain other human source material for 				which there is no approved test.  The FDA recommends such samples to be 			handled as specified in Centers for Disease Control's Biosafety Level 2 guidelines.

# 7.1.5	Calibration Information

# 7.1.5.1	The system must have valid calibration factors in memory before controls or 			patient samples can be run.

# 7.1.5.2	Under typical operating conditions the HDL assay must be calibrated every 28 days or with each new cartridge of reagent and also with certain parts replacements or maintenance procedures, as defined in the Unicel DxC 600/800 System Instructions for Use (IFU) manual. Alternatively, this assay has within-lot calibration available. Refer to the UniCel DxC 600/800 System Instruction for use for information on this feature.

# 7.1.5.3	For detailed calibration instructions, refer to the UniCel DxC 600/800 Instructions For Use (IFU) manual.

# 7.1.5.4	The system will automatically perform checks on the calibration and produce 			data at the end of calibration. In the event of a failed calibration, the data will 			be printed with error codes and the system will alert the operator of the 				failure. For information on error codes, refer to the UniCel DxC 600/800

System Instructions For Use (IFU) manual.

# 7.2	Testing Procedures

# 7.2.1	If necessary, load the reagent onto the system as directed in the  					Operations Manual.

# 7.2.2	After reagent load is completed, calibration may be required. Refer to the  				Operations Manual for details of the calibration procedure.

# 7.2.3	Program samples and controls for analysis as directed in the Operations 				Manual.

# 7.2.4	After loading samples and controls onto the system, follow the protocols for 			system operation as directed in the UniCel DxC 600/800 Instructions For		 		Use (IFU) manual.

# 8.0	Quality Control

# At least two levels of control material should be analyzed daily. In addition, these controls 	should be run with each new calibration, with each new reagent cartridge, and after specific 	maintenance or troubleshooting procedures as detailed in the SYNCHRON Operations 	Manual.  At this facility, the Biorad Chemistry Control levels 1 and 2 are used.  The controls 	should be prepared and used in accordance with the package inserts.

# 9.0	Calculations

# The UniCel  DxC 600/800 System performs all calculations internally to produce the final 	reported result. The system will calculate the final result for sample dilutions made by the 	operator when the dilution factor is entered into the system during sample programming.

# 10.0	Reporting Results

# 10.1	Reference Intervals (ranges)

# The reference interval listed below was taken from a Beckman study performed on 		SYNCHRON Systems and verified at this facility.

# Table 3:	 Reference Intervals

| # INTERVALS | # CARDIOVASCULAR RISK | # CONVENTIONAL UNITS | # S.I. UNITS |
| --- | --- | --- | --- |
| # SYNCHRON | # Low | # ≥ 60 mg/dL | # ≥ 1.55 mmol/L |
|  | # High | # < 40 mg/dL | # < 1.03 mmol/L |

# 11.0	Limitation

# Samples or control materials which contain acetic acid, detergents, or surfactants may 	inhibit the enzymes in the reagent and should not be used.

# 12.0	Interferences

# 12.1	The following substances were tested for interference with this methodology:

# Table 4 Interferences

| # Substances | # Source | # Maximum Level tested | # Observed Effect |
| --- | --- | --- | --- |
| # Hemoglobin | # RBC hemolysate | # 500 mg/dL | # Not Significant (NSI) |
| # Bilirubin | # Porcine | # 40 mg/dL | # NSI |
| # Triglyceride | # Human | # 900 mg/dL | # NSI |
| # Lipemia | # Intralipid | # 1500 mg/dL | # NSI |
| # Ascorbic Acid | # NA | # 20 mg/dL | # NSI |
| # Immunoglobulin IgG | # Human | # 5000 mg/dL | # NSI |

# 12.2	Inaccurate results (usually negative interference) may be produced in patient samples with elevated serum immunoglobulin levels.Error: Reference source not found In very rare cases gammopathy, especially monoclonal IgM (Waldeström’s macroglobulinemia), may cause unreliable results.

# 12.3	Inaccurate results (usually negative interference) may be produced in patients with 			elevated serum immunoglobulin levels.

# 12.4	Falsely low results may be obtained in patients with Type III hyperlipidemia.

# 13.0	Analytic Range (linearity)

# 13.1	The DxC 600/800 System method for the determination of this analyte 				provides the following analytical ranges:

# Table 5:	 Analytical Range

| # Sample Type | # Conventional Unit |
| --- | --- |
| # Serum or Plasma | # 5 – 135 mg/dL |

# 13.2	Samples with concentrations exceeding the high end of the analytical range should 		be diluted with saline and reanalyzed.

# 14.0	Detection Limit

# Limit of blank (LoB), limit of detection (LoD), and limit of quantitation (LoQ) data analysis was performed in accordance with the CLSI EP17-A2 guideline.21 The LoB corresponds to the concentration below which analyte-free samples are found with 95% confidence. The LoD corresponds to the sample concentration above the LoB which is detectable with 95% confidence. The LoQ is defined as the lowest amount of analyte in the sample that can be quantitatively determined with stated acceptable precision and trueness, under stated experimental conditions. A properly operating UniCel DxC 600/800 System should exhibit detection limit values equal to the following:

# The detection limit results using serum samples support the LoB, LoD, and LoQ specifications in the table below using the HDL reagent on a UniCel DxC 600/800 System.

Table 6 DETECTION LIMIT & RESULTS

|  | Specifications | Results |
| --- | --- | --- |
| LoB | ≤ 3 mg/dL (0.08 mmol/L) | 0.43 mg/dL (0.01 mmol/L) |
| LoD | ≤ 5 mg/dL (0.13 mmol/L) | 0.66 mg/dL (0.02 mmol/L) |
| LoQd | ≤ 5 mg/dL (0.13 mmol/L) | 3.26 mg/dL (0.08 mmol/L) |

# 15.0 	Equivalency

Equivalency was assessed by Deming regression analysis of patient samples to accepted clinical methods.

|  | Y (UniCel DxC 800 System) | = 1.015X - 0.9 |
| --- | --- | --- |
|  | Serum (in the range of 8.1 to 123.7 mg/dL): | = 131 |
|  | MEAN Y (UniCel DxC 800 System) | = 56.0 |
|  | MEAN X (AU680 OSR6x95) | = 56.1 |
|  | CORRELATION COEFFICIENT (r) | = 0.999 |

Performance data obtained on the UniCel DxC 800 System is representative of the data obtained on the UniCel DxC 600 System.

# 16.0	Precision

# 16.1	A properly operating UniCel DxC 600 System should exhibit precision values 			less than or equal to the following:

# Table 7 Precision Values

| # Type of Precision | # Sample Type | # 1 S.D. # mg/dL | # Changeover value # mg/dL | # % CV |
| --- | --- | --- | --- | --- |
| # Within-run | # Serum/Plasma | # 1.3 | # 42 | # 3.0 |
| # Total | # Serum/Plasma | # 1.7 | # 42 | # 4.0 |

DxC 600/800 System evaluated using the CLSI Approved Guideline EP5-A2 appears in the table below.Error: Reference source not found Each laboratory should characterize their own instrument performance for comparison purposes.

- Table 8 CLSI EP5-A2 Precision Estimate Method

| TYPE OF IMPRECISION | SAMPLE TYPE | No. Systems | No. Data Pointsf | Test Mean Value (mg/dL) | EP5-A2 Calculated Point Estimates |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | SD | % CV |
| Within-run | Human Sample 1 | 1 | 80 | 33.3 | 0.3 | 0.9 |
|  | Human Sample 2 | 1 | 80 | 56.1 | 0.6 | 1.1 |
|  | Human Sample 3 | 1 | 80 | 121.0 | 1.2 | 1.0 |
| Total | Human Sample 1 | 1 | 80 | 33.3 | 0.6 | 1.8 |
|  | Human Sample 2 | 1 | 80 | 56.1 | 1.0 | 1.8 |
|  | Human Sample 3 | 1 | 80 | 121.0 | 1.6 | 1.3 |

# 16.2	These degrees of precision and equivalency were obtained in typical testing procedures on the UniCel DxC 600/800 SYNCHRON System  and are not intended  to represent the performance specifications for this reagent.

# 16.3	For details of precision studies done at this facility, refer to instrument validation 			data.

# 16.4	For more detailed information on the UniCel DxC 600 System, refer to the appropriate SYNCHRON System manual.

# 17.0	Backup Procedure

# The DxC600i and DxC600 serve as backup for each other.    For technical assistance, call Beckman Coulter at 800-854-3633.

# 18.0	References

# 18.1	Beckman Coulter UniCel DxC 600/800 Instructions for Use (IFU) manual.		.

# 18.2	Beckman Coulter UniCel DxC 600/800 Operations Manual

18.3	SYNCHRON Systems Chemistry Information Sheet HDL Cholesterol.

# 18.4	Tietz, N. W., "Specimen Collection and Processing; Sources of Biological Variation", 		Textbook of Clinical Chemistry, 5th Edition, W. B. Saunders, Philadelphia, PA 			(2001).

# 18.5	Tietz, N. W., "Specimen Collection and Processing; Sources of Biological Variation", 		Textbook of Clinical Chemistry, 5th Edition, W. B. Saunders, Philadelphia, PA 			(2001).

DOCUMENT HISTORY PAGE

| Change type: New, Major, Minor etc. | Changes Made to SOP. (Describe) | Written by/ Revised by and Date | Medical Director &Signature | Date Approved | Date Change Implemented |
| --- | --- | --- | --- | --- | --- |
| New | N/A | Thomas Christopher 06/08/16 |  |  |  |