- Synonyms

TSH, Thyrotropin, hypersensitive hTSH, third generation TSH.

- Clinical Significance

Human thyroid-stimulating hormone is a glycoprotein hormone consisting of two noncovalently-bound subunits: an α subunit, which is nearly identical to the α subunits of human luteinizing hormone (hLH), human follicle-stimulating hormone (hFSH), and human chorionic gonadotropin (hCG), and a β subunit, which is responsible for immunological and biological specificity.

TSH, released from the anterior pituitary, is the principal regulator of thyroid function, stimulating the synthesis and release of thyroid hormones thyroxine (T4) and triiodothyronine (T3). T3 and T4 regulate biochemical processes that are essential for normal metabolism. The synthesis and secretion of TSH is stimulated by thyrotropin-releasing hormone (TRH), which is produced by the hypothalamus in response to low levels of circulating T3 and T4. In contrast, elevated levels of T3 and T4 suppress the production of TSH. Collectively, this negative feedback system is referred to as the hypothalamic-pituitary-thyroid axis. Any alteration in the function of this axis can influence the levels of TSH, T4, and T3 in circulation.

The principal clinical use for TSH measurement is for the assessment of thyroid status. TSH is measured in conjunction with thyroid hormones or antibodies to: 1) detect or exclude hypothyroidism or hyperthyroidism; 2) monitor T4 replacement treatment in hypothyroidism or antithyroid treatment in hyperthyroidism; 3) monitor TSH suppression in thyroid cancer patients on thyroxine therapy; and 4) assess the response to TRH stimulation testing.

Reference ranges for TSH may vary, depending on the method of analysis, and do not necessarily equate to cut-offs for diagnosing thyroid dysfunction.

As methods achieving third-generation sensitivity have become available, TSH measurements are also used to identify subclinical or latent hypothyroidism or hyperthyroidism. Third generation methods for TSH measurement are defined as achieving a functional sensitivity of 0.01-0.02 µIU/mL (mIU/L) with an interassay %CV ≤ 20%.4,5 Clinical laboratories may use these more sensitive and precise methods to distinguish different levels of TSH suppression associated with Graves’ disease and subclinical hyperthyroidism and to assist in the diagnosis of gestational and postpartum thyroid diseases.

3.0	Principle

3.1	The Access TSH (3rd IS) assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of human thyroid-stimulating hormone (thyrotropin, TSH, hTSH) levels in human serum and plasma using the Access Immunoassay Systems. This assay is capable of providing 3rd generation TSH results.

3.2	The Access TSH (3rd IS) assay is a two-site immunoenzymatic (“sandwich”) assay. A sample is added to a reaction vessel with mouse anti-hTSH-alkaline phosphatase conjugate, buffered protein solution and paramagnetic particles coated with immobilized mouse monoclonal anti-hTSH antibody. The hTSH binds to the immobilized monoclonal anti-hTSH antibody on the solid phase while the mouse anti-hTSH-alkaline phosphatase conjugate reacts with a different antigenic site on the hTSH.

3.3	After incubation in a reaction vessel, materials bound to the solid phase are held in a magnetic field while unbound materials are washed away. Then, the chemiluminescent substrate Lumi-Phos 530 is added to the vessel and light generated by the reaction is measured with a luminometer. The light production is directly proportional to the concentration of TSH in the sample. The amount of analyte in the sample is determined from a stored, multi-point calibration curve.

- Specimen

4.1	Serum (gel and no gel) and plasma (lithium heparin) are the recommended samples.  Collect all blood samples observing routine precautions for venipuncture.  Allow serum samples to clot completely before centrifugation.   Keep tubes stoppered at all times.

4.2	Within two hours after centrifugation, transfer at least 500L of cell-free sample to a storage tube. Tightly stopper the tube immediately.   Store samples tightly stoppered at room temperature (15 to 30°C) for up to 18 hours.

4.3	If the assay will not be completed within 18 hours, refrigerate the samples at 2 to 8°C.  If the assay will not be completed within seven days, freeze at -20°C or colder. Frozen specimens can be stored for 90 days before testing.

4.4	Do not thaw samples more than two times.

4.5	Use the following guidelines when preparing specimens, unless instructed otherwise in the product insert:  Ensure residual fibrin and cellular matter has been removed prior to analysis.

4.6	Follow blood collection tube manufacturer’s recommendations for centrifugation.

Variations in these products may exist between manufacturers and, at times, from lot to lot. Avoid assaying lipemic or hemolyzed samples.

- Reagent

  - Access TSH (3rd IS) Reagent Pack Cat. No. B63284: 200 determinations, 2 packs, 100 tests/pack.

R1a:	Paramagnetic particles coated with mouse monoclonal anti-human TSH antibody suspended in TRIS buffered saline, with surfactant, bovine serum albumin (BSA), < 0.1% sodium azide, and 0.1% ProClin** 300.

R1b	TRIS buffer saline with surfactant, BSA, protein (murine), < 0.1% sodium azide, and 0.1% ProClin 300.

R1c:	Mouse monoclonal anti-human TSH alkaline phosphatase conjugate in ACES buffered saline, with surfactant, BSA matrix, protein (murine), < 0.1% sodium azide, and 0.25% ProClin 300.

R1d:	Mouse monoclonal anti-human TSH alkaline phosphatase conjugate in ACES buffered saline, with surfactant, BSA matrix, protein (murine), < 0.1% sodium azide, and 0.25% ProClin 300.

  - Reagent Storage

    - Store upright and refrigerate at 2 to 10°C.

    - Refrigerate at 2 to 10°C for a minimum of two hours before use.

    - Stable until the expiration date on the label when stored at 2 to 10°C.

    - Stable at 2 to 10°C for 28 days after initial use.

5.2.5	Signs of possible deterioration are a broken elastomeric layer on the pack or control values out of range. If the reagent is damaged, discard the pack.

5.3	Note:	All antisera are polyclonal unless otherwise indicated.

5.4	Warnings and Precautions

5.4.1	For in vitro diagnostic use.

5.4.2	Patient samples and blood-derived products may be routinely processed with minimum risk using the procedure described. However, handle these products as potentially infectious according to universal precautions and good clinical laboratory practices, regardless of their origin, treatment, or prior certification. Use an appropriate disinfectant for decontamination.

    - Store and dispose of these materials and their containers in accordance with local regulations and guidelines.

    - Sodium azide may react with lead and copper plumbing to form highly explosive metal azides. On disposal of liquids, flush with a large volume of water to prevent azide build-up.

    - ProClin 300 is a potential skin sensitizer. Avoid spilling or splashing this reagent on skin or clothing. In case of contact with the reagent, flush thoroughly with soap and water.

  - Materials Required But Not Provided with Reagent Kit

5.5.1	Access TSH (3rd IS) Calibrators. Provided at zero and approximately 0.050, 0.30, 3.0, 15.0 and 50.0 µIU/mL  (mIU/L). Cat. No. B63285.

    - Quality Control (QC) materials.

5.5.3	Access Substrate Cat. No. 81906

5.5.4	Access Wash Buffer

6.0	Instrument and Equipment

Beckman Coulter Access (stand-alone model or part of the integrated system DxCi).

7.0	Procedure

  - Calibration

    - Calibration Information

An active calibration curve is required for all tests. For the Access TSH (3rd IS) assay, calibration is required every 28 days. Refer to the appropriate system manuals and/or Help system for information on calibration theory, configuring calibrators, calibrator test request entry, and reviewing calibration data.

7.2	 Testing Procedures

# 7.2.1 	Refer to the Instruction for Use Manual and Access Assay Manual for 				information on managing samples, configuring tests, requesting tests, 				principles of operation, system performance characteristics, operating 				instructions, calibration procedures, operational limitations and precautions, 			hazards, maintenance, troubleshooting and reviewing test results.

7.2.3	Mix contents of new (unpunctured) reagent packs by gently inverting pack several times	before loading on the instrument. Do not invert open (punctured) packs.

7.2.4	Use fifty-five (55) µL of sample for each determination in addition to the sample container and system dead volumes when requesting the assay (test name: TSH3). Use fifty (50) µL of sample in addition to the sample container and system dead volumes for each determination run with the special dilution feature (test name: TSH3d). Refer to the appropriate system manuals and/or Help system for the minimum sample volume required.

7.2.5	The system default unit of measure for sample results is µIU/mL.

- Quality Control

Quality control materials simulate the characteristics of patient samples and are essential for monitoring the system performance of immunochemical assays. Because samples can be processed at any time in a “random access” format rather than a “batch” format, quality control materials should be included in each 24-hour time period.  At this facility, the Biorad Immunoassay Plus Control levels 1 and 3 are used for the hTSH assay.  Follow manufacturer's instructions for reconstitution and storage.  Quality control results that do not fall within acceptable ranges may indicate invalid test results. Examine all test results generated since obtaining the last acceptable quality control test point for this analyte.

- Results

Patient test results are determined automatically by the system software using a weighted four-parameter logistic curve (4PLC) math model. The amount of analyte in the sample is determined from the measured light production by means of the stored calibration data. Patient test results can be reviewed using the appropriate screen.

Refer to the appropriate system manuals and/or Help system for complete instructions on reviewing sample results.

10.0	Expected Values

10.1	Beckman Coulter conducted a multicenter prospective study across geographically diverse locations to establish the central 97.5% reference interval in apparently healthy, euthyroid adults. Subjects with no known personal or family thyroid disease, goiter, chronic disease (including cancer, diabetes, autoimmune disease, or cardiovascular disease), acute bacterial or viral infection, or current use of prescription medication (excluding prenatal vitamins) or aspirin were enrolled in the study, following the guidance of both the National Academy of Clinical Biochemistry’s (NACB) Laboratory Medicine Practice Guideline, Laboratory Support for the Diagnosis and Monitoring of Thyroid Disease, and the Third National Health and Nutrition Examination Survey (NHANES III). The study included four populations: a general population of approximately equal numbers of males and non-pregnant females between the ages of 21-88; and pregnant females with approximately equal distribution across all three trimesters. Trimesters were defined according to American Congress of Obstetricians and Gynecologists guidelines. Approximately four hundred (400) subjects were enrolled in each population. After enrollment, subjects’ samples were screened for positive thyroid peroxidase antibody (TPOAb) and thyroglobulin antibody (TgAb) using the Beckman Coulter Access TPO Antibody and Access Thyroglobulin Antibody II assays prior to analyses. Samples with positive TPOAb or TgAb results (approximately 10%) were excluded from analysis of the TSH reference intervals.Serum samples were analyzed using multiple UniCel DxI 800 Access Immunoassay Systems with the Access TSH (3rd IS) assay, following the CLSI EP28-A3c guideline. The observed non-parametric ranges of TSH concentrations are shown below for each population tested.

| Population | Sample Size | Median (µIU/mL) | 97.5% Reference Interval (µIU/mL) |
| --- | --- | --- | --- |
| General Population (males and non-pregnant females, aged 21-88) | 367 | 1.48 | 0.45 - 5.33 |
| Pregnant Females, 1st Trimester | 318 | 1.13 | 0.05 - 3.70 |
| Pregnant Females, 2nd Trimester | 362 | 1.47 | 0.31 - 4.35 |
| Pregnant Females, 3rd Trimester | 335 | 1.61 | 0.41 - 5.18 |

10.2	For additional reference interval guidance, refer to the National Academy of Clinical Biochemistry (NACB) publication, Laboratory Support for the Diagnosis and Monitoring of Thyroid Disease, and the Third National Health and Nutrition Examination Survey (NHANES III).

10.3 	The evaluation of thyroid status should not depend on results from a single test. Complete thyroid status evaluation should include other thyroid function tests, including evaluation of thyroid autoantibodies (useful in the diagnosis of autoimmune thyroiditis), and the physician’s clinical evaluation.

10.4	At this facility, this assay TSH 3rd IS was correlated with the TSH assay (correlation coefficient = 0.9988).  The reference range was validated at 0.34-5.60 µIU/mL   established by Beckman Coulter for the TSH assay, a similar reference range used for the TSH 3rd assay on the Beckman instruments of 0.45 -5.33. The manufacturer’s suggested reference range for the new instrument is similar to the existing reference range, and was validated by running a minimum of 20 apparently normal samples. The results were within the existing limits and there was no need to change reference ranges

- Limitations

  - Samples can be accurately measured within the analytic range of the lower limit of detection and the highest calibrator value (approximately 0.01 - 50.0 µIU/mL [mIU/L]).

11.2	If a sample contains less than the lower limit of detection for the assay, report the results as less than 0.01µIU/mL.

11.3	If a sample contains more than the stated value of the highest Access TSH (3rd IS) Calibrator (S5), report the result as greater than that value (i.e., > 50.0 µIU/mL [mIU/L]). Alternatively, dilute one volume of sample with 9 volumes of Access Wash Buffer II or UniCel DxI Wash Buffer II.

11.4	Refer to the appropriate system manuals and/or Help system for instructions on entering a sample dilution in a test request. The system reports the results adjusted for the dilution.Dil-TSH (3rd IS) Special Dilution Feature for use on Access 2 and UniCel DxI systems: Samples containing hTSH concentrations greater than the concentration of the Access TSH (3rd IS) S5 calibrator can be processed using the Special Dilution Feature. When TSH3d is requested, the system automatically dilutes the sample using Wash Buffer II and reads the resulting dose off the TSH3 calibration curve.

| Test Name | Reportable Range (µIU/mL) | Sample Volume Required |
| --- | --- | --- |
| TSH3d | 37.5 - 500 | 50 µL |

11.5	For assays employing antibodies, the possibility exists for interference by heterophile antibodies in the patient sample. Patients who have been regularly exposed to animals or have received immunotherapy or diagnostic procedures utilizing immunoglobulins or immunoglobulin fragments may produce human anti-animal antibodies, e.g. HAMA, that interfere with immunoassays. Such interfering antibodies may cause erroneous results. Carefully evaluate the results of patients suspected of having these antibodies.

11.6	Other potential interferences in the patient sample could be present and may cause erroneous results in immunoassays. Some examples that have been documented in literature include rheumatoid factor, endogenous alkaline phosphatase, fibrin, and proteins capable of binding to alkaline phosphatase. Carefully evaluate results if the sample is suspected of having these types of interferences.

11.7	The Access TSH (3rd IS) results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information.

11.8	The Access TSH (3rd IS) assay does not demonstrate any “hook” effect up to 1,000 µIU/mL (mIU/L).

11.9	This assay is not validated for testing neonatal serum hTSH levels

12.0	Equivalency

A comparison of 155 samples using the Access TSH (3rd IS) assay on the UniCel DxI 800 Immunoassay System and a commercially available immunoassay kit gave the following statistical data using Passing-Bablok regression and Pearson’s correlation, following the CLSI EP9-A3 guideline:

| n | Range of Observations† (µIU/mL) | Intercept (µIU/mL) [95% CI] | Slope (95% CI) | Correlation Coefficient (r) |
| --- | --- | --- | --- | --- |
| 155 | 0.056 - 42.50 | -0.02 [-0.05 - 0.00] | 0.94 [0.92 - 0.97] | 0.98 |

†Observed concentration range of the Access TSH (3rd IS) assay.

# 13.0	Linearity

Based on CLSI EP6-A18, one high sample [approximately 50.0 µIU/mL (mIU/L)] and one low sample [< 0.01 µIU/mL (mIU/L)] were mixed to make nine sample concentrations evenly distributed across the analytical measuring range. Four replicates of the seven mixed samples, eight replicates of the low sample, and four replicates of the high sample were tested on one UniCel DxI 800 Immunoassay System. The Access TSH (3rd IS) assay demonstrated acceptable linearity throughout the analytical measuring range of 0.01 to approximately 50.0 µIU/mL (mIU/L).

14.0 	 Dilution Recovery

Four serum samples at concentrations above the Access TSH (3rd IS) S5 calibrator [approximately 50.0 µIU/mL (mIU/L)] were diluted 1/10 with Wash Buffer II. Twenty-four replicates for each sample were measured on one UniCel DxI 800 Immunoassay System, providing the following data:

| Sample | Target Concentration (µIU/mL) | Dil-TSH (3rd IS) Sample Mean Recovery (%) | Manual Dilution Sample Mean Recovery (%) |
| --- | --- | --- | --- |
| Sample 1 | 100 | 97 | 96 |
| Sample 2 | 200 | 98 | 95 |
| Sample 3 | 300 | 93 | 95 |
| Sample 4 | 400 | 93 | 93 |

15.0	Imprecision

The Access TSH (3rd IS) assay exhibits total imprecision ≤ 10% CV at concentrations > 0.02 µIU/mL (mIU/L) and total standard deviation (SD) ≤ 0.0029 µIU/mL (mIU/L) at concentrations ≤ 0.02 µIU/mL (mIU/L).

One study, using four serum-based samples on one UniCel DxI 800 Immunoassay System, generating a total of 40 assays, two replicates per assay, over 20 days with two runs per day, provided the following data, calculated based on the CLSI EP5-A3 guideline.

|  |  | Within-Run (Repeatability) |  | Between-Day |  | Between-Run |  | Total Imprecision (Within-Laboratory Precision) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample | Grand Mean (µIU/mL) (n=80) | SD (µIU/mL) | (%CV) | SD (µIU/mL) | (%CV) | SD (µIU/mL) | (%CV) | SD (µIU/mL) | (%CV) |
| Sample 1 | 0.02 | 0.0004 | 1.8 | 0.0008 | 3.7 | 0.0004 | 1.6 | 0.0010 | 4.4 |
| Sample 2 | 0.37 | 0.006 | 1.5 | 0.010 | 2.7 | 0.006 | 1.5 | 0.013 | 3.5 |
| Sample 3 | 4.71 | 0.13 | 2.7 | 0.11 | 2.4 | 0.008 | 0.2 | 0.17 | 3.6 |
| Sample 4 | 38.76 | 1.36 | 3.5 | 1.80 | 4.6 | 0.49 | 1.3 | 2.31 | 5.9 |

15.1	For details of precision studies done at this facility, refer to the validation data.

16.0	Analytical Specificity/ Interferences

16.1	Serum samples containing hTSH concentrations of approximately 0.30 µIU/mL and 5.0 µIU/mL were spiked with multiple concentrations of the substances below and run on one UniCel DxI 800 Immunoassay System. Values were calculated as described in CLSI EP7-A2.  Interference was determined by testing controls (no interfering substance added) and matched test samples (with interfering substance added). Of the compounds tested, none were found to cause significant interference (as defined by a shift in dose greater than 10%) using the highest test concentrations indicated in the table below.

| Substance | Highest Concentration Added |
| --- | --- |
| Acetaminophen | 200 µg/mL |
| Acetylsalicylic acid | 650 µg/mL |
| Bilirubin (conjugated) | 45 mg/dL |
| Bilirubin (unconjugated) | 40 mg/dL |
| Hemoglobin | 1000 mg/dL |
| Heparin (Sodium) | 3 U/mL |
| hGH (Human Growth Hormone) | 134 ng/mL |
| Human Serum Albumin (HSA) | 6000 mg/dL |
| Ibuprofen | 500 µg/mL |
| Multivitamin (Centrum Liquid) | 0.9% (v/v) |
| Triglycerides (Intra Lipid) | 3300 mg/dL |

16.2	A study was performed to evaluate the potential cross-reactivity of the assay with other substances that are similar in structure to hTSH. Serum samples containing hTSH concentrations of approximately 0.30 µIU/mL and 5.0 µIU/mL were spiked with multiple concentrations of the substances below and run on one UniCel DxI 800 Immunoassay System. Values were calculated as described in CLSI EP7-A2.

16.3	Samples containing substances at the concentrations listed below do not affect the concentration of hTSH reported.

| Substance | Highest Concentration Added (mIU/mL) | Cross-reactivity (%) |
| --- | --- | --- |
| hCG | 1,000,000 | < 0.010% |
| hFSH | 1,000 | < 0.10% |
| hLH | 3,000 | < 0.10% |

16.4	The Access TSH (3rd IS) assay is designed to have a Limit of Blank (LoB) of < 0.005 µIU/mL (mIU/L). In one study, LoB was tested using a protocol based on CLSI EP17-A2.21 A total of 360 replicates of four zero analyte samples were measured in three runs using multiple reagent pack lots and one calibrator lot on multiple UniCel DxI 800 Immunoassay Systems. This study determined the LoB for the Access TSH (3rd IS) assay to be 0.0004 µIU/mL (mIU/L), which supports the claim of < 0.005 µIU/mL (mIU/L).

16.5	The Access TSH (3rd IS) assay is designed to have a Limit of Detection (LoD) of ≤ 0.005 µIU/mL (mIU/L). In one study, LoD was tested using a protocol based on CLSI EP17-A2.21 A total of 675 replicates from five low-level samples were measured using multiple reagent pack lots and one calibrator lot in five runs on multiple UniCel DxI 800 Immunoassay Systems. This study determined the LoD for the Access TSH (3rd IS) assay to be 0.001 µIU/mL (mIU/L), which supports the claim of ≤ 0.005 µIU/mL (mIU/L).

16.6	The Access TSH (3rd IS) assay is designed to have a Limit of Quantitation (LoQ) of ≤ 0.01 µIU/mL (mIU/L) at ≤ 10% between-run CV. In one study, LoQ was tested using a protocol based on CLSI EP17-A2.21 A total of 945 replicates of seven samples were measured using multiple reagent pack lots and one calibrator lot in five runs on one UniCel DxI 800 Immunoassay System. LoQ was determined as the lowest concentration with a between-run imprecision of 10% CV. This study determined the LoQ for the Access TSH (3rd IS) assay to be 0.001 µIU/mL (mIU/L), which supports the claim of ≤ 0.01 µIU/mL (mIU/L) at ≤ 10% between-run CV.

- Backup Procedure

There are two Access systems at this facility.  The instruments serve as backup for each other. For technical assistance on the DxCi or Access, call Beckman Coulter at 800-854-3633.

- References

- 18.1	Tietz, N. W., "Specimen Collection and Processing; Sources of Biological Variation", Textbook of Clinical Chemistry, 3rd Edition, W. B. Saunders, Philadelphia, PA (2003).

- 18.2	Tietz, N. W., ed., Fundamentals of Clinical Chemistry, 5th Edition, W. B. Saunders, Philadelphia, PA (2001).

  - Beckman Coulter Access Assay Manual, 2010.

  - Beckman Coulter Access Instruction for Use Manual, 2007.

  - Beckman Coulter Technical Documents, http://www.beckmancoulter.com

Document History Page

| Change type: New, Major, Minor etc. | Changes Made to SOP - describe | Name Responsible Person/date | Medical Director Reviewed/Date | Laboratory Director/ Designee Reviewed/ Date | Date Change Implemen-ted |
| --- | --- | --- | --- | --- | --- |
| Major | New TSH (3rd IS) assay replaces hTSH assay | Thomas Christopher 09/14/17 | DR JAMES KEEFE 12/19/17 | Thomas Christopher 12/15/17 | 05/01/2018 |