Purpose:

This process describes the schedule for the calibration, cleaning, and maintenance of the Helmer Ultra CW cell washer.

| TASK | DAILY | WEEKLY | MONTHLY | ANNUALLY |
| --- | --- | --- | --- | --- |
| Inspect the tubing and drain and clear obstructions if necessary. |  |  |  |  |
| Inspect the tubing connections and secure them if necessary. |  |  |  |  |
| Clean and dry the interior after normal usage to prevent corrosion and contamination. ► For more information and instructions, see Section 6.2.2, “Cleaning the interior.” |  |  |  |  |
| Flush the system. ► For more information and instructions, see Section 6.2.3, “Flushing the system.” |  |  |  |  |
| Clean the fill ports on the rotor. ► For more information and instructions, see Section 6.2.4, “Cleaning the fill ports.” |  |  |  |  |
| Check the saline volume setting and calibrate it if necessary. Frequency varies by length of service. ► For more information and instructions, see Section 5.2, “Calibrating the saline volume.” |  |  |  |  |
| Inspect the rotor for wear, corrosion, and damage. Replace the rotor if these conditions exist, or after the rotor has been in use for 5 years. |  |  |  |  |
| Inspect the tube holders for wear and damage. Replace tube holders if they are worn or damaged, or after they have been in use for 2 years. ► For more information and instructions, see Section 6.4, “Replacing tube holders.” |  |  |  |  |
| Clean the exterior. ► For more information and instructions, see Section 6.2.1 “Cleaning the exterior.” |  |  |  |  |
| Replace the supply and drain tubing. ► For more information and instructions, see Section 3.7.1, “Installing and removing the supply and drain tubing.” |  |  |  |  |
| TASK | DAILY | WEEKLY | MONTHLY | ANNUALLY |
| Replace the pump tubing. ► For more information and instructions, see Section 6.3, “Replacing the pump tubing.” |  |  |  |  |
| Replace the tube holder inserts for 10 mm x 75 mm if applicable. |  |  |  |  |

References

Helmer Ultra CW Operator’s Manual

AABB Standards for Blood Banks and Transfusion Services, Current Edition.