| University of Washington, Harborview Medical Center 325 9th Ave. Seattle, WA, 98104 Transfusion Services Laboratory Policies and Procedures Manual | Original Effective Date: August 1, 2011 | Number: 3013-1 |
| --- | --- | --- |
|  | Revision Effective Date: | Pages: 4 |
| TITLE: Responding to Temperature Alarms in Monitored Storage Devices |  |  |

Purpose:

To provide instructions for responding to temperature alarms in blood and tissue products storage devices in order to ensure the products remain at the required temperature range.

Temperature Alarm Guidelines

| Monitored Storage Unit | Required (Acceptable) Temperature Range | Low Alarm Activation Temp | High Alarm Activation Temp |
| --- | --- | --- | --- |
| Blood refrigerators | 1° C - 6° C | 1.5° C | 5.5° C |
| Platelet incubator | 20° C - 24° C | 20.5° C | 23.5° C |
| Plasma Freezer | -18° C and lower | None | -20° C |
| Tissue freezers (including bone) | -40° C and lower | -96° C | -65° C |

| Step | Action | Related Documents |
| --- | --- | --- |
|  | NOTE: Alarms must be addressed immediately and steps taken to correct the problem. If the problem cannot be corrected within 30 minutes for blood components or 15 minutes for tissue, products must be moved to an alternate monitored storage unit. |  |
| Blood Refrigerator and Freezer Alarms |  |  |
| 1. | - For alarms that occur with an open door, close the door as soon as feasibly possible and mute the alarm. If temperature is quick to stabilize then no further action is required. |  |
| 2. Step | If the door is closed when the alarm sounds and the internal temperature is within acceptable range silence the alarm while investigating the cause: Blood storage refrigerators: - Check the battery. Replace if needed. - Verify the temperature probe is not out of the solution bottle. - Begin corrective action according to the SOP and the appropriate Operating Manuals. - Call TSL Manager or Medical Director (if time appropriate). Tissue freezers: Refer to Procedure C. Tissue Freezer Alarms, Procedure: Quality Control and Maintenance of Temperature Monitored Tissue Freezers and the Operating Manual. Action | Operators Manuals Refrigerator and Freezer QC and Maintenance Related Documents |
| 3 | If the alarm continues to sound and the reason cannot be determined, silence the alarm and contact the lead tech, manager, or appropriate service department. - Portable refrigerators: Notify Clinical Engineering to respond STAT (206-744-3496). - All other blood storage equipment: Notify Facilities Engineering to respond STAT (206-744-3191). - After hours, page the engineer on call through the hospital operator. NOTE: Engineering is expected to respond within 30 minutes of receiving a STAT call. |  |
| 4. | Complete an Equipment Out of Service Form. Leave a copy in the TSL review file holder and attach the original to the unit. (Procedure: Requesting Equipment Repair). | Removing Equipment from Service |
| 5. | Document the problem on the Refrigerator/Freezer Daily QC Log. | Refrigerator/ Freezer Daily QC Log |
| 6. | Continue to monitor the internal temperature: - As long as the internal thermometer readings are within the acceptable range: - Take their temperature every 30 minutes and record on a Manual Refrigerator/Freezer Temperature Log. Document the reason in the “Comments/Corrective Action” section. - Avoid excessive opening and closing of the door(s) to maintain internal temperature. - It is not necessary to move components or tissue at this time. - If the internal thermometer readings are not within acceptable range: - Blood storage refrigerators: - Recheck the internal temperature every 15 minutes. - If the internal temperature has not returned to the acceptable range within 30 minutes or the unit has failed completely, move the components to an alternate unit. If no alternate is available, return components to PSBC (Procedure: Packing Components for Return to PSBC). - Tissue (including bone) freezers: - Recheck the internal temperature after 15 minutes. - If the internal temperature has not returned to the acceptable range after 15 minutes or the unit has failed completely, move the tissue to an alternate freezer. If no alternate is available, immediately notify: - First call TSL Manager or Medical Director. - Risk Management (744-9574). - Infection Control (744-9560). - Neurosurgery Attending (for autologous cranial bone flaps on call through the hospital operator. - Orthopedic Attending (for all other autologous bone, if any) on call through the hospital operator. Leave completed Manual Refrigerator/Freezer Temperature Logs in the TSL review file holder. | Manual Refrigerator/ Freezer Temperature Log |
| Step | Action | Related Documents |
| 7. | Complete an on-line incident (PSN) and an internal QIM report whenever blood components or tissue (including bone) are wasted due to a malfunctioning storage unit. | Quality Improvement Monitoring Form |
| 8. | Document equipment problems and their resolution in the Communication Book and include in the shift change report. |  |
| Tissue Freezer Alarms |  |  |
| 1. | The tissue freezer alarm system is shown in the table below. When an alarm is active, the message appears in the LED message center. - Press the MUTE key to silence the audible alarm for the ringback period. - The visual alarm will continue until the freezer returns to a normal condition. - When an alarm condition occurs and then returns to normal, the freezer automatically clears the alarm condition and the message center. - When multiple alarm conditions occur, active messages are displayed in the message center one at a time, updating at 5 second intervals. Pressing Mute during multiple alarms causes all active alarms to be muted and ring back in 15 minutes. |  |
| 2. | Probe failures: - If an error is detected with the control probe (PROBE 1 FAIL), the high and low stage compressors will run continuously. As a result, the cabinet temperature will decrease until it reaches the lowest temperature that the refrigeration system can maintain. - If an error is detected with the heat exchanger probe (PROBE 2 FAIL), the freezer will cycle properly at its temperature set point using a 5 minute step start between the high and low stage compressors. - If an error is detected with the condenser probe (PROBE 3 FAIL), there is no impact on the performance of the freezer. However, the hot condenser alarm may also occur when the condenser probe fails. Call the TSL Manager or Medical Director (when time appropriate) regarding any probe failure alarms. Technical service is required. | Operators Manuals Refrigerator and Freezer QC and Maintenance |
| Platelet Incubator/Agitator Alarms |  |  |
| 1. | For alarms that occur with an open door, close the door as soon as feasibly possible and mute the alarm. If temperature is quick to stabilize then no further action is required. |  |
| 2. | - If the door is closed when the alarm sounds and the internal temperature is within acceptable range, silence the alarm and investigate possible causes. - Begin corrective action according to the SOP and the Operators Manual. - Recalibrate as necessary. | Platelet Incubator QC and Maintenance |
| 3. | If the alarm continues to sound and the reason cannot be determined within 30 minutes, and the temperature is out of acceptable range or the agitator has stopped working, then silence the alarm. - Call TSL Manager or Medical Director (if time appropriate). - Call Facilities and Engineering to respond STAT (206-744-3191). - After hours, page the engineer on call through the hospital operator. - Remove the apheresis units and place into back up unit. |  |
| 4. | Complete the Equipment Out of Service Form. Leave a copy in the TSL review file holder and attach the original to the unit. | Removing Equipment from Service |
| 5. | Document the problem on the daily QC log. | Platelet Incubator Daily Maintenance Form |
| 6. | Complete an on-line (PSN) and internal QIM incident report whenever blood components are wasted due to a malfunctioning storage unit. |  |
| 7. | Document equipment problems and their resolution in the TSL Communication Book include in the shift change report. |  |

References

AABB Standards for Blood Banks and Transfusion Services, Current Edition

Manufacturer Operator Manuals for Helmer i Series Refrigerators and Freezers

Manufacturer Operator Manual for Helmer Horizon Platelet Incubator

Manufacturer Operator Manual for VWR Tissue Freezers