UW Medicine - Pathology

6000-02-01-01

Peloris Cassette Loading Procedure

Adopted Date:4/6/2011 Review Date: Revision Date:

PURPOSE

To provide instruction on how to load cassettes onto the processing instrumentation.

PROCEDURE

Materials:

- Gloves

- Lab Coat or Vinyl apron and sleeve protectors

- Eye protection

- Gross room camera

- Cassettes in tray with lid on

- Peloris Processor

Procedure:

- Load cassettes in to trays as follows; bottom to top, left to right.

- Place cassettes on the camera table and take picture.

- Print out picture and import to PP [TS01, spec.A].

- Place the closed cassette tray (lid on) into an empty retort on selected Peloris and close.

- Hit "Protocols" selection at top of screen and then hit "Time".

[UWMC-2 hour or UWMC-8 hour].

6.   Under "Control" column, hit "Load".

7.   Screen prompt will ask you to select retort A or B (note: will indicate if a retort is being  used).  After selection "Main Screen" will reappear.

8.   Look on the far side of the Main screen of selected retort for the "Column" icon and hit "run".

9.   Enter number of cassettes and hit "okay". A prompt will appear here to load cassettes. DOUBLE CHECK here that cassette tray lid is on.

10. Hit "Done".

11. Prompt for scheduling will appear to enter end times and dates.  Enter as follows;

a) Hit "ASAP" for the 2-hour runs to start immediately

b) Hit "Edit" for delayed runs and enter the desired end time and date here.

12. DOUBLE CHECK end time and date and hit "Start".

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________