UW Medicine - Pathology

6000-02-03-01

H&E Slide Staining and Turnout Procedure

Adopted Date: 04/07/11 Review Date: Revision Date: 05/03/11

PURPOSE

Slides are stained and coverslipped, prior to delivery.  If slides are hand labeled due to printing issues, they will need to have labels printed from PowerPath.  The slides are sorted by SU number, and by priority (rush, biopsy, surgical, recut, teaching), before being delivered to the assigned attending or resident with accompanying paperwork.  Proper PPE should be worn at all times when using stainer or coverslipper.

PROCEDURE

Materials:

- Slides

- Slide racks

- Prisma Stainer

- Coverslipper

- Pallets

- PCRs

- Power Trax

- Eye protection

- Lab coat

- Gloves

1.  Load the slide staining racks with all slides facing the same direction

2.  Apply clip and hook into loading station

3.  Stainer will automatically "add basket" on monitor when front door is opened.  If not, press add " basket" and click on appropriate load station (1-3). Station is highlighted dark blue on monitor when occupied.

4.  Stainer will move slides through 21 staining steps and then into unload station (xylene) and alarm will sound.  Press "exit" on monitor.  Open front door and pull slides in rack from unload station with slides facing forward into coverslipper and remove staining clip. Monitor will read "Did you remove slides from unloading station," Press "yes" on monitor.

5.  Close door and press "start" on coverslipper.

6.  Beep alarm will sound when slides are coverslipped.  Remove slide rack, replace rack, close cover and press "verify" on coverslipper.

7.  Place slides in alpha-numeric order into a slide pallet, skipping a space between each case. If applicable, add frozen section or touch prep slides for case.  (note: print labels for FS or TP slides as needed)

8.  Pull the matching PCR paperwork for each case and place (in order of the cases) into the pallet.

9.  Open Power Trax and scan each individual slide under the "complete/ship" field.

10. Next, scan the pallet barcode for the Power Trax "pallet # (note: if more than  one pallet is used for a case, scan the top pallet and rubber band all the pallets together.

11. In "ship to" field, indicate the recipient physician or lab.  If no barcoded mailbox is available for recipient, mark "deliver directly to user" box.

12. Next, select "complete order and ship".

13.Pallets are then to be delivered to the intended recipient or mailbox.

14.For slides going to mailboxes (i.e. am turnout), upon arrival to the faculty mailbox area, log into Power Trax.

15.Select the "mailbox arrival" icon on the far right side of the screen.

16. 1st, scan the barcode on the pallet and then scan the receiving physician's box barcode.

17. Hit "confirm delivery" and continue this until all pallets have been distributed.

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________