UW Medicine - Pathology

6000-02-07-12

Corning Hot Plate/Stirrer Usage and Maintenance Procedure

Adopted Date: 04/25/2011 Review Date: Revision Date:

PROCEDURE

To ensure that the proper usage and maintenance are begin upheld for the Corning hot plate and stirrer.

PROCEDURE

Materials:

- Corning Hot plate/stirrer

- AC Power Connection

- Glassware vessel

- Chemicals/powders

- Spatula

- Stir bar

Stirring instructions:

- Fill vessel with solution to be stirred.

- Place stir bar into vessel.

- Place vessel in the center of the top surface.

- Turn Stir Control Knob until the Stirring Speed Display shows the desired speed.  The speed setting can be adjusted according to the table  below;

Operating Range (RPM):                            Adjustable Increment, (RPM):

60-100				      5

100-200				      10

200-400				      20

400-1150				      50

- When desired stirring is complete, turn the Stir Control Knob to the OFF position.  Allow stir bar to cease rotation before removing the vessel from the unit.

- Heating and stirring can be performed simultaneously with PC-420D and PC-620D models which offer both heating and stirring functions.

Note:  Please note that the viscosity of the stirred material will affect the ability of the stir bar to remain coupled to the rotating ring magnet.   Materials of high viscosity must be stirred at slower speed settings.

Heating instructions:

- Fill vessel solution to be heated.

- If using a PC-420D or PC-620D and the stirring function, place stir bar into vessel.

- Place vessel in the center of the top surface.

- Turn Heat Control Knob until the Heating Temperature Display shows the desired temperature.  The temperature setting can be adjusted in 5˚C increments.

  - Flashing display:  The number shown on the Heating Temperature Display will FLASH when actual heating temperature is not at the set temperature.

  - Constant Display: The number shown on the Heating Temperature Disply will remain constantly ON when the actual heating temperature is at the set temperature.

  - Hot Top Indicator:  The Hot Top indicator will be ON at all times when the temperature of the top surface is too hot to touch (greater than approximately 60˚C).

  - The Hot Top Indicator will FLASH when the Heat Control Knob is turned OFF, but the top surface is still too hot to touch.

  - The Hot Top Indicator will be OFF when the temperature of the top is less than approximately 60˚C.

Maintenance:

- Ceramic Top Plate:

- These products are supplied with Corning's proprietary Pyroceram top that easy to clean and highly resistant to scratches, corrosion, and chemical attack.

- The ceramic top may break during use if not properly maintained.

- Keep the ceramic top clean.  A nonabrasive cleaner may be used to clean the ceramic top.

- Inspect the ceramic top or damage during cleaning.

- Discontinue product use if the ceramic top is chipped, etched, or shows excessive scratching.  A replacement top can be ordered.

- General:

- It is important to keep this product dry and clean.

- Remove minor exterior liquid spills promptly.

- Clean exterior surfaces with a nonabrasive cleaner.

- If liquid or wet solid material gets inside the product, immediately	disconnect the power to the product and discontinue use.

REFERENCE

Corning Instruction manual

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________

Supervisor / Personnel