UW Medicine - Pathology

2000-06-01-02

Balance QC Procedure

Adopted Date: 01/29/95 Review Date: 06/05/09 Revision Date: 04/28/11

PURPOSE

To monitor quality control of the Mettler Balances in the IHC Laboratory.

PROCEDURE

Mettler Balance model # AB104-S:

Serial Number:  1121430091

UWMC Number: 17850

Instillation Date: 12/24/02

PM Number: n/a

Mettler Balance model # PM 2000 (top loading):

Serial Number:  1114243882

UWMC Number: 00960429

Instillation Date: 1/19/96

PM Number: B 2907

Preventative Maintenance Procedure:

- Daily (or before each use):

  - Level the balance.

  - Clean any spills

- Weekly:

  - Clean doors, door slides, and interior of balance.

- Annually:

  - Scientific Instruments performs preventative maintenance on the instrument every 12 months doing the following:

    - Check physical condition of power cord and plug

    - Check and record ground integrity

    - Inspect mechanical integrity, clean external controls, glass, pan, housing, etc. where applicable.

    - Clean internal chassis.

    - Inspect balance weights for proper seating, function, cleanliness.

    - Check level

    - Check damping, sensitivity, corner load, and zero range.

    - Check for accuracy of scale using calibrated weights / standards, and record.

    - Annotate sticker and records.

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________

IHC Supervisor