UW Medicine - Pathology

2000-07-01-01

Her2 Gap Analysis Procedure

Adopted Date: 06/11/07 Review Date: 06/05/09 Revision Date: 2/28/13

PURPOSE

To compare current laboratory practice in Her2 testing by IHC and FISH with guidelines proposed by ASCO/CAP.

PROCEDURE

In preparation for implementation of lab accreditation standards that require adherence to ASCO/CAP recommended testing guidelines for Her-2-Neu, the following 'gap' analysis was performed:

- The complete document of the 'ASCO/CAP Guideline Recommendations for Her2 Testing in Breast Cancer' was reviewed by the Medical Director.

- Laboratory protocols for both IHC and FISH detection of Her2 expression and gene amplification, respectively, were reviewed in the context of proposed guidelines.

- Current methodology for technical validation of both IHC and FISH based procedures were evaluated.

- Current interpretative criteria were reassessed in the context of proposed guidelines.

Results of procedural analysis:

- Currently laboratory protocols for Her2 testing by IHC are in substantial compliance with fixation (10% NBF; fixation time generally less than 48 hours), retrieval (citrate pH 6.0 buffer for 20  minutes for IHC or Epitope Retrieval buffer 1 for 20 min) and staining protocol (polyclonal anti-HER2 DAKO) requirements, with independent validation of routine deviations from standard recommended protocols.

- Current laboratory protocols for Her2 testing by FISH are in substantial compliance with both manufacturers' recommendations and with proposed guidelines.

- Because almost all FISH are run after or concurrent with Her2 testing by IHC (including referral cases - IHC in the latter cases for performed for validation purposes only and not billed to patients), it is feasible to cross-validate FISH and IHC techniques at the 95% concordance level required by guidelines and within expectations of cross-platform concordance as indicated by the literature (summarized in the guideline document).

- Effective calendar year 2007, our laboratory in all appropriate CAP surveys for proficiency testing of Her2 analysis.  These surveys can be used as an adjunct to cross-platform validation.

- Interpretative criteria for a positive result in both IHC and FISH methodologies have not yet been changed to reflect guidelines.

No other 'gaps' have been identified.

Implementation Time:

The following timeline has been established for implementation of ASCO/CAP guidelines.

March 1, 2007: review existing procedure for test validation, performance and interpretation.

July 1, 2007: implement periodic correlation of IHC and FISH results to establish corss-platform methodology validation; review relevant CAP survey results.  Develop strategies for improving concordance if 95% target is not met (either by direct comparison of 3+ IHC with gene amplification as a measure of positive assay values and correlation of 0-1+ IHC with absence of gene amplification as a measure of negative assay values, or by failure to show concordance with all 10 samples in the Her2 IHC CAP Survey sample); use external lab standard if concordance falls below 90%.

October 1, 2007: enforce guideline-based interpretative criteria; employ archived CAP Survey materials or representative 'in-house' samples to assess diagnostic / interpretative competence of faculty engaged in the evaluation of Her2 IHC or FISH samples.

November 15, 2007: target date for complete compliance in principle with ASCO/CAP guidelines.

Independent validation of IHC using "consensus controls", CAP Survey Her2 2007/2008.

REFERENCES

Wolff AC, Hammond ME, Schwartz JN, et al.: American Society of Clinical Oncology / College of American Pathologists guideline recommendations for human epidermal growth factor receptor 2 testing in breast cancer.  Arch Pathol Lab Med 2007; 131:18-43.

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		___________________________

IHC Supervisor