Purpose:

This procedure describes the daily maintenance to be performed on the TANGO instrument to assure that it is maintained properly according to the manufacturer’s recommendations.

Procedure:

| Step | Action | Related Documents |
| --- | --- | --- |
| General Statements: |  |  |
|  | - Required Materials are: - Tango Maintenance Form - Additional TANGO Reagent Form - Lint free wipes - CLR (Clinical Laboratory Reagent) Water - Disinfectant | Tango Maintenance Form Additional TANGO Reagent Form |
|  | - Use only the Daily Maintenance section of the Tango Maintenance Form. - Complete the Additional Tango Reagent Form in its entirety |  |
|  | - Record the date and initials of person performing the daily maintenance as it is completed. |  |
| Rinsing Pipettor Needles/Cleaning Washer Needles |  |  |
|  | - The TANGO can be automatically programmed to perform this function. - Select the ‘Auto Rinsing’ button to open the ‘Automatic Rinsing’ window. - Define the intervals for automatic rinsing. - For automatic rinsing, the system automatically switches over to the cleaning solution in the Cleaning Solution container. - It is recommended that the Auto-Rinsing repetition be set to 1 day at a time that is convenient to workflow. - Select the ‘Active’ button so that the check mark appears in the box. This activates the Auto Rinsing function. |  |
|  | - View Event Log in the lower left hand corner of the screen to verify auto rinsing took place. Alternately, the presence of a checkmark indicating the rinse is “Active” may be verified in the Auto Rinse dialog. |  |
|  | - Place a checkmark in the appropriate box on the Daily Maintenance section. |  |
| Checking the Suspension Cup |  |  |
| A. | - If the instrument is turned off, remove the acrylic glass cover. If the instrument is in ‘Ready’ status, follow Removing the Covers SOP to remove the cover. | - Removing the Covers SOP |
| B. | - Remove the gray plastic cover. |  |
| C. | - Without removing the cup, verify it is empty and inspect for cracks, liquid or residue. |  |
| Step | Action | Related Documents |
| Checking the Suspension Cup cont. |  |  |
| A. | - If the cup is dirty or cracked, it needs to be replaced. Follow Monthly Maintenance SOP. | - Monthly Maintenance |
| B. | - The plastic splash bowl should be dry or contain only a small amount of residual liquid. |  |
| C. | - Replace gray plastic cover. |  |
| D. | - Reinsert acrylic glass cover. |  |
| E. | - Place a checkmark in the appropriate box of the Daily Maintenance section. |  |
| Checking tubing line for bubbles |  |  |
| A. | - Observe tubing that runs from the top of the analyzer to the pipette needles. - Small bubbles are normally present in the tubing. - Large bubbles indicate leaking syringes or empty containers. |  |
| B. | - Check fluid containers to ensure there is sufficient fluid. Empty waste container as necessary. | - Managing Liquid Containers |
| C. | - Check syringes for leaking. If leaking is observed or bubbles persist, call Technical Service. |  |
| D. | - Place a checkmark in the appropriate box of the Daily Maintenance section. |  |
| Cleaning the TANGO surfaces |  |  |
| A. | - Clean up any splatters/spills immediately. |  |
| B. | - The surface of the TANGO can be disinfected with a suitable disinfectant. |  |
| C. | - After disinfection, the surfaces should be wiped again with CLR water to remove any residual disinfecting solution. |  |
| D. | - Place a checkmark in the appropriate box of the Daily Maintenance section. |  |
| Checking Pipettor needles |  |  |
| A. | - Check the pipettor needles for damage. The pipettor needles should be observed during pipetting for the following: - The needles should be straight. - The flow of water during washing should be straight down. - There should be no dripping from needles after initialization or auto-rinsing. - Liquid should flow through the needles during washing. |  |
| B. | - Gently wipe the exterior of the needles with a lint free wipe moistened with DI water. |  |
| C. | - Place a checkmark in the appropriate box of the Daily Maintenance section. |  |
| Check Strip Waste |  |  |
| A. | - If full, see Emptying Strip Waste SOP. | - Emptying Strip Waste SOP |

| Step | Action | Related Documents |
| --- | --- | --- |
| Document Reagent Lot Numbers in Use |  |  |
| A. | - The on-board reagents can be printed by selecting a CID of a specimen with current date. - In main menu select Lab Journal and enter CID of specimen in search box. - Touch specimen selected to highlight and then Export. - In the Export screen select the report for Lot Information by Date. Date should be current date. - Select Export key to get printout. |  |
| B. | - Other lot numbers in use will need to be filled out on Additional Tango Reagent Form (F-5501). | Additional TANGO Reagent Form |
| C. | Perform Reagent Dating Verification In the appropriate column: - Record open/draw date and expiration date/time from the reagent bottle or sample label. - Mark YES or NO for acceptability |  |
| D. | - Add to the form throughout the day as lot numbers are opened or patient samples are selected as controls. |  |
| E. | - File in Tango Daily QC binder after two tech ID sign and date for performance and review. - TS Lead or designee will perform a daily review. TS Manager performs a monthly review. |  |

References:

TANGO User Guide, Version 3.0.2