UW Medicine - Pathology

400-11-01-08

Array Scanning and Processing Procedure for constitutional targeted microarray analysis

Adopted Date: 10/12/09 Review Date: 07/2010 Revision Date: 03/2013

PURPOSE

To scan two-color NimbleGen arrays with Agilent Microarray Scanner and prepare 55K oligo microarray data to load onto Genoglyphix for case review

PROCEDURE

# Equipment

- Agilent 3um scanner

- Computer with software below

- Color Printer

- SciGene Ozone Filter box

- Microarray slide holder

- Microarray slide

# Software

- Agilent Scan Control

- NimbleScan

- DNA Analytics

- NGpackager

- Excel

- www.genoglyphix.com

  - Procedure

Part I: Loading and Scanning Slides

Note: Keep arrays in a dark desiccator until you are ready to scan them. When handling slides, wear powder-free gloves and use care to touch only the slide’s edges

- Loading Slides

- Turn on the Agilent Scanner and the SciGene Ozone Filter for about 10 minutes before start scanning

- Load slide into slide holder with proper orientation (barcode should be on the back side.  DO NOT touch the surface of the slide)

- Load slide holder into the slot of the carousal.  (use gloves when open and close scanner to minimize fingerprint and dusts)

- Record slide’s position

- Close the cover and close door of the Ozone box

- Scanning Slides

- Open Agilent Scan Control (It takes about 5 minutes for the scanner to be ready)

- Select the starting and ending slot to scan

- Profile should be set as <Default>

  - Slide ID:  <Auto detect>, Channels: R+G, Scan Region: Full Slide (71 x 21.6 mm), Resolution: 3um, TIFF: 16 bit, R PMT: 100%, G PMT: 100%, XDR: <NO XDR>, Output Path: D:\ScanData

- Click on “Scan Slot 1-xx” (lower right corner) to start scanning

- After the scanning is completed, a summary of the scanning process will pop up.  Make sure there is no critical error (red text) before move on to next step.

- Close Agilent Scan Control

- Sources of Error for Part I: Loading and scanning slides

- Not using proper handing technique when handing slides

- Loading slide in the wrong orientation

- Carousel is not properly in place

- Scanner has not warm up

- Scanning at the incorrect resolution

# Part 2:  NimbleScan to Burst, Align, and Extract Features

Note:  NimbleScan is picky.  Make sure to close the software and open it again after each step.

- Rename images

- Open “D:\Scan Data” folder and search for scanned tiff files

- Original file name should be: slideID_date_S01_Green/Red.tif

- Rename it as: slideID_532.tif for Green Signal and slideID_635.tif for Red

  - If scanner did not pick up slideID (barcode), enter it manually

- Burst Multiplex Images…

- Open NimbleScan software

- Click on file -> Burst Multiplex Images…

- Click on “add images” from pop up screen and select all images to be burst and Add to batch

- Layout file = “SignatureSelect2p03plex.ncd” in the same folder as “.ndf”

- Select the output folder and click on Burst to start

- Close NimbleScan

- Crop image

- Open NimbleScan software

- Click on file -> open

- Open the folder where the tiff images are stored

- Select one color channel at a time

- Select the appropriate design file (.ndf) (design file is located on the desktop inside the SignatureSelect2p03plex folder.  Or can be downloaded on www.genoglyphix.com)

- Click open

- Crop the image (select the region of interest, then click on edit-> crop)

- Save image

- Repeat 1-8 for the rest of the image files

- Close NimbleScan

- Align images

- Open NimbleScan software

- Click on file->open

- Choose bursted image (one channel at a time).  (note: bursted image end with A01, 2, 3)

- Choose the design file (.ndf) (091104_HG18_CGX3_55kv1_HX3.ndf).

- Click OK, then Click OPEN

- Click on View -> auto contrast/brightness to enhance image quality

- Check and record the image quality for physical abrasions, integrity of the array, and global intensity of the array features.

- Click on Analysis -> auto align to align the grid to the image

- Check and record alignment values.  (note: lower left corner, this value should be <0.2.)

- Check the green dots using navigation panel on the top left corner to check for correct alignment.

- Save aligned image.  (aligned image will have an “Aligned” description in the open image pop-up)

- Align another image by clicking on File->New Images (instead of File->open)

- Repeats step 3 to 11 for the rest of the unaligned images

- Close NimbleScan

- Feature Extraction

- Open NimbleScan

- Go to Analysis -> CGH -> segMNT…

- Add all the aligned files from “part D”

- NDF design file should be <CGH_55Kv1_Hx3_091104_HG18\091104_HG18_CGX3_55kv1_HX3.ndf> and position file should be  <091104_HG18\091104_HG18_CGX3_55kv1_HX3.pos>

- Make sure the data channels ratios is the correct one (usually 532/635)

- Deselect the following from “Choose the desired algorithm settings…”

  - Averaging windows: 1x & 10x

  - Generate GFF

  - PDF Formats: Single panel, rainbow & Multi-panel

- Choose the destination folder… (note: if the folder is changed, the software will not let you click Run.  Repeat step 2-5 and the Run option should be active.)

- Click on Run to generate…segMNT.txt

- Close NimbleScan

- Sample Tracking

- Open NimbleScan

- Go to Analysis -> Sample Tracking…

- Add files -  Add only the Green channel (532 file)

- Check the destination file

- Run.  This will generate a stc.xls file.  (double check to make sure everything matches up)

- Close NimbleScan

- Sources of Error for Part 2:  NimbleScan to Burst, Align, and Extract Features

  - Make sure to close and reopen NimbleScan after each process

  - Bursted file is not aligned when running feature extraction

  - Data channel might have switched

  - Use the incorrect *.ndf and *.ncd file

  - Cropping off part of the image

  - Not saving image after each step

  - Selection of the appropriate algorithm settings

# Part 3:  GGX file converter using NGpackager

Note: Software will overwrite existing file without any warning.

- NGpackager  – converts *.txt file to *.ggx file

- Open “ngpackage.exe” file from the NGpackager_ folder on the desktop

- Click Choose Probe Data and select an output file (segMNT.txt file) created from Feature Extraction using NimbleGen Software

- Click Great GGX File.

- When it show Done, Repeat step 2-3 for the others files

- Click Quit to exit.

- Sources of Error for Part 3:  GGX file converter

- Not selecting the right file format for the right program

- Will overwrite existing file without any warning

  - Output format: Complete Genome

  - Select file location: example of file name: “SlideID_Interval Report.xls”

  - Click on save

- Repeat Part A  for the other *.ggx files

- Close without saving

# Part 4:  Loading files into Genoglyphix

- Log into www.genoglyphix.com

- Click on Add Subject and enter subject information

- Click on Add New Dataset GGX file and select *.ggx file from Part 3A, Enter the unique slideID in the slideID box.

- Click on submit data (this will upload the 2 files into the genoglyphix server)

- Click on Process GGX files to create dataset once the upload is complete

  - Repeat step 4-6 to add new dataset

- Check and record the standard Deviation.  (Cut off value should be <0.2)

- Mark and initial Data reviewed on Genoglyphix.

- Inform technologists that Data is ready for case review

REFERENCES

- NimbleGen Arrays User’s Guild: CGH Analysis v5.1

- Agilent G2565CA Microarray Scanner System User Guide

- Agilent DNA Analytics 4.0 CGH Module User Guide

- www.genoglyphix.com

Written By:					Director Approval:

(Signature and Date)				(Signature and Date)

_______________________		________________________

UW Medicine - Pathology

Cytogenetics - UWMC

SIGNATURE PAGE FOR POLICIES AND PROCEDURES

Procedure / Policy Title: 	Array Scanning and Processing Procedure

Procedure / Policy Number: 	400-11-01-03

I, ___________________ approve __________________ _to sign as Designee,on

___________, 20____.

__________________   	______________________		_______________

Medical Director 	  	Medical Director Signature or		Date Reviewed

Print or Designee		Designee

| STAFF NAME: (printed) | STAFF SIGNATURE | DATE REVIEWED |
| --- | --- | --- |