Purpose

To describe the steps for monitoring and maintaining storage units when the TempTrak is not transmitting the temperature data to the server for the department to monitor on a continuous basis.

Policy

- All HMC Transfusion Service (TSL) storage units are monitored 24/7 by staff.

- Temperature readings must agree within + 2°C.

- HMC TSL maintains an audible alarm system as a back up to the Temp Trak electronic alarm system.

| Step | Action | Related Documents |
| --- | --- | --- |
| Out of Service – Temp Trak system is not available |  |  |
| Start Chart Recorder | - Place a new chart recorder on all affected equipment. - Press and hold the Chart Change Switch (#3) for 1 second until pen begins to move to the left of the chart. - Unscrew the knob at the center - Choose the correct chart paper - Position the new chart so the correct time line coincides with the time line groove on the chart plate. - Replace the knob and hold paper while securing. - Push the chart change switch (#3) until pen begins to move back onto the chart. - Check that the pen is marking. - Document change on the unit’s temperature log sheet. - Portable 3 chart recorder: - Press the pen home button: the pen begins to move to the right of the chart - To replace the knob, push the pen home button: the pen begins to move back onto the chart - Once TempTrak resumes, remove chart and adhere to corresponding Downtime Daily QC Log. - Submit completed log sheets to be reviewed by Quality Coordinator. | Downtime Daily QC Logs |
| Record Temps | - Fill out daily temperature log sheet for each corresponding piece of equipment. - Record temperatures on log sheets of each piece of equipment. Temperatures should include digital temperature, chart recorder reading, and any internal thermometer temperatures - All temperatures must agree within 2°C. - Ensure chart recorder is working properly and the LED green light is steady. - Change the chart recorder battery, as needed. - Test the alarms and record on temperature log sheets in the alarm test column. | Downtime Daily QC Logs |

| Exceptions | - Humidity and Room Temperature are monitored by Temp Trak. - It is not necessary to monitor either parameter when Temp Trak is not available. |  |
| --- | --- | --- |
| Chart Recorders: Temp Trak is not available |  |  |
| Change Chart Recorder Graph Paper | Weekly: - Label new paper prior to replacing with HMC TS, start date, stop date, refrigerator or freezer number, and Tech ID. - Remove and replace chart recorder graph paper (see Step 1) - Document change on temperature log sheet. - Submit for review the previous week’s chart recorder data. Any deviations should be noted on graph paper. |  |
| Adjust Chart Recorder | - Wrong day/time recording? Reposition the paper to correct day/time. - Pen not tracing well? Ensure that the pen is in contact with the paper. If out of ink, change the chart recorder pen. - Snap the plastic stylus around the outer edge of the metal pen arm. - Remove pen cap by gently lifting the pen arm away from the chart paper then take cap off to expose the fiber tip. Carefully place pen on the paper. - Adjustments can be made by slightly bending the metal pen arm in the center toward the paper. - If chart recorder still isn’t working properly, but all other temperatures are within acceptable limits, the equipment must be monitored by documenting the temperatures every 4 hours with the Manual Refrigerator and Freezer Temperature Log. |  |
| Chart Recorder needs calibration | - Place a certified thermometer in the solution bottle with the recorder probe. - Allow to stabilize and then compare thermometer to pen location on chart. - If an adjustment is required, press and hold for 5 seconds either the left top button or right top button to move the pen the corresponding direction to match the thermometer temperature. |  |
| Chart Recorder Battery | - Replace chart recorder 9-volt battery as necessary. Document in the comments section of the temperature log for that equipment. |  |

References:

Standards for Blood Banks and Transfusion Services, Current Edition. American Association of Blood Banks. AABB Press, Bethesda, MD

Helmer i Series Refrigerator Service Manual

Helmer i Serier Freezer Service Manual

Jewett Refrigerator Service Manual

Helmer Platelet Incubator Operation Manual

Helmer Flatbed Agitator Operation Manual

Helmer Chart Recorder Operation Manual