Platelet Incubator Maintenance Form

Incubator No._________________________________     Year: ________________________________

Programmed Alarm Parameters

| Storage Equipment | Low Alarm Setting | High Alarm Setting | Delay in Agitation |
| --- | --- | --- | --- |
| Platelet Incubator | 20.5°C | 23.5°C | 1 Minute |

Monthly Audible and Temp Trak Alarm testing is recorded on the Monthly Storage Unit  QC  form.

| Date | Reason | Auto Alarm Test |  | Agitator Motion Alarm Delay |  |  | Condensor Cleaning | OK? Y/N | Tech | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | High Alarm | Low Alarm | Stop Time | Time in alarm | Actual Delay |  |  |  |  |

Additional Testing

| Date | Reason | Open Door Alarm Test |  | Power Failure Alarm | OK? Y/N | Tech | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Alarm sounds? Y/N | Time Delay | Alarm sounds? Y/N |  |  |  |

Calibration of iSeries Temperature Monitor and Temperature Controller

| Date | iCenter temp | Nist Temp | Calibration required? Y/N | Temp Controller | Nist Temp | Calibration required? Y/N | OK?Y/N | Tech | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

Annual Maintenance – performed by HMC Engineering

| Inspect motor bearings Y/N | Inspect wheels Y/N | Lubricate Y/N | Tech | Date |
| --- | --- | --- | --- | --- |

Reason Codes:

Q = Quarterly, PC = Program Change, RTS = Return to Service

Comment Code Key:

RPR=Repeat	LE=Late Entry	CE=Corrected Entry		CAL=Recalibrated		RS=Referred to Supervisor

NC= Not corrected			OTR=Explain		EB=Explain on back of page

Reviewed by:____________________________________Date:____________________________________