| University of Washington, Harborview Medical Center 325 9th Ave. Seattle, WA, 98104 Transfusion Services Laboratory Policies and Procedures Manual | Original Effective Date: August 1, 2011 | Number: 3013-2 |
| --- | --- | --- |
|  | Revision Effective Date: March 17, 2014 | Pages: 4 |
| TITLE: Responding to Temperature Alarms |  |  |

Purpose:

To provide instruction in response to temperature alarms, for the storage units of blood and tissue components, to ensure the products remain within the required temperature range.

Temperature Alarm Guidelines

| Monitored Storage Unit | Required (Acceptable) Temperature Range | Low Alarm Activation Temp | High |
| --- | --- | --- | --- |
| Blood Refrigerators | 1° C - 6° C | 1.5° C | 5.5° C |
| Reagent Refrigerator | 2° C - 8° C | 2.5° C | 7.5° C |
| Platelet Incubators | 20° C - 24° C | 20.5° C | 23.5° C |
| Plasma Freezers | -18° C and lower | None | -20° C |
| Ultralow Freezers | -40° C and lower | -96° C | -65° C |
| Room Temperature | 19.4 – 23.9oC | 19.4 oC | 23.9 oC |
| Humidity | 15 – 55% | 15% | 55% |

Alarms

Transfusion Services Laboratory has a redundant alarm system for all blood, bone, and reagent storage devices. Room temperature and humidity are also monitored by the Temp Trak System. - Temp Trak is a computer monitoring system with pager notification. - Audible alarms are in place for all units. - Alarms must be addressed immediately and steps taken to correct the problem. - Products must be moved to an alternate monitored storage unit if the problem cannot be corrected within: - 30 minutes for refrigerated components - 15 minutes for frozen components and room temperature components - Bone freezer contents must be moved if temperature reaches -40oC - Reagent and patient sample storage is less critical. - Room Temperature and Humidity are reported to Engineering for resolution. - Alarms are documented in Temp Trak and/or equipment QC logs. - All measures taken are recorded. - Relocation of product is documented. - Operator Manuals and Maintenance and Troubleshooting procedures are followed: - Maintenance and Troubleshooting of Platelet Incubators - Maintenance and Troubleshooting of Refrigerators and Freezers - Maintenance and Troubleshooting of Ultralow Freezers

| Step | Action | Related Documents |
| --- | --- | --- |
| - General Response |  |  |
| Open Door | - Close the door as soon as feasibly possible - Mute the audible alarm. - If Temp Trak pager sounds, respond to Temp Trak alarm event. - If temperature is quick to stabilize then no further action is required. | - Using the Temp Trak System to Monitor Blood Storage Devices |
| Closed Door | If the door is closed when the alarm sounds and the internal temperature is within acceptable range silence the alarm while investigating the cause: Blood storage refrigerators/freezers: - Check the alarm and/or system battery. Replace if needed. - Verify the temperature probe is not out of the solution bottle or has been damaged. - Begin corrective action according to the SOP and the appropriate Operating Manuals. - Call TSL Manager or Medical Director (if time appropriate). | Operator Manuals |
| 2nd Alarm | If the alarm continues to sound and the reason cannot be determined, silence the alarm and contact the Lead Tech, Manager, and/or appropriate service department. - Portable refrigerators: Notify Clinical Engineering to respond STAT (206-744-3496). - All other storage equipment: Notify Facilities Engineering to respond STAT (206-744-3191). - After hours, page the Nursing Supervisor through the hospital operator. They will contact the engineer on call to respond to TSL. NOTE: Engineering is expected to respond within 30 minutes of receiving a STAT call. |  |
| Remove from Service | Remove the equipment from service. | Removing Equipment from Service |
| Temp Trak | - Document the problem in TempTrak system. - Include temperature(s) of the digital and thermometers in the storage unit. - Clear the Alert in order to activate the alarm pager | TempTrak SOP |
| Acceptable Internal Temperatures | Continue to monitor the internal temperature: As long as the internal thermometer readings are within the acceptable range: - Record the temperature every 30 minutes. Document the reason in the “Comments/Corrective Action” section. - Avoid excessive opening and closing of the door(s) to maintain internal temperature. - It is not necessary to move components or bone at this time. | Manual Storage Unit Temperature Log |

| Step | Action | Related Documents |
| --- | --- | --- |
| General Response |  |  |
| Unacceptable Internal Temperatures | If the internal thermometer readings are not within acceptable range: Record the internal temperature every 15 minutes. Blood storage Refrigerators: - If the internal temperature has not returned to the acceptable range within 30 minutes or the unit has failed completely, - Move the contents to an alternate unit. Plasma or Platelet Incubators: - If the internal temperature has not returned to the acceptable range after 15 minutes or the unit has failed completely, - Move the contents to an alternate unit. If NO alternate is available: - Contact the TS Manager and PSBC Inventory Management - Make arrangements to return components to PSBC Ultralow Freezers (Bone and Tissue Storage Units) - If no alternate is available, immediately notify: - First call TSL Manager or Medical Director. - Risk Management (744-9574). - Infection Control (744-9560). - Neurosurgery Attending (for autologous cranial bone flaps on call through the hospital operator. - Orthopedic Attending (for all other autologous bone, if any) on call through the hospital operator. - Leave completed Manual Refrigerator/Freezer Temperature Logs in the TSL review file holder. | Packing Components for Return to PSBC Manual Storage Unit Temperature Log |
| Wastage Reporting | Complete an on-line incident (PSN) and an internal QIM report whenever blood components or bone are wasted due to a malfunctioning storage unit. | Quality Improvement Monitoring Form |
| Room Temp and Humidity | Report directly to Facilities Engineering. | Removing Equipment from Service |
| Notifications | Document equipment problems and their resolution in the Communication Book and include in the shift change report. |  |

| Bone Freezer Alarms |  |  |
| --- | --- | --- |
| LED Message Center | The tissue freezer alarm system is shown in Operators Manual. When an alarm is active a message appears in the LED message center. - Press the MUTE key to silence the audible alarm for the ring back period. - The visual alarm will continue until the freezer returns to a normal condition. - When an alarm condition occurs and then returns to normal, the freezer automatically clears the alarm condition and the message center. - When multiple alarm conditions occur, active messages are displayed in the message center one at a time, updating at 5 second intervals. Pressing Mute during multiple alarms causes all active alarms to be muted and ring back in 15 minutes. | Operator Manuals |
| Probe Failures | Probe failures: - If an error is detected with the control probe (PROBE 1 FAIL), the high and low stage compressors will run continuously. As a result, the cabinet temperature will decrease until it reaches the lowest temperature that the refrigeration system can maintain. - If an error is detected with the heat exchanger probe (PROBE 2 FAIL), the freezer will cycle properly at its temperature set point using a 5 minute step start between the high and low stage compressors. - If an error is detected with the condenser probe (PROBE 3 FAIL), there is no impact on the performance of the freezer. However, the hot condenser alarm may also occur when the condenser probe fails. Call the TSL Manager or Medical Director (when time appropriate) regarding any probe failure alarms. Technical service is required. | Operator Manuals |

References

AABB Standards for Blood Banks and Transfusion Services, Current Edition

Manufacturer Operator Manuals for Helmer i Series Refrigerators and Freezers

Manufacturer Operator Manual for Helmer Horizon Platelet Incubator

Manufacturer Operator Manual for VWR Tissue Freezers