Purpose:

This process provides instructions for how to QC and maintain manual testing stations in TSL.

Process:

| #### Required Supplies |  |  |  |
| --- | --- | --- | --- |
| Cleaning/Disinfection | Labelling | Handling | Routine Reagents |
| - Disinfectant - Lint-free cloths - Tissues - Paper towels - Gauze squares | - Indelible black markers - Indelible ink pens | - Parafilm - 12 × 75 glass tubes - Plastic lids to fit 12 × 75 tubes - Blood bank transfer pipettes - Segment opening devices - Sample centrifuge with balance tubes - Calibrated serologic centrifuge - Calibrated cell washer - Squeeze bottle of 0.9% saline | Routine reagent rack with - anti-A - anti-B - A1 cells - B cells - anti-D - screening cells 1, 2, 3 - LISS - anti-IgG - antihuman globulin (AHG) control cells - ABO + Rh Control |
| Safety | General |  |  |
| - Eye splash protection - Scissors - Gloves - Biohazard containers - Lab coat | - 37 C dry bath - Computer terminal(s) - Laser printer - Label printer - Agglutination viewer and extra bulbs - Testing rack |  |  |

Procedure A:  Start of Shift

|  |  | Action |  |  | Related Documents |
| --- | --- | --- | --- | --- | --- |
|  | - Follow these instructions at the start of every shift for all benches. |  |  |  |  |
|  | - Prepare the working disinfectant solution or obtain disinfectant wipes. |  |  |  |  |
|  | Action |  |  | Related Documents |  |
|  | - Dampen all non-metal surfaces with disinfectant or wipes. - Allow surfaces to air dry. |  |  |  |  |
|  | - Restock bench with essential supplies. |  |  |  |  |
| 5 | At the start of EACH shift for all benches: |  | Then |  |  |
|  | Verify the functioning and calibration date of all equipment |  | Heat block: - Read and record temperature. Confirm acceptable fluid level in test tube. Adjust as per procedure, if found out of acceptable limits. Cell Washer: - Daily checks - Weekly checks - Monthly checks - Annual checks - Confirm schedule and perform as needed. Serofuge and Specimen Centrifuge: - Clean regularly as needed. | - Thermo Scientific Heat Blocks Operation & Maintenance Procedure - Heat block Daily QC Form - Helmer Ultra CW Operation, Cleaning & Maintenance Procedure - Cell Washer Daily QC Form - Helmer Ultra CW Maintenance Schedule - Helmer Serofuge Operation & Maintenance Procedure - Manual Reagent QC Procedure - Daily Manual Testing Reagent QC Record - Bench Daily Reagent & Equipment QC Form - Removing Equipment from Service |  |
|  | Verify QC and acceptability of reagents |  | - Confirm reagent QC has been performed by comparing lot numbers to lots listed on the Daily Reagent QC form - Perform QC if indicated - Confirm saline squeeze bottle is in date. - Discard any outdated reagents. - Replace any empty reagents. - Follow up on low reagent supplies |  |  |
|  | Document completion of reagent QC for rack assigned to bench: - Bench designation/tech ID - Y or N for equipment in-date and reagent QC - Remove from service any equipment found to be out of validation or performing incorrectly. |  |  |  |  |
| 6 | Perform additional maintenance duties as specified on the CT Shift Responsibilities Checklist. - Duties are rotated monthly by shift - Tech is responsible for duties if not performed by assigned shift. |  |  | CT Shift Responsibilities Checklist |  |

Procedure B: End of Shift

|  |  | Action |  | Related Documents Title |
| --- | --- | --- | --- | --- |
| 1 | Follow these instructions at the end of every shift. |  |  |  |
| 2 | - Dampen all non-metal surfaces with disinfectant solution or wipes. - Allow to air dry. |  |  |  |
| 3 | Ensure there are sufficient amounts of reagents in the rack for the next user. |  |  |  |
| 4 | Refill saline squeeze bottle, if needed. |  |  |  |
| 5 | Fill any depleted cleaning, labeling, or handling supplies (from the Required Supplies list). |  |  |  |
| 6 | Store patient samples and paperwork. |  |  |  |
| 7 | Empty and replace full biohazard containers. |  |  |  |
| 8 | Leave the workstation in a clean and usable condition. |  |  |  |