Refrigerator and Freezer Maintenance Form

Year: ________________

Equipment No._________________________________ Type of Equipment: _______________________

Programmed Alarm Parameters

| Storage Equipment | Low Alarm Setting | High Alarm Setting |
| --- | --- | --- |
| Blood Storage Refrigerator | 1.5oC | 5.5°C |
| Reagent Storage Refrigerator | 2.5oC | 7.5°C |
| Plasma Freezer | N/A | -20°C |
| Ultra Low Freezer | -90°C | -65°C |

| Date | Reason | High Alarm Test |  |  |  | Low Alarm Test |  |  |  | OK? Y/N | Tech ID | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Digital | Chart | NIST/ Fluke | Temp Trak | Digital | Chart | NIST/ Fluke | Temp Trak |  |  |  |

All recordings must agree within 1.0oC.

Date:___________________________Tech:__________________________________

Reason Codes:

Q = Quarterly, PC = Program Change, RTS = Return to Service

Comment Code Key:

RPR=Repeat	LE=Late Entry	CE=Corrected Entry		CAL=Recalibrated		RS=Referred to Supervisor

NC= Not corrected			OTR=Explain		EB=Explain on back of page

Reviewed by:____________________________________Date:____________________________________