## Slide 1: Serofuge and Cell Washer Operation and Maintenance

## Slide 2: Basics
Centrifuges must be balanced before operation. NOTE: Balance tubes should be used for a limited time due to stress breakage.
Tach and Timer QC is performed by Scientific Instruments
Programmed settings must be checked prior to each operation cycle
Blood spills must be cleaned immediately

## Slide 3: For All TSL Staff Let’s Get Started
EBA 20 Sample Centrifuge

## Slide 4: EBA 20 Sample Centrifuge
Power Switch is on the back
START
Time
Speed
Red Stop Key
Open Lid button

## Slide 5: Programming
Up/Down buttons allow changes in speed and time
TANGO samples
  SOFT Spin to standardize robotic sampling
  Setting is 35 RPM for 5 minutes
Manual Bench Testing
  Any time and speed will do

## Slide 6: EBA-20 Cleaning and Maintenance
Daily
Clean housing and surface
Mild (pH 5-8) soap and water
  Green Top Sani-Wipes OK
Clean obvious blood spots immediately
Do Not use Purple Top Sani-Cloth wipes
  Destroy plastics
  Harm porous surfaces

## Slide 7: EBA 20 Cleaning and Maintenance
As Needed
Remove rotor head
  Wrench is stored in Fellowes black tool kit under Helmer Thawers
Clean with mild detergent.
Rinse with water
Check parts for wear and corrosion.

## Slide 8: For MLS Staff Testing Bench Equipment
EBA 21 Sample Centrifuge
Ultra CW Cell Washers
Stat Spins

## Slide 9: EBA 21 Testing Serofuge
Clean daily with mild soap and water or Green Top Sani-Wipes

## Slide 10: Programmable
Serofuge Programs: Press UP/Down buttons
Program 1 Routine tube testing 2900 rpm / 20 seconds
Program 2 Buffered eluate 3400 rpm / 1 minutes
Program 3 anti-Jka (BioRad) recommends 1 min spin
2900 rpm / 1 minute

## Slide 11: EBA 21 Serofuge
Head lifts out easily for rapid cleaning of spills

## Slide 12: Emergency Lid Release – EBA 21
Turn off Power
Lid May Fail to Open following a Power Failure:
Turn off Power.
Make sure rotor has stopped.

## Slide 13: EBA 21 Parts and Manual
Locate the pin driver in the equipment drawer under the SOP binders
  Bright Orange Handle

## Slide 14: Emergency Lid Release – EBA 21
Move the serofuge to the front of the counter and locate the emergency release hole.
(Not wise to tip the serofuge back this far except for illustration!)
Insert the pin driver, pushing upwards and at the same time rotate the lid-lock handle.
Open the lid.

## Slide 15: Helmer Ultra CW Cell Washer
Programmable Serofuge and Cell Washer

## Slide 16: Programmable
Serofuge Programs
SPIN 2900 20 seconds
Cell Washer Programs
Program 1 4 cycle wash
CLEAN program Weekly Decontamination
Other Programs
Program 2 - 5 Used for calibrations and for 10mm tubes

## Slide 17: Daily QC
Inspect tubing and drain. Clear obstructions.
  At Bench C and D, empty waste daily
Inspect tubing connections and secure
Clean and Dry interior after normal usage
Check Saline Volume Delivery

## Slide 18: Saline Volume Delivery
Place volumetric cylinder under the saline spout.
Hold Saline button until CALIBRATE appears
Press CHECK until pump starts and saline is entering the cylinder

## Slide 19: Weekly QC
Flush the system with 10% bleach
Clean the fill ports on the rotor
Why does the procedure call for
Tubes in every other tube holder during
CLEAN program?

## Slide 20: Why are tubes required during Clean?
Tube contents are flunge out and up, cleaning the upper part of the bowl and lid with bleach solution.

## Slide 21: Monthly QC
Inspect rotor for wear, corrosion and damage
Replace if indicated or at 5 years
Inspect tube holders for wear and damage
Replace if indicated or at 2 years
Clean the exterior

## Slide 22: Annual QC
Replace supply and drain tubing.
Replace pump tubing

## Slide 23: STATSpin Express 3
Speed is NOT adjustable
Used primarily for STAT Manual Testing orders
Not used for Eluates
Not Programmable
Buttons for 2, 3 or 5 minutes spin

## Slide 24: Removing the Bowl
Large Hex required

## Slide 25: Remove all Tube Holders 1st
Center ring screws off
After removal, head will lift out.