**Purpose:**

To ensure that the cover is opened and closed properly when necessary to gain access to internal systems.

**Procedure:**

|  |  |  |
| --- | --- | --- |
| **Step** | **Action** | **Related Documents**  |
| 1 | Ensure system is in ‘ready’ status.Note: Opening the cover during ‘active’ status will stop all mechanical movements and cause a fatal system error |  |
| 2 | Touch the ‘Maintenance’ button in the Main Menu to open the ‘Maintenance’ menu |  |
| 3 | Touch the ‘Open Front Cover’ button; the ‘Lift Front Cover’ dialog will open. |  |
| 4 | Open the front cover by lifting it upwards. The ‘OK’ button in the ‘Lift Front Cover’ dialog will be deactivated when the Front Cover is open. |  |
| 5 | Close the front cover by gently moving it downwards. |  |
| 6 | Touch ‘OK’ in the ‘Lift Front Cover’ dialog to confirm that the front cover has been closed properly. |  |

**References**

TANGO Infinity User Manual, Version 1.2.1