Purpose:

This document describes the process on how to report and track the completion of work requests for  equipment, critical in daily usage, and in order to assure appropriate revalidation before placing back into service.

Procedure:

| Step | Action | Related Documents |
| --- | --- | --- |
| 1 | Staff should troubleshoot before placing a service call | QP: Equipment Management Maintanence and Troubleshooting of Storage Units Use and Maintenance of the Tempcheck Device |
| 2 | - Portable Refrigerators: Notify Clinical Engineering to respond STAT (206-744-3496) - All other blood storage equipment: Refrigerators and Freezers. Notify Facilities Engineering to respond STAT (206-744-3191) - Centrifuges, cell washers, plasma thawers, platelet rotators, and small equipment are repaired by Scientific Instruments. - Main phone line 206-543-5580 - Emergency pager 206-991-4030 - After hours, page the Nursing Supervisor through the hospital operator. The Nursing Supervisor will contact the Engineer on call. - NOTE: Engineering is expected to respond within 30 minutes of receiving a STAT call. |  |
| 3 | Prior to the completion of an Equipment Out of Service Form, all equipment must be thoroughly cleaned with a disinfectant cleaner, before repair is attempted by a technician. - Complete Notice of Laboratory Equipment Decontamination form, which is found at www.ehs.wa.edu/forms. - Complete an Equipment Out of Service Form. - Temp Trak: Document incident in Temp Trak for all monitored equipment (Add Note). Do NOT clear alarm. - Submit CSS request electronically | Equipment out of Service Form Notice of Lab Equipment Decontamination Form Table A: Completing Clinical Support Services (CSS) Centralized Work Request |
| Step | Action | Related Documents |
| 4 | Record the call date/time on the Equipment Out of Service Form and make a photocopy of the completed form. - The original gets taped to the equipment so it is visible to staff. - The copy goes to the Lead Technologist/TSL Manager. |  |
| 5 | When an Engineer or Technician arrives, record the response date/time on the Equipment Out of Service Form. |  |
| 6 | The Equipment Out of Service Form accompanies any equipment removed from the lab for repair. - Maintain a copy in TSL. - Instruct the Engineer/Technician to complete the Clinical Engineering/Facilities/Scientific Instruments Sign-Off box, and result the resolution field once the equipment is returned to service. |  |
| 7 | Upon completion of the repair: - Clinical Engineering OR Scientific Instruments will return equipment with a completed Equipment Out of Service Form or inform the lab when the equipment can be repaired and returned to service. - Lead Technologist or designee will review completed Equipment Out of Service Forms for completeness of documentation including action taken. - Lead Technologist, Compliance Analyst, and/or Manager will evaluate whether revalidation is required. |  |
| 8 | A signature by lab staff is required to authorize “Return to Service” to approve the work order is completed, and equipment is returned. |  |
| 9 | If no revalidation is required, return equipment to service as follows: - For minor repairs of small equipment, do daily QC tasks and monitor under normal working situations for expected results with documentation. - For repairs of temperature monitored equipment, perform the quarterly QC task as stated in the corresponding SOP. Document and adhere to corresponding daily logs. - Allow tissue storage freezers come to temperature (-80˚C) and perform the quarterly QC. Document and adhere to corresponding daily logs. | Maintenance and Troubleshooting of Blood Storage Devices |
| 10 | Store Equipment Out of Service forms and any additional documentation in the appropriate equipment folder in TSL. |  |

Table A: Completing Clinical Support Services (CSS) Centralized Work Request

| Step | Action |
| --- | --- |
| 1 | Log into HMC Intranet |
| 2 | Click on Workplace Services tab |
| 3 | Click on Engineering Work Request (CSS Service Requestor) |
| 4 | Click on Submit Work Request |
| 5 | Complete the fill in information: - Hospital - Harborview Medical Center - Building - Center Tower - Floor - Basement - Room # - BCT 67 or BCT 38 - Requestor’s Name - your first and last name - Contact Phone - use TSL main line - Contact email – use your email - Department Name - TSL - Department Mailbox - 359743 - Account or Budget # - 31-7202 - Short description of the issue – Begin typing and select from the menu the closest option which applies to the issue - Additional information or important details – additional information i.e. fridge number, etc. - Submit - Select Print and save with Equipment out of Service Form taped on the equipment |

References:

Standards for Blood Banks and Transfusion Services, Current Edition. AABB Press, Bethesda, MD.

Roback J (ed). Technical Manual, Current Edition. AABB Press, Bethesda, MD