Purpose	To provide detailed lab specific information for  β-Hydroxybutyrate LiquiColor performed at Einstein Medical Center Philadelphia.

Intended Use	To measure the presence and degree of ketosis in serum or plasma using the ARCHITECT c Systems.

Clinical Utility	Ketosis is a common feature in acutely ill patients. In subjects

suffering from starvation, acute alcohol abuse, or diabetes

mellitus, ketosis can result in severe life threatening metabolic

acidosis. The presence and degree of ketosis can be

determined by measuring blood levels of ß-hydroxybutyrate.

Ordinarily, ß-hydroxybutyrate is the ketoacid present in the greatest amount in serum. It accounts for approximately 75% of the ketone bodies which also contain acetoacetate and acetone. During periods of ketosis, ß-hydroxybutyrate increases even more than the other two ketoacids, acetoacetate and acetone, and has been shown to be a better index of ketoacidosis including the detection of subclinical ketosis.

In diabetics, the measurement of ß-hydroxybutyrate as well as the blood glucose is needed for the assessment of the severity of diabetic coma and is essential for the exclusion of hyperosmolar non-ketotic diabetic coma. Moreover, the insulin requirements are often based on the extent of the existing hyperketonemia shown by the blood levels of ß-hydroxybutyrate is therefore extremely important in the assessment of ketosis.

Enzymatic quantitation of ß-hydroxybutyrate by ß-hydroxybutyrate dehydrogenase has been reported. In the Stanbio method, ß-hydroxybutyrate (D-3-hydroxybutyrate) in the presence of NAD gets converted to acetoacetate and NADH at pH 8.5 by ß-hydroxybutyrate dehydrogenase (D-3-hydroxybutyrate dehydrogenase). At this pH, the reaction is favored to the right. The NADH produced reacts with INT in the presence of diaphorase to produce color at 505 nm.

Methodology		Enzymatic ß-Hydroxybutyrate Method

D-3-Hydroxybutyrate + NAD → Acetoacetate+NADH+H+

NADH + INT (oxidized) →   NAD + INT (reduced)

Specimen 		Serum or plasma collected with EDTA, heparin, or

Collection/ 		sodium fluoride can be used in the assay.

Preparation

Sample Stability	Serum or plasma ß-hydroxybutyrate levels are stable for 7 days at 2-8 °C.

Interfering 		No significant changes in values were observed when the

Substances	following analytes were added to serum containing 0.5 mM ß-hydroxybutyrate.

% Recovery

Glucose (2000 mg/dL)				 96

Acetoacetic acid (5 mM)				 96

Creatinine (5 mg/dL) 				           106

Ascorbate (3 mg/dL) 				           106

Bilirubin (10 mg/dL) 					 96

Uric Acid (16 mg/dL) 				           102

Triglycerides (417 mg/dL) 			           104

Cholesterol (314 mg/dL) 				 94

Lactic dehydrogenase (1515 U/mL) 			 93

Sodium lactate (96 mg/dL) 				 99

Reagent 		Reagents are supplied ready to use. Each box of reagent Preparation 		comes with one R1 bottle, and two R2 bottles.  Pour R1 into

a 20mL wedge and label with Reagent Name, date and technologist initials.  Pour both R2 bottles into a 20mL wedge and label with Reagent Name, date and technologist initials.

Reagent 		Reagents are stable stored at 2-8°C until expiration date

Storage/		on their respective labeling. Once opened, contamination

Stability		must be avoided.

Precautions	For In Vitro Diagnostic Use Only. Avoid skin contact with the reagents. If this occurs wash immediately with water.

Quality Control	A minimum of two levels of controls spanning the medical decision range are to be run once every 24 hours of assay use and when the standard calibrations are performed.  Refer to Chemistry QA Manual Quality Control Procedure CHQA01-002 and appendix CHQA01-002 Appendix A, CHQA01-002 Appendix B, CHQA01-002 Appendix C for detailed instructions, materials in use and LIS QC accession numbers.

Calibration		Calibration is stable for approximately 72 hours.  Calibration

is required with each change in reagent lot number or when QC is out of range.  Verify calibration with at least two levels of controls to the established quality control requirements.  Each box of reagent comes with a vial of calibration material.  Refer to CH03-100 Appendix A for Calibrator material.

Limitations		Lactic dehydrogenase and lactate have been shown to interfere with the assay. The incorporation of oxalic acid in this reagent eliminates this interference as reported.

Sensitivity		  Concentrations of ß-hydroxybutyrate of 0.18, 0.28, and 0.38

mM (1.8, 2.9, and 3.9 mg/dL) can be clearly distinguished at

the 99.9% confidence limit.

Dilution		Specimens with Beta Hydroxybutyrate values exceeding

4.5 mmol/L are flagged and may be diluted using the the Manual Dilution Procedure.

Manual dilutions should be performed as follows: Use water to dilute the sample. The operator must enter the dilution factor in the patient or control order screen. The screen uses this dilution factor to automatically correct the enzyme activity value by multiplying the result by the entered factor. If the operator does not enter the dilution factor, the result must be multiplied by the appropriate dilution factor before reporting the result.

Performance 		ß-hydroxybutyrate is linear to 4.5 mM (46.8 mg/dL)

Characteristics

Unit of Measure    mmol/L

Analytical 	    The Analytical Measurement Range of the assay is

Measurement 		0 to 4.5mmol/L

Range

Normal 	    0.02 mmol/L – 0.27 mmol/L

Range

| Reporting Results | Results are interfaced to the LIS and verified by the tech or auto-verified by the LIS as appropriate. Results higher than 13.5 mmol/L are reported as >13.5 mmol/L |
| --- | --- |

Approval Signatures:

| Date | Printed Name | Signature |
| --- | --- | --- |
| 2/3/2016 | Jennifer Lore, MFS, MT Chemistry Supervisor |  |
| 2/3/2016 | Nancy A. Young, M.D., FCAP # Medical Director |  |

History Review

| Date Reviewed | Reviewed By | Revisions |
| --- | --- | --- |