I.	PURPOSE

To confirm any WBC below 2.0x103 or above 25,000/cumm by slide review, manual methodology, or by dilution.

II.        CONTROLS Cell-Chex Levels I and II controls are used at Elkins Park for Manual WBC

counts. (See HEM30-002.02 for storage, stability, and preparation information)

Frequency use of control assay:

- Run Level 1 and Level 2 controls prior to patient analysis.

- Process the control the same as a patient sample

- Document quality control on HEM30-003 / HEM30-004 Form E2.

III.	PROCEDURE

- WBC < 2.0x103

- Smears are made and evaluated for a "WBC count" when outside of acceptable automated range.

- Scan to confirm 1st time WBC value of <2.0x103 footnote comment confirmed by “manual methodology and slide reviewed.”

B.	WBC Above 25,000:

The linearity of the DxH800 WBC is 400,000/cumm.

The linearity of the LH500 WBC is 200,000/cumm

WBCs over the limits of linearity are confirmed by dilution:

- Dilute blood using 1 part blood to 4 parts DILUENT

- Manual Dilution on Coulter

- Mix well and process sample within 30 minutes of dilution on the Coulter (run twice).

To ensure accurate dilution was made, check that the RBC and HGB are ±0.09/cumm (RBC) and ±0.2 (HGB) when multiplied by 5.

- Multiply the diluted WBC by 5 and record this value, once confirmed for dilution accuracy.  If the x5 dilution does not adequately lower the count within linear range, adjust the dilution and multiply by the appropriate factor.  Footnote WBC CONFIRMED BY DILUTION.

- COMPETENCY ASSESSMENT:

Annual Competency assessments for all trained operators.  For new employees, 6 months after initial training, and second 6 month after training then annually thereafter.  Documentation is placed in the employee’s file.

### V.	REFERENCES

- Beckman Coulter UniCel DxH 800 user manual; December 2009.

- NCCLS.  Clinical Laboratory Technical Procedure Manuals-Third Edition; Approved Guideline. (GP2-A3, 1996).

Approval Signatures:

| Date | Printed Name | Signature |
| --- | --- | --- |
| 12/28/2015 | David Hinkle, MT, ASCP Hematology Supervisor |  |
| 12/28/2015 | Vanessa Rawlings Elkins Park Supervisor |  |
| 12/28/2015 | Vivian Arguello, MD Section Director of Hematology |  |
| 12/28/2015 | Nancy A. Young, MD, FCAP Medical Director |  |

History Review

| Date Reviewed | Reviewed By | Revisions |
| --- | --- | --- |