1.0	Purpose

1.1	To ensure the continual and accurate operation of automatic temperature recording devices

2.0	Scope

2.1	This procedure provides the steps necessary to ensure continuous monitoring of the temperatures of Blood Bank refrigerators, freezers, and incubators.

3.0	Specimen

3.1	Patient specimens are not required for this procedure.

4.0 	Equipment

4.1	Automatic temperature recorder charts

4.1.1	Blood Refrigerators range: 		1C to 6C

4.1.1.1 	Reagent Refrigerator range:     	2to 6C.

4.1.2	Freezer range:	< -18C

4.1.3	Platelet incubator:    20C to 24C

4.2	Calibrated thermometer graduated in the range of intended use

4.3	Automatic temperature recorder ink

4.4	10% glycerol

4.5	50% Ethylene Glycol

5.0	Quality Control

5.1	The temperature reading on the recorder shall not exceed a 2C variance from the temperature indicated on the internal thermometer.

5.2	The temperature charts are reviewed weekly by the Lead Tech or Supervisor.

6.0	Procedure

6.1 	The sensor probes are immersed in a container or solution that has the same heat transfer characteristics of the component being stored.

6.1.1	Immerse refrigerator sensor probes in 10% glycerol solution

6.1.2	Immerse the sensor probe from plasma freezers in 50% ethylene glycol.

6.1.3	Attach the sensor probe in the platelet incubator to the side of the interior of the incubator.

6.2	Change the automatic temperature recorder chart every Monday morning. (The recorder chart has a seven (7) day cycle.

6.3	Remove the old chart from the recorder.

6.4 	Date and initial the old chart.

6.5	Date and initial the new chart.

6.5.1	Insure that the new chart has been previously stamped with the Temperature Record Stamp.

6.6	Replace the old chart with the blank chart. Align the chart wheel to the present day and time.

6.8	Notify the Lead Tech/Supervisor if any of the following conditions exist:

6.8.1	 Discrepancy between recorder and internal thermometer exceeds 2 C.

6.8.2	 Motor mechanism not functioning as indicated by the incorrect date or time on the chart.

6.8.3	 Recorder pen fails to write.

6.8.4	 Unusual temperature patterns recorded on the chart.

6.11	Document pronounced temperature variations directly on the chart, including reason, initials, and date.  (e.g. “alarm check,”  “door ajar”)

7.0 	Interpretation

7.1 	This section does not apply to this procedure.

8.0	Reporting Results

8.1 	This section does not apply to this procedure.

9.0	Procedure Notes

9.1	When temperatures exceed the allowable variation:

9.1.1	Check to see if the door is ajar.

9.1.2	Check that the fan is operational and that air is circulating.

9.1.3	Keep the door closed whenever possible.

9.1.4	Notify the Lead Tech or Supervisor.

    - Notify EMCP Maintenance at extension 6-6840

9.5.1.	EMC-EP maintenance at extension 6655

    - Transfer the contents of the refrigerator or freezer to another monitored refrigerator or freezer or temporary storage container

      - Place a “BB05-005 Form M :Do Not Use” form on equipment

9.1.6.2     Fill out all information on the BB QC05-005 Form L:  Equipment Repair Form

9.2	If the Automatic Temperature Recorder is not functioning (clock mechanism or thermometer device):

9.2.1	Record temperature visually observed from interior thermometer at least every four (4) hours onto BBQC05-005 Form N: Emergency Temperature Monitoring Form.

9.2.2	If temperature exceeds the allowable limits, follow section 9.1.

10.0	Limitations of the Procedure

10.1	This section does not apply to this procedure.

- References

  - Reference 1

  - Reference 2

12.0	Records

12.1	Charts from the automatic temperature recorders are retained in the Blood Bank for ten years.

13.0	Attachments/Appendix/Forms/Documents

13.1	BBQC05-005 Form L :Equipment Repair Form

13.2. 	BBQC05-005 Form N: Emergency Temperature Monitoring Form

13.3	BB05-005 Form M :Do Not Use” form

Approval Signatures

| # Date | # Printed Name | # Signature |
| --- | --- | --- |
| 6/20/14 | Pettina Walton Blood Bank Supervisor | New Leadership-N/A |
| 2/24/16 | Vanessa Rawlings Laboratory Supervisor Elkins Park |  |
| 6/20/14 | Manjula Balasubramanian, MD Section Director, Blood Bank | New Leadership-N/A |
| 6/20/14 | Nancy Young, MD FCAP Medical Director | New Leadership-N/A |

Annual Review:

| Date Reviewed | Review By | Revisions |
| --- | --- | --- |
| 09/27/2013 | Elizabeth Pailloz | Added range for Reagent Frig 2-8C 6.1.1. deleted 250ml volume glycerol 9.1.5 Clarified EMCP and EMC-EP |
| 11/25/2013 | E. Pailloz | Document control forms |
| 6/20/14 | Pettina Walton Tara Cameron | New Leadership added. No Revisions. |
| 4/29/15 | Pettina Walton Tara Cameron | No revisions |