- PURPOSE

  - Procedure to operate the Sysmex XN-10 analyzers.

- PRINCIPLE

The Sysmex XN-3000 is an integrated system that incorporates two hematology analytical modules as well as an automated slidemaker/stainer.

The analytical module (XN-10) is a quantitative automated hematology analyzer for in vitro diagnostic use in determining 31 whole blood diagnostic parameters and 7 body fluid diagnostic parameters.  Examination of the numerical and/or morphological findings of the complete blood count by the physician are useful in the diagnosis of disease states such as anemias, leukemias, allergic reactions, viral, bacterial, and parasitic infections.

The analyzer performs hematology analysis according to the hydrodynamic focusing (DC Detection), flow cytometry method (semiconductor laser), and SLS-hemoglobin method.

The device counts and sizes red blood cells (RBC) and platelets (PLT) using electronic resistance detection.  Hematocrit (HCT) is measured as a ratio of the total RBC volume to whole blood using cumulative pulse height detection.  Hemoglobin (HGB) is converted to SLS-hemoglobin and read photometrically.

The white blood cell (WBC) count, differential (DIFF), reticulocytes (RET) nucleated red blood cells (NRBC) and fluorescent platelets (PLT-F) are all evaluated using flow cytometry with a semiconductor laser exploiting the differences in cell size, complexity and RNA / DNA content.  Forward scattered light provides information on blood cell size and Lateral Scattered Light provides information on the cell interior such as the size of the nucleus.  Lateral fluorescent light intensity increases as the concentration of the stain becomes higher.  By measuring the intensity of the fluorescence emitted, information is obtained on the degree of blood cell staining.  Fluorescent light is emitted in all directions.  The XN detects the fluorescent light that is emitted sideways.

The Sysmex SP-10 is a fully automated hematology slide preparation and staining system.  Whole blood specimens are mixed and aspirated and a wedge type blood smear is prepared using hematocrit information from the Sysmex XN to determine optimum smearing criteria. The dried smear is automatically loaded into an individual slide cassette and is then advanced to the staining area. In the staining area, stain and buffer are dispensed into the cassette at operator-defined intervals.

The system also provides a manual mode operation where pre-made smears may be added to be stained. The unit is self-monitoring and alarms when operation is interrupted.

Slides prepared by the Sysmex SP-10 are used for differentiation and morphologic evaluation of cellular elements of whole blood.

- SPECIMEN AND WORKSHEET SPECIFICATIONS

- Collection Tubes

  - Becton Dickinson (B-D) #367861 purple-top tube containing 7.2 mg EDTA (K2).  Approximate draw 4 ml.

  - Becton Dickinson (B-D) #365974 Microtainer, purple top, containing EDTA (K2).Approximate draw 250-500 microliters, capillary collection.

  - Becton Dickinson (B-D) #367878 green-top tube containing 95 USP units of sodium heparin.  Approximate draw 6.0 mL.

    - When running body fluid specimens aliquot the specimen into a 12X75 plastic tube labeled with at least two patient identifiers.

NOTE:  Sodium Citrate may be used when EDTA platelet clumping or platelet satellitism is noted on the EDTA specimen. Use Sodium Citrate results for platelet counts and WBC counts, and result all other parameters from original tube. Multiply instrument PLT and WBC result by 1.1 to correct for anticoagulant dilution.

- Specimen Handling

  - Specimen volumes required

    - Optimal draw is a 12 x 75 tube filled to capacity

    - A minimum of 1 mL of whole blood is required for sampler analysis.

    - Manual analysis whole blood mode

      - Closed tube – 1 mL

      - Open tube – 300 μL

      - Open microtube – 160 μL

    - Manual analysis body fluid mode

      - Closed tube – 1 mL

      - Open tube – 300 μL

      - Open microtube – 160 μL

    - Manual analysis – SP-10

      - Closed tube smear and staining – 1 mL is optimal, 200 μL is aspirated.

      - Microtainer mode – 300 μL minimum volume, 60 μL is aspirated.

- Stability/Storage

  - Stored at 4-8oC, EDTA blood samples with normal results may be analyzed up to 24 hours without significant loss of differential stability.

  - Sample stability at room temperature is 24 hours.  Samples stored at room temperature may exhibit an increase in MCV after 24 hours, which may be minimized by refrigeration.

  - Allow refrigerated samples to come to room temperature, and mix well for a minimum of 30 minutes and mix well (at least 20 times) prior to analysis.

NOTE:  Do not place CBC and Diff samples on a mechanical rocker.  Constant rocking may alter white cell membranes, resulting in false interpretive messages

- Specimen acceptability

  - Unacceptable specimens including those listed below must be redrawn:

    - A 2.5 mL. EDTA tube filled less than one-half full is unacceptable

    - Clotted samples or those containing clots, fibrin strands, or platelet clumps.  All specimens will be checked visually for obvious clots prior to sampling by the analyzer.

    - Grossly hemolyzed samples

    - Samples drawn above an IV line

    - QNS specimens

    - Characteristics that may affect test results: lipemia, icterus, and cold agglutinins.

- Patient Preparation

  - None required.

WARNING:  All patient specimens should be considered potentially infectious and must be handled with precautions used for human blood, as described in CDC recommendations and in compliance with the Federal OSHA Blood borne Pathogen Standard, 29 CFR part 1910.1030.  Follow specimen handling as outlined by laboratory safety policy.

Recommended:  Wear gloves and a lab coat. Wear safety glasses if there is a risk of splash.

## STANDARDS AND CONTROLS

- XN CHECK

  - Manufactured by Streck, available as a tri-level package.

  - 8 mL Whole blood commercial control per vial used to monitor performance of the XN analyzers.

  - Formulation

    - Consists of human red and white blood cells with a platelet component suspended in fluid medium.

  - Storage

    - Store vials at 2-8oC

    - Do not freeze or expose to excessive heat.

  - Stability

    - Unopened and properly stored, XN CHECK is stable until the expiration date printed on the unopened vial.

    - Open vial stability is 7 days when promptly refrigerated after each use.

    - Record the date on each vial upon opening or cap piercing.

    - Heat or freezing can damage XN CHECK without gross visible changes.  Moderate hemolysis can be normal.  Deterioration is suspected when the mean of the control results is not within the assay expected ranges after appropriate troubleshooting.

    - If deterioration is suspected, call the Sysmex Technical Assistance Center.  1-888-879-7639 (1-888-8SYSMEX)

- XN CHECK BF

  - Manufactured by Streck, available as a bi-level package.

  - Body Fluid commercial control used to monitor performance of the XN analyzer body fluid analysis mode.

  - Formulation

    - Each vial contains 3 mL of control material

  - Storage

    - Store vials at 2-8oC

    - Do not freeze or expose to excessive heat.

  - Stability

    - Unopened and properly stored, XN CHECK BF is stable until the expiration date printed on the unopened vial.

    - Open vial stability is 30 days when promptly refrigerated after each use.

    - Record the date on each vial upon opening or cap piercing.

    - Heat or freezing can damage XN CHECK BF without gross visible changes.  Deterioration is suspected when the mean of the control results is not within the assay expected ranges after appropriate troubleshooting.

    - If deterioration is suspected, call the Sysmex Technical Assistance Center.  1-888-879-7639 (1-888-8SYSMEX)

  - SP-10 Quality Control

    - Daily, examine a stained smear from the routine workload for smear and stain quality.

    - Document results on appropriate log HEM50-004 FORM C

WARNING:	POTENTIALLY INFECTIOUS MATERIAL.

The human blood used in XN CHECK is non-reactive for Hepatitis B Surface Antigen and negative for antibodies to HIV-1, HIV-2, and Hepatitis C Virus using FDA specified techniques.  However, no current tests can assure the absence of these pathogens.  XN CHECK should be considered potentially infectious and must be handled with precautions used for human blood as described in CDC recommendations and in compliance with the Federal OSHA Blood borne Pathogen Standard, 29CFR, 1910.1030.

- Calibrators

- XN CALTM: for use in calibrating the analyzer for WBC, RBC, HGB, HCT, PLT,  and RET

  - XN CAL Storage and Stability

    - Store the calibrator in a dark refrigerator at 2-8oC

    - Unopened and properly stored, XN CAL is stable until the expiration date printed on the unopened vial.

    - Open vial stability is 4 hours.

- XN CALTM PF: for use in calibrating the analyzer for PLT-F (platelet count obtained from the PLT-F channel)

  - XN CAL PF Storage and Stability

    - Store the calibrator in a dark refrigerator at 2-8oC

    - Unopened and properly stored, XN CAL PF is stable until the expiration date printed on the unopened vial.

    - Open vial stability is 4 hours.

WARNING:	POTENTIALLY INFECTIOUS MATERIAL.

The human blood used in XN CHECK is non-reactive for Hepatitis B Surface Antigen and negative for antibodies to HIV-1, HIV-2, and Hepatitis C Virus using FDA specified techniques.  However, no current tests can assure the absence of these pathogens.  XN CHECK should be considered potentially infectious and must be handled with precautions used for human blood as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29CFR, 1910.1030.

- CALIBRATION and PRECISION

- Initial calibration is performed during installation by the Sysmex Field Service Representative.  Perform calibration as needed, e.g., when QC data is fluctuating.  However, if the abnormality in the QC analysis data was caused by an error in the analyzer, degradation of the reagent, or degeneration of the control blood, do not perform calibration.  Calibrators traceable to reference methods are used in the calibration of the analyzer.

- The laboratory must verify calibration every six months or on an "as-needed" basis to ensure accuracy of system.  Calibration verification is also required if one or more of the following occur:

  - Critical parts are replaced.

  - Controls show an unusual trend or are outside of acceptable limits and cannot be corrected by maintenance or troubleshooting.

  - When advised by Sysmex Field Service Representative.

- Calibration verification may be performed by review and documentation of commercial control and X-BarM QC data, proficiency testing results and patient control testing results.  The operator may calibrate the following parameters using XN CAL and XN CAL PF calibrator:  WBC, RBC, HGB, HCT, PLT, PLT-F and RET.

Before calibration, ensure that the XN is both clean and precise.

- Precision Check

  - Perform routine maintenance on the analyzer and perform a background count to ensure counts are within acceptable limits.

  - Verify that there is sufficient volume of all reagents.  Precision and Calibration procedures will be aborted if the XN runs out of reagent.

  - Obtain a sample of fresh normal whole blood.  Do not use commercial controls or calibrators for precision.  The blood donor specimen should:

    - Be from a healthy person who is not taking any medication

    - Have morphologically and numerically normal CBC.

    - Be drawn in a potassium EDTA anticoagulant tube using proper collection technique.

    - Have a minimum of 2.5 mL of sample.

  - On the main unit, check the Status indicator LED.  Confirm the LED is green indicating the analyzer is Ready

  - If the tube holder has not ejected out, press the mode switch

  - Select the Change Analysis Mode button on the control menu and select Whole Blood

  - Select [OK] to close the dialog box

  - Select the Analyzer menu button on the control menu

  - Select [Calibration] – [Precision Check]

  - Mix the vial containing the sample – 10 end-over-end inversions confirming cell button is dispersed

  - Place the vial in the sample tube holder

  - Press the start switch on the analyzer

  - The analysis is automatically performed 11 times consecutively with the tube holder pulled into the analyzer

    - The tube holder will slide out when analysis is complete

  - The results are displayed in the [Precision Check] analysis dialog box.

    - If the analysis results do not satisfy conditions for normal results, or if results are outside acceptable limits, the test numbers of the tests that must be repeated are displayed.  Select and redo the manual analysis.

  - When all analysis results satisfy the conditions, select [OK] in the dialog box.

  - Select [Yes] to record passing precision results in the precision check history.

NOTE:  If an error occurs during analysis and the analysis can no longer continue, stop precision check.  Once the error is cleared, redo the manual analysis.

- Calibration – XN CAL

  - On the main unit, check the Status indicator LED.  Confirm the LED is green indicating the analyzer is Ready

  - If the tube holder has not ejected out, press the mode switch

  - Select the Change Analysis Mode button on the control menu and select Whole Blood

  - Select [OK] to close the dialog box

  - Select the Analyzer menu button on the control menu

  - Select [Calibration] – [Calibrator Calibration]

  - Mix the vial containing the calibrator according to package insert

  - Place the vial in the sample tube holder

  - Press the start switch on the analyzer

    - The analysis is automatically performed 11 times consecutively with the tube holder pulled into the analyze

    - The tube holder will slide out when analysis is complete

  - The results are displayed in the [Calibrator Calibration] analysis dialog box.

  - If the analysis results do not satisfy conditions for normal results, or if results are outside acceptable limits, the test numbers of the tests that must be repeated are displayed.  Select and redo the manual analysis.

  - When all analysis results satisfy the conditions, select [Calibration] in the dialog box.

  - Select [OK] to display results in the [Calibrator Calibration] execution dialog box.

  - Select the check box to include the calibration parameter in the calibration exercise, clear the check box to exclude the parameter in the calibration exercise.  If a parameter meets all of the following criteria, the check box will automatically be selected:

    - 80% < New Rate < 120%

    - New Rate – Current Rate <  +5

      - Range Value < Max Range

    - Acceptable Limit < Delta Percent < Service Limit

  - If a parameter meets all of the conditions and the Delta Percent is less than the Acceptable Limit, it is excluded from calibration as there is no need for calibration.

  - If a parameter does not meet all of the conditions and the Delta Percent is greater than the Acceptable Limit, the calibration cannot be performed.  Calibration is performed with the parameter excluded.

  - Selecting the check box enables you to manually enter a value in [New Rate (%)].  A range of 80% to 120% may be entered.

  - Select [OK] to update the compensation rates.  The calibration process is logged in the calibrator calibration history.

- Calibration – XN CAL PF

  - On the main unit, check the Status indicator LED.  Confirm the LED is green indicating the analyzer is Ready

  - If the tube holder has not ejected out, press the mode switch

  - Select the Change Analysis Mode button on the control menu and select Whole Blood

  - Select [OK] to close the dialog box

  - Select the Analyzer menu button on the control menu

  - Select [Calibration] – [Calibrator Calibration (PLT-F)]

  - Mix the vial containing the calibrator according to package insert

  - Place the vial in the sample tube holder

  - Press the start switch on the analyzer

  - The analysis is automatically performed 11 times consecutively with the tube holder pulled into the analyzer

    - The tube holder will slide out when analysis is complete

    - The results are displayed in the [Calibrator Calibration (PLT-F)] analysis dialog box.

  - If the analysis results do not satisfy conditions for normal results, or if results are outside acceptable limits, the test numbers of the tests that must be repeated are displayed.  Select and redo the manual analysis.

  - When all analysis results satisfy the conditions, select [Calibration] in the dialog box.

  - Select [OK] to display results in the [Calibrator Calibration (PLT-F)] execution dialog box.

  - Select the check box to include the calibration parameter in the calibration (PLT-F) exercise, clear the check box to exclude the parameter in the calibration exercise.  If the parameter meets all of the following criteria, the check box will automatically be selected:

    - 80% < New Rate < 120%

    - New Rate – Current Rate <  +5

    - Range Value < Max Range

    - Acceptable Limit < Delta Percent < Service Limit

  - If the parameter meets all of the conditions and the Delta Percent is less than the Acceptable Limit, it is excluded from calibration as there is no need for calibration.

  - If the parameter does not meet all of the conditions and the Delta Percent is greater than the Acceptable Limit, the calibration cannot be performed.  Selecting the check box enables you to manually enter a value in [New Rate (%)].  A range of 80% to 120% may be entered.

  - Select [OK] to update the compensation rate.  The calibration process is logged in the calibrator calibration history.

NOTE:  If an error occurs during analysis and the analysis can no longer continue, stop precision check.  Once the error is cleared, redo the manual analysis.

- QUALITY CONTROL

Quality control is performed in order to monitor an analyzer’s performance over time.  XN CHECK  and XN CHECK  BF is the material used to monitor the performance of the XN analyzer.  Quality control should be run in accordance to licensing agency regulations.  It should be noted that for troubleshooting purposes, additional control runs may be necessary.  To QC the SP-10, examine a stained smear from the routine workload for smear and stain quality on a daily basis.  Document results on appropriate log.

- XN CHECK  Commercial Controls Instructions for Use

- Remove vials from refrigerator and allow them to come to room temperature (18-25oC), for approximately 15 minutes.

- Mix vials by gentle end-to-end inversion until the cell button in the bottom of the vial is completely suspended.

- XN CHECK  BF Commercial Body Fluid Controls Instructions for Use

      - Remove vials from refrigerator and allow them to come to room temperature (18 – 25oC) for approximately 15 minutes.

      - Mix vials by gentle end-to-end inversion until the cell button in the bottom of the vial is completely suspended.

- Manual analysis – SP-10

  - Closed tube smear and staining – 1 mL is optimal, 200 μL is aspirated.

  - Microtainer mode – 300 μL minimum volume, 60 μL is aspirated

WARNING:	POTENTIALLY INFECTIOUS MATERIAL.

The human blood used in XN CHECK is non-reactive for Hepatitis B Surface Antigen and negative for antibodies to HIV-1, HIV-2, and Hepatitis C Virus using FDA specified techniques.  However, no current tests can assure the absence of these pathogens.  XN CHECK should be considered potentially infectious and must be handled with precautions used for human blood as described in CDC recommendations and in compliance with the Federal OSHA Bloodborne Pathogen Standard, 29CFR, 1910.1030.

- Frequency of Control use and review

XN CHECK control levels:  1,2,3 will be run on 1st shift at 10am.

XN CHECK control levels:  1,2,3 will be run on 2nd shift at 6pm.

XN CHECK control levels:  1,2,3 will be run on 3rd shift at 2am.

XN CHECK BF control levels:  1,2 will be run on 1st shift in the Manual BF mode as needed.

XN CHECK BF control levels:  1,2 will be run on 2nd shift in the Manual BF mode as needed.

XN CHECK BF control levels:  1,2 will be run on 3rd shift in the Manual BF mode as needed.

SP-10 QC slide will be evaluated daily on First shift.

NOTE:  You can periodically display a message to prompt the user to perform quality control tasks through the QC Settings Menu.

- Registering and modifying a QC file – lot information input

  - Select [QC File] Icon

  - Select TAB for analyzer from bottom of QC File screen

  - Select File number to be registered.

  - Select [Register] button on toolbar

    - Enter lot information

      - Material

      - Lot Number

      - Expiration Date

      - Select [Restore]

    - Browse XN QC Limits folder on XN-IPU Desktop

    - Select file for QC to be registered

    - Select Open.

    - Sysmex Range Limit %’s will automatically upload to the file

    - Repeat for each level of XN CHECK, XN CHECK BF to be registered and for each module in the XN configuration

  - To modify an existing QC File, select the QC File and [Modify] from the toolbar.  Update the Lot No, Exp. Date as appropriate.

  - Perform parallel studies between production lot and new lot prior to production lot expiration.

- XN CHECK QC Analysis

  - Place the vial containing control blood in the rack.

  - Place rack on sampler unit; sampler unit will auto-start.

  - Results will be plotted on the L-J Chart as well as the Radar Chart for review.

- XN CHECK BF Analysis

  - Check the Status indicator LED on the analyzer to confirm analyzer is in ready state.

  - If the tube holder is not ejected, press the mode switch.  Tube holder will slide out.

  - Select the Change Analysis Mode button on the control menu.

  - Select [Body Fluid] mode.  Analyzer will automatically perform Autorinse.

  - Select [OK]

    - Place thoroughly mixed vial in tube holder, press start switch.

  - If vial barcode is unreadable, select the analyzer menu button on the control menu.

    - Select [QC Analysis]

    - From the list of QC files, select the file to be analyzed.  Judgment dialog box will open automatically.

    - Place thoroughly mixed vial in tube holder, press start switch.

  - When analysis is complete, analysis results are displayed.  User should review results and either accept or cancel the run.  Accepting the run will transfer the results to the L-J Chart and the Radar Chart for review.

- Auto set Targets

  - Parallel test new controls by analyzing the chosen levels of control, selected per lab policy QC protocol, a minimum of twice a day for 5 days prior to expiration or previous lot.  After a minimum of 10 data points are accumulated, auto set the targets.

  - Select QC Chart

  - Select [Range] and set cursors so that every data point is included

  - Select [Register]

  - Highlight all parameters and select [Auto Setting]

  - Confirm that the check box for TARGET ONLY is set.  Do not select the check box for LIMIT.

  - Select [OK]; the target for each parameter will be calculated and set for the duration of the QC lot.

  - Repeat steps for each new lot of QC being moved into production.

  - Confirm the target set falls within the range of means provided on the XN Check assay sheet provided.

  - Reviewing Quality Control Results

    - QC File screen

      - Allows for review of the latest QC results in Radar Chart format for the QC file that is selected in the list.

      - Any point exceeding the upper or lower limit is marked with a red “X”.

    - QC Chart screen

      - Allows for review of detailed graph data of all QC runs for selected file.

      - Analysis data is plotted cumulatively and displayed in the chart area as a line graph.

      - Any point exceeding the upper or lower limit is marked with a red “X”.

      - User must scroll up and down through the chart to view all parameters for each run.

      - Select [Range] to set a main cursor and a sub-cursor so that data between the two cursors can be manipulated.

      - Statistics may be analyzed over any selected range.

        - Targets may be auto-set for the selected range.

        - To cancel range mode, select [Range] on the toolbar again or exit QC Chart mode.

      - QC charts may be overlaid on top of each other for comparison.

      - Select [Compare QC Files] to view QC charts registered to a single analyzer.  This will compare the new lot with the current lot.

      - Select [Compare Analyzers] to compare QC files for the same material registered to different analyzers.

  - When QC fails control limits follow the instructions in procedure HEM20-004.01 Hematology Quality control program.

- Quality Control Management

  - From the QC Chart view, select the [Manage] button on the toolbar.

  - Specify whether a QC run should be excluded from quality control

    - Select [Not Manage] to exclude data from the following:

    - Statistical computations (SD, Mean, CV)

    - Variable target computation

    - Number of data points = n

    - An open circle will be displayed on the L-J Chart when the QC run is not managed or excluded and is not connected by a line to the adjacent QC runs.

    - A comment may be added to the QC data selected by the cursor

    - Select [Input Any Comment] to input a free text comment.

    - Select [Fixed Comments] to use a comment from a list of preset comments in the QC settings menu.

    - Select [OK]

    - A comment bubble will be displayed when a comment exists for a QC run.

    - The comment will be visible in the comment display area when the cursor is placed on the QC run.

- Recording and Storage of QC Data

Monthly the lead Technologist or designee will print the QC data.  The data is automatically transmitted to Sysmex Insight, and a monthly report will be printed upon completion of the peer group evaluation.  The peer group evaluation will be reviewed and signed by the supervisor.

  - Printing and saving QC Data

  - Select QC Files Icon and highlight file to output.

  - Select QC Chart Icon

  - Set Range of points to output by clicking [Range] and capturing the points with the cursors

  - Select [output] to print the selected chart to either GP or LP

  - Select [file] to save the data to removable media

- InsightTM Quality Assurance Program (QAP)

The Insight account number is 26604

The XN serial # is 23153, and SN# 23157

The lead technologist is responsible for saving the data to a USB memory device and submitting by due date in the event of an SNCS connection failure.

  - Each lot has 2 data submission dates, approximately every 30 days for the 84-day dated product.

  - Data may be managed in the XN-IPU and/or in Insight.  See Insight User Manuals.

  - Insert flash drive into USB port on the IPU’s hard drive.

  - Select the QC file you want to output, click [File], [Output in Sysmex Insight].  Save the file to the flash drive.

  - Repeat for each file needing Insight submission.

  - Properly eject the flash drive from the IPU.

  - At a networked PC, establish connection with the Insight program via www.sysmex.com/us and submit the data.  Contact the Insight team with questions at: 1-888-879-7639 (1-888-8SYSMEX).

- X-barM Moving Patient Averages

  - Activating / deactivating X-barM Control

- Select the analyzer menu button on the control menu

- Select X-barM Setting

- Click [Execute] to perform X-barM Control, Click [Cancel] to deactivate.

- Click [OK]

- X-barM Moving Patient Averages

  - Establishing X-barM Limit%

    - Sysmex data center suggests using 200 data points representing 4000 samples in 20 patient size batches.  Data will be collected over multiple reagent lots and over at least one month including all types of patient samples normally encountered.

  - Batch size and review frequency

    - Complete this section with your lab’s batch size and chart review frequency.  X-barM can be monitored in lieu of a retained patient sample for a longitudinal control if 100 or more patients are run each day.  Common batch size is 20; however, the Sysmex data center suggests using a larger batch size to allow about six points to be plotted per 24 hour period.  Include when and whether X-barM will be turned off for specific groups of patient specimens to avoid QC error messages related to population shifts.

Our batch size for X-barM is to be determined as per process outlined above patient samples per batch.  Each point on the X-barM graph represents one batch.

Supervisor or Lead Technologist will review X-barM charts every 14 days.

  - Activating / deactivating X-barM Control

    - Select the analyzer menu button on the control menu

    - Select X-barM Setting

    - Click [Execute] to perform X-barM Control, Click [Cancel] to deactivate.

    - Click [OK]

- Equipment and Materials

  - Supplies

          - De-ionized water

          - Lint-free plastic lined lab wipes

          - Gauze

          - Test tubes

          - Plastic squeeze bottles

          - CELLCLEAN® AUTO

          - Sysmex reagents

          - Commercial controls; XN CHECKTM, XN CHECKTM BF

          - Alcohol prep pads, isopropyl.  Used to clean SP-10 spreader glass

          - Microscope slides, frosted with rounded / clipped corners

            - 76 x 26 mm; 0.9 – 1.2 mm thick

        - Sysmex Reagents

          - Sysmex reagents and CELLCLEAN AUTO are used on the Sysmex XN-Series modules.

          - All reagents are used at room temperature and are to be used within the manufacturer’s expiration date on each container.

          - Record date received and date opened on container.

          - All reagents are azide free and are intended for in vitro diagnostic use only.  Do not ingest.

XN REAGENTS				OPEN EXPIRATION

CELLPACK DCL				60 Days

CELLPACK DST				60 Days

CELLPACK DFL				60 Days

SULFOLYSER				60 Days (1.5L)

90 Days (5.0L)

Lysercell WNR				60 Days

Fluorocell WNR				90 Days

Lysercell WDF				90 Days

Fluorocell WDF				90 Days

Fluorocell RET				90 Days

Fluorocell PLT				90 Days

SP REAGENTS

Stain –Wright Stain

Buffer – pH 6.8

Methanol

CELLPACK DCL		60 Days

  - Diluents

          - CELLPACK DCL: Whole blood diluent for use in hematology analyzers and for use as a rinsing agent for the spreader glass, sample pipette, and piercer on the SP-10.

            - CELLPACK DCL Storage

              - Store at 2o-35oC away from direct sunlight.

              - If frozen, thaw and mix thoroughly before using.

            - CELLPACK DCL is clear and colorless.  If it is showing signs of contamination or instability such as cloudiness or discoloration, replace.

            - CELLPACK DCL Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 60 Days.

          - CELLPACK DST (DST): Concentrated diluent of reagent for use in hematology analyzers.

            - CELLPACK DST Storage

              - Store at 2o-35oC away from direct sunlight.

              - If frozen, thaw and mix thoroughly before using.

            - CELLPACK DST is clear and colorless.  If it is showing signs of contamination or instability such as cloudiness or discoloration, replace.

            - CELLPACK DST Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 60 Days.

          - CELLPACK DFL (DFL):  Whole blood diluents for use in hematology analyzers; used in combination with Fluorocell™ RET for the analysis of reticulocytes, or with Fluorocell PLT for the analysis of platelets by flow cytometry method using a semiconductor laser.

            - CELLPACK DFL Storage

              - Store at 2o-35oC away from direct sunlight.

              - Do not use the reagent if it is suspected to have frozen.

              - Replace the reagent if it is showing signs of contamination or instability such as cloudiness or discoloration.

            - CELLPACK DFL Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 60 Days.

        - Lysing Reagents

          - SULFOLYSER (SLS): Reagent for the automated determination of hemoglobin concentration of blood.  Sulfolyser is a lysing reagent that releases the hemoglobin to be measured by the SLS hemoglobin method.

            - SULFOLYSER Storage

              - Store at 1o-30oC away from direct sunlight.

              - Allow the container to equilibrate to environmental temperature (15-30o) prior to use.

              - Do not use the reagent if it is suspected to have frozen.

              - Replace the reagent if it is showing signs of contamination or instability such as cloudiness or discoloration.

            - SULFOLYSER Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 60 Days (1.5L) or 90 Days (5L).

          - Lysercell WNR: Reagent product to be combined and used with Fluorocell WNR.  By hemolyzing red blood cells with Lysercell WNR and by differentiating white blood cells (non-basophil), basophils, and nucleated red blood cells with Lysercell WNR and Fluorocell WNR, the white blood cell count, basophil count, basophil percentage, nucleated red blood cell count, and nucleated red blood cell percentage are analyzed.

            - Lysercell WNR Storage

              - Store at 2o-35oC away from direct sunlight.

              - Allow the container to equilibrate to environmental temperature (15-30o) prior to use.

              - Do not use the reagent if it is suspected to have frozen.

              - Replace the reagent if it is showing signs of contamination or instability such as cloudiness or discoloration.

            - Lysercell WNR Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 60 Days.

          - Lysercell WDF: Reagent product to be combined and used with Fluorocell WDF.  By hemolyzing red blood cells with Lysercell WDF and dying the white blood cell component with Fluorocell WDF, the counts and percentages of neutrophils, immature granulocytes, lymphocytes, monocytes, and eosinophils are analyzed.

            - Lysercell WDF Storage

              - Store at 2o-35oC away from direct sunlight.

              - Allow the container to equilibrate to environmental temperature (15-30o) prior to use.

              - Do not use the reagent if it is suspected to have frozen.

              - Replace the reagent if it is showing signs of contamination or instability such as cloudiness or discoloration.

            - Lysercell WDF Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 90 Days.

        - Staining Reagents

          - Fluorocell WNR: Used to stain the nucleated cells in diluted and lysed blood samples for determination of white blood cell count, nucleated red blood cell count and basophil count in blood.

            - Fluorocell WNR Storage

              - Store at 2o-35oC in a dark place.

              - Do not use the reagent if it is suspected to have frozen.

            - Fluorocell WNR Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 90 Days.

          - Fluorocell WDF: Used to stain the leukocytes in diluted and lysed blood samples for determination of differential count in blood.

            - Fluorocell WDF Storage

              - Store at 2o-35oC in a dark place.

              - Do not use the reagent if it is suspected to have frozen.

            - Fluorocell WDF Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 90 Days.

          - Fluorocell RET: Used to stain the reticulocytes in diluted blood samples for the assay of reticulocyte count, reticulocyte percent in blood.

            - Fluorocell RET Storage

              - Store at 2o-35oC in a dark place.

              - Do not use the reagent if it is suspected to have frozen.

            - Fluorocell RET Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 90 Days.

          - Fluorocell PLT: Used to stain the platelets in diluted blood samples for the assay of platelet counts in blood.

            - Fluorocell PLT Storage

              - Store at 2o-35oC in a dark place.

              - Do not use the reagent if it is suspected to have frozen.

            - Fluorocell PLT Stability

              - Unopened, it is stable until expiration date printed on the container.

              - Opened, stable for 90 Days.

        - Cleaning Agent

          - CELLCLEAN AUTO: Detergent for fully automated hematology analyzers.  To be used as a strong alkaline detergent to remove lysing reagents, cellular residuals, and blood proteins remaining in the hydraulics of the analyzer.  For use as a cleaning fluid for the hematology analyzers and the SP-10.

            - CELLCLEAN AUTO Storage

              - Store at 1-25oC, away from direct sunlight.

              - Do not use the reagent if it is suspected to have frozen.

            - CELLCLEAN AUTO Stability

              - Unopened, it is stable until expiration date printed on the container.

        - Stain / Buffer for SP-10

          - Romanowsky stain (Wright or Wright-Giemsa)

            - Used to fix and stain blood cells for the purpose of differentiation and morphologic evaluation

          - Phosphate Buffer pH 6.8

WARNING:	Stain contains methanol. Methanol is flammable and poisonous. Potential human carcinogen. May be fatal if ingested. Harmful if inhaled. Causes irritation to eyes, skin and respiratory tract.

          - Methyl Alcohol (Methanol), anhydrous for SP-10

            - Used for cleaning of the staining system and cassettes.  Also may be used for optional staining pre-fix.

WARNING:	Methanol is flammable and poisonous. Potential human carcinogen. May be fatal if ingested. Harmful if inhaled. Causes irritation to eyes, skin and respiratory tract.

          - Deionized water, pH ~7.0

- XN Reagent Replacement

  - When the reagent runs out during analysis, the analysis is paused and an error message appears in the analyzer area of the Control menu.

  - Display the [Reagent Replacement] dialog box to replace the reagent.

  - Select the help button on the control menu

  - Select [Execute]

        - Remaining Reagent Volume indicator appears

          - Replacing a new diluent / hemolytic agent

            - Display the [Reagent Replacement] dialog box

            - Remove the cap from the new reagent container

            - Confirm the reagent has not expired

            - Input the reagent code (barcode)

            - Place the cursor in the reagent code field

            - Scan the reagent code on the outer box of the new reagent with the hand-held barcode reader or manually enter the reagent code

            - Select [OK]

            - Remove the cap from the old reagent container.

            - Pull out the dispensing set straight up.

            - Insert the dispensing set straight into the new container.

            - Close the cap.

            - Select [Execute]

            - Reagent replacement starts.  When complete, the dialog box closes automatically.

          - Replacing CELLPACK DST with an RU-20

            - Display the RU-20 Maintenance menu.

            - Select [Replace Reagent]

            - Remove the cap from the new reagent container.

            - Confirm that reagent has not expired

            - Input the reagent code (barcode)

            - Place the cursor in the reagent code field.

            - Scan the reagent code on the outer box of the new reagent with the hand-held barcode reader or manually enter the reagent code.

            - Select [OK]

            - Remove the cap from the old reagent container

            - Pull out the dispensing set straight up.

            - Insert the dispensing set straight into the new reagent container.

            - Close the cap

            - Select [Execute]

            - Reagent replacement starts.  When complete, the dialog box closes automatically.

          - Replacing Dye

            - Display the [Reagent Replacement] dialog box.

            - Prepare the new reagent cartridge.

            - Confirm the reagent has not expired.

            - Open the top front cover.

            - Pull up the cover from the reagent that is to be replaced.

            - When the dye solution cover is pulled up, a Help dialog box appears in the IPU screen.

            - Remove the old reagent cartridge from its holder.

            - Install the new reagent cartridge into the holder

              - Make sure the color of the label on the new reagent cartridge matches the color of the dye cover and install.  Analyzer will beep as confirmation of new reagent installation.

              - If the wrong reagent is installed, the analyzer beeps repeatedly and the Help dialog box appears in the IPU screen.

            - Pull down the cover on the reagent until you hear a click.

              - When the cover is pulled down, the Help dialog box closes automatically.

              - The ID of the new reagent is read automatically and the information is registered.

            - Close the top front cover.

              - Reagent replacement starts.

              - When complete, the reagent replacement window closes automatically.

          - SP-10 Reagent Replacement

The following is a list of replacement messages and the reagent requiring replacement:

| Message | Reagent |
| --- | --- |
| *DCL not filled | CELLPACK DCL |
| *Stain 1 not filled in Chamber 1 Stain |  |
| *Stain 1 not filled in Chamber 2 Stain | 2nd stain (if using 2 stain method) |
| *Rinse water not filled(internal chamber not filled) | Deionized water |
| Replace Rinse water (external container empty) | Deionized water |
| Replace buffer | Buffer |
| Replace methanol | Methanol |

* Reagents with internal chambers. Other reagents use bottle sensors.

- When a reagent container is empty, an alarm sounds and a dialogue box displays. Press [OK] to silence the alarm and close the dialogue box.

- Press [Help] icon and follow the corrective action message.

- When replacing a reagent with an internal chamber, press [OK] to clear the action message and reset. For reagents with bottle sensors, the error clears when the reagent is replaced or filled.

- Replace reagent using clean technique. Avoid placing spout kit or sensor on a contaminated surface.

## INSTRUMENTATION

A.	Identification

|  | NAME | SERIAL # | SYSTEM ID# | “LIVE” DATE |
| --- | --- | --- | --- | --- |
| # AEMC | XN-3000L | 23153 |  | 6/7/2016 |
|  | XN-3000R | 23157 |  | 6/7/2016 |

- Operating Procedure

  - Start-Up Procedure

          - Checks prior to turning on

            - Visual inspections of analyzer / system / reagents

              - Place completed samples into final storage area for the lab

              - Remove any items that may interfere with operations

              - Gather and re-locate all empty racks to designated processing or sample loading area

              - If applicable, verify waste container is empty

              - Verify network / host connections are properly working

              - Ensure that the towers (slide supply cassettes) have sufficient slides. Fill with glass slides.

                - Remove the tower to be filled.

                - Remove the metal insert from the end of the tower.

                - Fan the slides to prevent them from adhering to each other and place them with the frosted end up and towards the open end of the tower.

                - Replace the metal insert and replace the tower with the frosted end of the slides towards the back of the analyzer.

                - Verify sufficient reagent supply is nearby

                - Fill the cassette supply table with clean, dry single cassettes. The Sysmex logo should be forward and the notch at the bottom must be away from you (or to the left).  The supply table holds up to 100 cassettes.  A minimum of 8 cassettes are required for start-up.

          - Turning ON the entire system

            - Verify that all power switches for each device are in the ON position

            - Press the start-up switch on the sampler to power ON the entire system

          - Log on to the XN-IPU

            - When the logon dialog box appears, enter user name and password

          - Analyzers and SP-10 self-checks

            - XN: Initialization of the mechanical parts; Rinse; Temperature stabilization; Background Check (up to 3 times)

| XN Acceptable Background Counts |  |
| --- | --- |
| Parameters | Acceptable Limit |
| WBC-N | 0.10 x 103/ μL |
| WBC-D | 0.10 x 103/ μL |
| RBC | 0.02 x 106/μL |
| HGB | 0.1 g/dL |
| PLT-I | 10 x 103/ μL |
| PLT-F | 3 x 103/ μL |

            - SP-10: System check to evaluate internal stored data files; shutdown check to determine whether shutdown was performed properly, a mechanical initialization sequence.

          - Analyze Quality Control Material

      - Patient Sample Processing

          - System Analysis (sampler analysis)

            - Make sure the analyzer and the sampler are in READY state

            - Check that tube holder has retracted into the analyzer, press mode button if necessary

            - Place sample(s) in rack(s) in right sampler pool (analyzer side)

            - Rack(s) will auto-start.

            - Samples will run, results will be displayed in the IPU.

            - On-Board rules engine will determine repeat or reflex testing

            - Rack will run in reverse to perform repeat or reflex testing.

            - If smear is required, rack will be transported to SP-10 via analysis line and samples will be aspirated by SP-10.

            - If no smears are required, rack will be transported to the left sampler pool without stopping at the SP-10.

            - Remove the rack from the left sampler pool when analysis in completed.

          - Manual Analysis – XN

            - Check the status of the analyzer.  Confirm the analyzer is ready.

            - Press the mode switch to eject the tube holder.

            - Select the Change Analysis Mode button on the control menu

            - Select analysis mode

              - [Whole blood] is selected when whole blood is being analyzed

              - [Low WBC] Select this to perform low WBC analysis on whole blood

              - [Pre-Dilution] select when running 1:7 pre-diluted blood.

              - Select [OK]

              - Select Manual Analysis button on the control menu

              - Input sample ID or select [Read ID]

              - Select [OK]

              - Properly mix the specimen and place in the tube holder

                - If running microtainer, remove the cap using caution to avoid splattering.

              - Press the start switch on the analyzer

                - The tube holder will slide in and the sample will be aspirated

                - When the analysis is complete, the tube holder slides out

              - Remove the sample, repeat steps for additional samples

              - Review results in IPU to determine whether repeat or reflex testing is required.  Rerun sample if required.  Make smear if required.

          - Body Fluid Analysis – XN

            - Check the status of the analyzer.  Confirm the analyzer is ready.

            - Press the mode switch to eject the tube holder.

            - Select the Change Analysis Mode button on the control menu.

            - Select [Body Fluid]

            - Select [OK]

              - The analyzer will automatically perform a background check up to three times

            - Select the Manual Analysis button on the control menu

            - Input the sample ID or select [Read ID]

            - Select [OK]

            - Properly mix the specimen and place in tube holder.

              - If running microtainer, remove the cap using caution to avoid splattering

            - Press the start switch on the analyzer

              - The tube holder will slide in and the sample will be aspirated

              - When the analysis is complete, the tube holder slides out

            - Remove the sample

            - Perform Background check prior to running additional samples if indicated

            - Return analyzer to Whole Blood mode prior to running whole blood samples.

          - Off-line analysis; The sampler for the analyzer, or the sampler for the SP-10 is separated from the transport line of the overall system and operated as a standalone device

            - Press mode switch on the sampler

            - Verify sampler is in READY state

            - Place the rack in the right pool of the sampler for the analyzer that you wish to use.

            - Transport begins automatically

            - Remove the rack after analysis is complete

            - Press the mode switch on the sampler

          - SP-10 Manual Mode – Smear and Stain

            - Press [Conv. Int.] on the SP-10 main menu screen

            - Press [Interrupt]

            - Select [Return]

            - Select [Manual] on the SP-10 main menu screen

            - Op Mode is set to [Smr + Sta], Smpl. Tube is set to [Closed]

            - Input Specimen information, Sample ID, HCT, select number of slides to be made

            - Thoroughly mix the sample and place in 10th rack position

            - Place the rack so that the sample aligns with the tube gripper and that the left end of the rack fits the label on the sampler

            - Select [Start]

              - Analysis will begin

              - When the tube is returned to the rack, remove the rack

            - Press [Return]  [Conv. Int.] [Stop Int.]

          - SP-10 Manual Mode – Stain Only

            - Select [Manual] mode

            - Select [Op. Mode], [Stain].  Do not proceed until [START] button is green.

            - Place labeled, unstained blood films into cassettes at the front of the cassette supply table on the right side of the analyzer.  If multiple slides are to be stained, place them in consecutive cassettes.

            - Press [Start]

            - The cassettes will be fed to the stain table and the smears will be stained.  An empty cassette will follow to indicate the end of the run.

          - SP-10 Micro Mode

            - Select [Manual]

            - Choose [Op. Mode], [Smr.+Stain] and set [Smpl. Tube] to [micro]

            - Input Specimen information, Sample ID, HCT

            - Place the thoroughly mixed uncapped microtainer in the micro collection sample tube holder.

            - Select [Start]

            - Micro tube will be lowered into position and sample will be aspirated.

            - When aspiration is complete, micro tube will be returned to home position and should be removed.

          - SP-10 Smear Only – No staining occurs

            - Smear mode may be used in System, Single, or Manual Modes. To access Smear Mode:

              - Press [Settings] on the main screen. (A password may be required.)

              - Press [Select], [Cond.], [Mode].

              - Press [Op. Mode] and select [Smear]. Press [RETURN] and [YES] to accept the settings.

                - To use Smear Only in System Mode

                - Place bar coded samples in a Sysmex rack.

                - Place the rack in the right pool of the Sampler Unit.

                - Racks will auto-start.

                - Racks are transported to the XN analyzer and then to the SP-10 where a smear will be prepared when appropriate criteria are met.

                - To use Smear Only in Off Line Mode: The sampler for the analyzer, or the sampler for the SP-10 is separated from the transport line of the overall system and operated as a standalone device

                - Press mode switch on the sampler

                - Verify sampler is in READY state

                - Place the rack in the right pool of the sampler for the analyzer that you wish to use.

                - Transport begins automatically

                - Remove the rack after analysis is complete

                - Press the mode switch on the sampler

                - To use Smear Only in Manual Closed Mode:

                - Press [Conv. Int.] on the SP-10 main menu screen

                - Press [Interrupt]

                - Select [Return]

                - Select [Manual] on the SP-10 main menu screen

                - Op Mode is set to [Smear], Smpl. Tube is set to [Closed]

                - Input Specimen information, Sample ID, HCT, select number of slides to be made

                - Thoroughly mix the sample and place in 10th rack position

                - Place the rack so that the sample aligns with the tube gripper and that the left end of the rack fits the label on the sampler

                - Select [Start]

                - Analysis will begin

                - When the tube is returned to the rack, remove the rack

                - Press [Return] [Conv. Int.] [Stop Int.]

                - Remove the rack when sampling is complete.

Return Setting to SMEAR + STAINING

Press [Settings], [Select], [Cond.], [Mode], [Smr + Sta.]. Press [RETURN] and [YES].

Note:   If setting is left at Smear, the system will perform smear only in all modes.

- Maintenance

Document all maintenance procedures on HEM50-01.001 FORM B for the, SP-10.  Maintenance performed on the XN will be automatically tracked in the maintenance history, but must be documented on HEM50-004.01 FORM A.

  - Cleaning the XN Daily

          - Confirm analyzers are at ready.

          - Confirm tube holders are retracted into the analyzers.

          - Open the Analyzer menu

          - Select maintenance

          - Select cleaning

          - Place tube of CELLCLEAN AUTO in the tube holder

          - Press start button

          - Analyzer will perform cleaning cycle for 20 minutes, and will automatically return to ready state upon completion.

  - In lieu of cleaning cycle a Shutdown can be performed as needed.

          - CELLCLEAN AUTO is used to shut down, or cleans the entire system.  Refer to the XN-3000 Instructions for Use for detailed, illustrated procedures.

          - Confirm analyzers, sampler unit and SP-10 are at ready.

          - Obtain 2 empty racks

            - Place one tube of CELLCLEAN AUTO in rack one, position 8.  This rack will shut down the SP-10.

            - Place 2 tubes of CELLCLEAN AUTO in rack two, positions 9 and 10. This rack will shut down the XNs.

            - Place racks on sampler unit, sampler unit will auto-start.

            - Document the XN maintenance on HEM50-004.01 FORM A

            - Document shutdown on the SP maintenance log HEM50-004.01 FORM B.

      - SP-10 Daily

          - Clean Spreader Glass:  Power must be on to perform this maintenance – may be performed prior to Shutdown, or after Start-up.

            - Press [Maint.] on the main screen. (Maintenance button is not available during routine operation.)

            - Press [Spreader Glass] and the “Spreader Glass Replace” screen displays.

            - Press [OK] to move the smear unit forward.

            - Remove the left tower for easier access to the spreader glass.

            - Wipe the spreader in one direction with an alcohol prep pad.

            - Replace the tower so that the frosted end of the slides are towards the back of the analyzer.

            - Press [OK] to return the smear unit to the home position.

            - Press [OK] to reset the spreader glass cycle counter or [CANCEL] to allow the cycle count to continue.

            - Press [RETURN].

          - Clean Single Cassettes

            - Place cassettes in a bin with open end up.

            - Pour methanol over the cassettes, filling them.

            - Swish the methanol and pour off into designated container for reuse.

            - Invert cleaned cassettes on absorbent material to dry.

- Note:	Methanol may be reused for cleaning cassettes up to three (3) times.  Discard when appropriate, according to laboratory policy.

          - SP-10 Shutdown

            - Perform Shutdown 1 and 2 Daily (alternate Days)

              - Press [SHUTDOWN] on the main screen.

              - Press [Shutdown 1 or 2]

              - The shutdown screen displays the number of cassettes and amount of methanol required for the shutdown process. Ensure that required amounts are available.

              - Place a tube of CELLCLEAN AUTO in position 10 of a Sysmex rack.

              - Place the rack so that the tube is lined up with tube gripper.

              - Press [OK].

              - When the process completes, the SP-10 turns off automatically.

              - To restart the SP-10, press the green button on the right side.

      - SP-10 Weekly

          - Clean DI water/Buffer containers

            - If re-usable containers for deionized water and/or buffer are used, empty weekly.

            - Rinse with methanol and allow to dry.

            - Fill with fresh deionized water or buffer.

      - As Needed Maintenance

              - Refer to the XN-3000 Instructions for Use for detailed and illustrated instructions for performing as needed maintenance.

- PROCEDURAL NOTES AND CALCULATIONS

  - If making a dilution of a patient specimen and running in XN Whole Blood mode, multiply the parameters by the dilution factor

  - If correcting the HGB or HCT due to interfering substances, recalculate and correct the affected indices:

          - MCHC   = HGB / HCT x 100

          - MCH   = HGB / RBC x 10

          - MCV      = HCT / RBC x 10

        - If a sodium citrate tube is used for EDTA induced platelet clumping, multiply the citrate platelet count and WBC count by 1.1 to correct for anticoagulant dilution. Use all other parameter results from the original EDTA tube.

        - Lipemic Specimens- Refer to HEM01-006.

        - Cold Agglutinins – Refer to HEM01-008.

        - Use the Help function on the SP-10 when errors and messages display.  Use the error icon on the XN to display help menu.

        - While slides are being processed on the SP smear table, the START key may not be available for manual mode processing

        - During normal processing of slides on the SP-10, Maintenance., Settings, and Shutdown functions are not available.

        - Current settings for XN and SP-10 should be recorded and maintained in the XN-Series Resource Manual and the SP-Series Implementation Manual.

        - Current on-board rules should be exported and saved on external storage device.  A printout of the rules should be inserted in the XN-Series Resource Manual.

        - Do not place samples on a mechanical rocker.  Excessive mixing may alter white cell membranes resulting in false interpretive messages.

        - For troubleshooting specifics refer to the Sysmex XN-3000 Instructions for Use.

- REPORTING RESULTS

A.	Normal Values

WBC:     			 --	4,000 - 11,000/cumm

RBC:

Male	--	4.7 - 6.1 million/cumm

Female	--	4.1 - 5.0 million/cumm

Hemoglobin:

Male	--	14.0 - 18 grams

Female	--	12.0 - 16 grams

Hematocrit:

Male	--	39 - 50%

Female	--	35 - 43%

MCV:

Male	--	81-96 cubic microns

Female	--	81-96 cubic microns

MCH:

Female   	--	28 – 32 pico grams

Male      	--    	25 – 32 pico grams

MCHC:      	--		32 - 36%

RDW:       	--		11.4 – 14.9%

MPV:       	--		7.4 – 10.4 cubic microns

LYMPH%     --		11 - 40%

MONO%		--		2 - 11%

GRAN%		--		44 - 79%

EOS%		--		0 – 5%

BASO%		--		0 – 1%

RETIC	       --		0.0-2.0%

Normal Differential: (Manual)

Band Cells	--	                        	3 - 11%

Segmented Neutrophils         	30 - 60%

Lymphocytes	--	              	16 - 49%

Monocytes	--	                       	 3 - 14%

Eosinophils	--	                        	 1 - 5%

Basophils	--	                        	 0 - 1%

Normal Platelet Count:	--	140,000 - 400,000/cubic millimeter

B.	Normal Mean Values Infants, Children, Adolescents

| ## Age | ## WBC (x103/mm3) | ## Hgb (g/dL) | ## Hct (%) | ## RBC (x106/mm3) | ## MCV (fL) |
| --- | --- | --- | --- | --- | --- |
| 1 d | 5.0 – 25.0 | 14.3 – 18.9 | 43 – 53 | 3.80 – 5.10 | 98 – 110 |
| 2 d | - | 17.3 – 21.5 | 49 – 65 | 4.80 – 6.00 | 98 – 117 |
| 7 d | - | - | - | - | - |
| 29 d | - | 10.7 – 15.3 | 33 – 46 | 3.75 – 4.95 | - |
| 1 y | 5.0 – 15.0 | 11.3 – 12.5 | 32 – 39 | 4.04 – 4.84 | 80 – 100 |
| 5 y | - | 11.7 – 13.7 | 34 – 40 | - | - |
| 10 y | - | 12.0 – 14.4 | 36 – 43 | 4.30 – 5.30 | - |

C.	Normal Mean Platelet Values, Infants

First day  	-- 192,000/cumm

2 - 3 days	-- 192,000/cumm

4 - 8 days	-- 213,000/cumm

9 - 13 days	-- 248,000/cumm

14-60 days -252,000/cumm

D.	Critical/ Urgent Values

Critical HGB        < 6.0 and ≥20.0

<31 days 		≤ 8.0

Platelets 		< 20,000 and ≥ 900,000

<31 days 		≤50,000 and ≥ 900,000

WBC 	≤2.0 and ≥ 100,000

Any blasts

Any parasite or organism seen.

- All critical/urgent values must be called and documented with the name (Last Name and First name/initial) of the individual receiving the result, extension called, date/time called, and the technologist initials reporting the critical as a result comment in ARE (Accession Result Entry).

- The clinical section supervisor or designee is notified if the appropriate person cannot be reached.

- REPORTING ABNORMAL RESULTS.

- Abnormal bloods

Any time abnormal bloods are recognized, the analysis should be verified by additional reference methods using other equipment or techniques for verification.

1. 	Leukocytes

a. 	Discrepancies may obviously appear in leukemia. Although the linearity ranges on the Sysmex XN3000 has been documented over 440,000, any WBC above 440,000, must be diluted to determine the correct WBC count. Also, correction of the RBC count and recalculation of indices based on the RBC count are then required.

b.	It should also be noted that carryover may be a factor if in the sequence of samples. This can be corrected if each of the suspect samples is analyzed twice and only the second value is reported.

c.	Platelets the size of small lymphocytes and megakaryocytes may be counted as WBC’s. If more than 5 giant platelets or megakaryocytes are seen per 100 WBC’s counted on a differential, a manual WBC must be performed.

.

2.	Erythrocyte parameters

a.	Agglutination of erythrocytes, no matter what the etiology, may produce decreased RBC counts, along with apparent macrocytosis. Discrepancies should be suspected when the MCH and MCHC appear to be outside the normal ranges. Additional studies should be done.

b.	Large numbers of large platelets or large numbers of white cells may interfere with the RBC count, MCV, RDW, and the associated calculated parameters.

3.	Thrombocyte parameters

a.	 Debris or particles from cellular breakdown in the platelet size range may be counted as thrombocytes. Rerunning as an optical platelet may obtain more accurate results.

- Platelet count linearity is 5,000,000. Higher counts are to be reported as greater than 5 million after being confirmed on a stained smear and called to the floor.

- LIMITATIONS OF PROCEDURE

- XN-Series Acceptable Linearity

| Parameter | Range | Units |
| --- | --- | --- |
| WBC | 0.01-440.0 | x103/μL |
| RBC | 0-8.60 | x106/μL |
| HGB | 0-26.0 | g/dL |
| HCT | 0-75.0 | % |
| PLT, PLT-F | 0-5000 | x103/μL |
| RET% | 0-25 | % |
| NRBC% | 0-600 | /100 WBC |
| WBC-BF | 0-79 | x103/μL |
| RBC-BF | 0-5.4 | x106/μL |

- Parameters that exceed these limits are flagged with @ beside the result.  The sample must be diluted, rerun and multiplied by the dilution factor.

- Note the use of dilution for linearity on the patient report.

- Possible Sample Interferences

- Specimens must be free of clots and fibrin strands.

- Marked changes in plasma constituents, (e.g., low sodium, extremely elevated glucose) may cause cells to swell or shrink.  The blood to anticoagulant ratio is important.

- Red cell fragments, microcytic RBC's, or white cell cytoplasmic fragments may interfere with automated platelet counts.  A fluorescent platelet may be performed to avoid this interference.

- Cold agglutinins produce spurious macrocytosis, elevated MCH's MCHC's, falsely decreased RBC counts and HCT's.  Rare, warm agglutinins produce the same spurious results as a cold agglutinin.

- Extremely elevated WBCs may cause turbidity and falsely increase the hemoglobin, in addition to RBC and HCT values.

- Red cells- If hemolysis is occurring in vivo, the instrument RBC may be flagged as low, reflecting the true circulating cells. If, however, the hemolysis is in vitro, the specimen may give falsely low RBC results. Cell counts due to in vitro hemolysis do not represent the number of circulating red blood cells.

- Giant platelets and clumped platelets may falsely elevate the WBC count and falsely decrease the platelet count. Platelet clumping and/or "platelet satellitism" can occur in specimens collected in EDTA.  This may falsely elevate the WBC count and falsely decrease the platelet count.

  - When results are flagged for platelet clumps check specimen for clots.

  - Vortex specimen for 2 minutes

  - Repeat analysis in open mode using the Platelet-F test order.

  - If flag persists perform smear review

- Abnormal paraproteins found in Multiple Myeloma patients can falsely increase the HGB.  To correct HGB perform plasma replacement.

- Severely icteric samples may falsely elevate the HGB value and related indices.  Make a 1:5 dilution with CELLPACK.

- Rocking specimen excessively, may affect the WBC differential.

- Megakaryocytes may falsely increase WBC counts on automated hematology analyzers.  Confirm count by performing smear review.

## REFERENCES

- Sysmex XN-3000 Instructions for Use (North American Edition), Sysmex Corporation, Kobe, Japan.

- Sysmex XN series Administrator’s Guide (North American Edition), Sysmex Corporation, Kobe, Japan

- Sysmex SP-10 Instructions for Use (North American Edition), Sysmex Corporation, Kobe, Japan.

- Clinical and Laboratory Standards Institute (CLSI).  Laboratory Documents: Development and Control; Approved Guideline; Fifth Edition. (GP2-A5, 2006).

- Sysmex America Inc., Lincolnshire, IL. XN CAL, XN CAL PF Hematology Calibrators:  Calibrators for Sysmex Hematology XN-Series Analyzers, package insert.

- Sysmex America Inc., Lincolnshire, IL. XN CHECK Hematology Control for Sysmex XN-Series Analyzers package insert.

- Sysmex America Inc., Mundelein, IL.  Sysmex Insight  Participant Overview Guide.

- Koepke, John.  Practical Laboratory Hematology.  Churchill Livingstone Inc. 1991.

p. 24-25, 36-39.

- Cornbleet J., Spurious results from automated hematology cell counters.  Lab Medicine. 1983;8:509-514.

- Sysmex Reagents of America, Inc.  MSDS sheets and reagent product inserts.

- College of American Pathologists (CAP) Hematology-Coagulation Checklist, July 2012.

- Stewart, Charles and Koepke, John.  Basic Quality Assurance Practices for Clinical Laboratories, Van Nostrand Reinhold, 1989, p 189.

- Gulati GL, Asselta A, Chen C. Using vortex to disaggregate platelet clumps, Laboratory Medicine, 28:665, 1997.

- Zhou X, Xiaoli W. Amikacin Can Be Added to Blood to Reduce the Fall in Platelet Count, American Journal of Clinical Pathology, 136:646-652, 2011.

Sysmex XN-3000, CELLPACK DCL, DST, DFL, FLUOROCELL, LYSERCELL, XN CHECK, XN CHECK BF, XN CAL, XN CAL PF, and Sysmex Insight are trademarks of the Sysmex Corporation.

All other trademarks are the property of their respective owners.

Approval Signatures:

| Date | Printed Name | Signature |
| --- | --- | --- |
| 6-6-2016 | David Hinkle, MT, ASCP Hematology Supervisor |  |
| 6-6-2016 | Vivian Arguello, MD Section Director of Hematology |  |
| ## 6-6-2016 | ## Nancy A. Young, MD, FCAP ## Medical Director |  |

History Review

| Date Reviewed | Reviewed By | Revisions |
| --- | --- | --- |