### PURPOSE

To confirm by alternate methodology or by dilution any platelet count above linearity of 4,701,000/cumm on the DxH800 and above linearity of 2,000,000 on the LH500, and to correlate the count with an estimate on a stained smear.

II.	PROCEDURE

A.	Platelet counts below 50,000 for the first time – confirm with footnote “SCAN” (smear scanned), results consistent with automated count OR FOOTNOTE PLATELET CLUMP PRESENT.

1.	Refer to procedure Manual Platelet Count, for methodology

B.	Platelet Counts Above 4,701,000/cumm on the DxH800 and above 2,000,000 on the LH500.

1.	Manual/Automated methodology

Perform platelet count manually (see above for reference) and enter footnote (Verified By Alternate Methodology) or confirm by SCAN smear scanned, results consistent with automated count – and result as >4,701,000/cumm (DxH800) or >2,000,000 (LH500) per instrument linearity.

2.	Dilution:

- Using a volumetric pipette make a 1:3 dilutions of 1 part blood to 2 parts Diluent

- Make sufficient dilution for a duplicate run (use 200 mcL blood to 400 mcL diluent).

- Mix well and sample on the Coulter twice.

- To assure accurate dilution was made, check that the RBC when multiplied by 3 is ±0.09 of the undiluted sample and that the Hgb is ±0.2 (when multiplied by 3) of the undiluted sample.

- Multiply the diluted platelet count by 3 and enter this value (once confirmed as in step "d").  Footnote the following message to quantitative platelet count” confirmed by dilution”.

III.		REFERENCE

- Beckman Coulter UniCel DxH 800 user manual; December 2009.

- NCCLS.  Clinical Laboratory Technical Procedure Manuals-Third Edition; Approved Guideline. (GP2-A3, 1996).

Approval Signatures:

| Date | Printed Name | Signature |
| --- | --- | --- |
| 10/26/15 | David Hinkle, MT, ASCP Hematology Supervisor | New Leadership- NA |
| 4/12/16 | Vanessa Rawlings, MHA, MT Elkins Park Supervisor |  |
| 11/18/13 | Vivian Arguello, MD Section Director of Hematology | New Leadership- NA |
| 6/1/12 | Nancy A. Young, MD, FCAP Medical Director | New Leadership- NA |

### History Review

| Date Reviewed | Reviewed By | Revisions |
| --- | --- | --- |
| 1/31/2014 | Jaclene Kokoszka | No revisions. |
| 10/6/2014 | Christine Eagle | New leadership - NA |
| 1/6/2015 | Christine Eagle T. Cameron | No revisions. |
| 10/26/2015 | David Hinkle | New leadership - NA |
| 1/6/2016 | David Hinkle | No Revisions. |