### PURPOSE

To confirm by alternate methodology or by dilution any platelet count above linearity of 5,000,000/cumm, for platelet, and 5,500,000 for Platelet-F and on the Sysmex XN-3000 and above linearity of 5,000,000 on the Sysmex 1000i, and to correlate the count with an estimate on a stained smear.

II.	PROCEDURE

A.	Platelet counts below 50,000 for the first time – confirm with footnote “SCAN” (smear scanned), results consistent with automated count OR FOOTNOTE PLATELET CLUMP PRESENT.

1.	Refer to procedure Manual Platelet Count, for methodology

B.	Platelet Counts Above 5,000,000/cumm, for platelet, and 5,500,000 for Platelet-F on the Sysmex XN-3000 and above 5,000,000 on the Sysmex 1000i.

1.	Manual/Automated methodology

Perform platelet count manually (see above for reference) and enter footnote (Verified By Alternate Methodology) or confirm by SCAN smear scanned, results consistent with automated count – and result as >5,000,000/cumm, for platelet, and >5,500,000 for Platelet-F on the Sysmex XN-3000 and >5,000,000 on the Sysmex 1000i.

2.	Dilution:

- Using a volumetric pipette make a 1:7 dilutions of 1 part blood to 6 parts DCL

- Make sufficient dilution for a duplicate run (ex.  Use 100 mcL blood to 600 mcL DCL).

- Mix well and sample on the Sysmex twice.

- To assure accurate dilution was made, check that the RBC when multiplied by 3 is ±0.09 of the undiluted sample and that the Hgb is ±0.2 (when multiplied by 3) of the undiluted sample.

- Multiply the diluted platelet count by 3 and enter this value (once confirmed as in step "d").  Footnote the following message to quantitative platelet count” confirmed by dilution”.

III.		REFERENCE

- Sysmex XN-3000 instructions for use , May 2014

- Sysmex Xs-1000i instructions for use, January 2014.

- NCCLS.  Clinical Laboratory Technical Procedure Manuals-Third Edition; Approved Guideline. (GP2-A3, 1996).

Approval Signatures:

| Date | Printed Name | Signature |
| --- | --- | --- |
| 7-22-2016 | David Hinkle, MT, ASCP Hematology Supervisor |  |
| 7-22-2016 | Vanessa Rawlings, MHA, MT Elkins Park Supervisor |  |
| 7-22-2016 | Vivian Arguello, MD Section Director of Hematology |  |
| 7-22-2016 | Nancy A. Young, MD, FCAP Medical Director |  |

### History Review

| Date Reviewed | Reviewed By | Revisions |  |  |
| --- | --- | --- | --- | --- |
| 7-22-2016 | David Hinkle Vanessa Rawlings | Major revision updated for Sysmex platelet linearity. Updated dilution Criteria. |  |  |