EINSTEIN MEDICAL CENTER-Hematology/Chemistry

subject:	Hematology/ Chemistry STAFF MEETING

ATTENDEES:	 Jennifer, David, Anil Smauel, Annette brabham; Ngoc To; Ngoc Lu; Andrea Barrette; Phyllis Braxton; Monica Kwon; karen chappelle; courtney le; shanice bobb; loretta hunt;lisa frisco; al schlank; dr. arguello

date:  September/October 15thth, 2019

AGENDA

| CATEGORY | TOPIC | ##### ANNOUNCEMENT / UPDATE | DISCUSSION |
| --- | --- | --- | --- |
|  | Mission Story | - Does anyone have a story they would like to share that connects to Einstein’s Mission Statement? | Nothing Discussed |
| TECHNICAL CHEMISTRY | - Honesty - Errors - Concentrated Buffer - C Side Mixers - Weekly Maintenance - Documentation - Pending logs - Fluid slide reminder - Off shift training - Procedures - Supplies - Pathology review | - We need to be honest with our mistakes/errors. If we mislabel a sample in the lab, we must complete a IQE immediately after notifying the nurse. These mislabel samples must be announced during our Daily Safety Huddles. We cannot hide our own errors but always announce when nurses make the error. - If the track has an error and samples spill on the track, or we have a lab accident of any type and a patient sample is compromised we need to let the nurse know immediately. We can’t hold off notify the nurse as we spend hours searching for a sample. - Preparation of the Concentrated Buffer. Please be careful with preparing the material. If it is prepared incorrectly it will affect the patient results. We have to shut the machine down, remove all the onboard buffer and flush the lines to clear it out. This has happened a few times. Also when preparing the solution, you need to indicate the date and your initials on the side of the buffer bottles. We did put in for a new ARM for the buffer, the new ARM will service both analyzers and will allow us not to have to make up any buffer manually again. - Check the Mixers on the C8000 instruments daily. Ensure that the green rod is attached. We were having trouble with QC and in the troubleshooting discovered that the entire rod was mixing. - When you perform weekly maintenance and you clean the mixers, the mixers come out so you can wipe them down. While the mixer is out, please use a Q-tip and distilled H2O and clean out the well. - We need to ensure that maintenance is being completed daily on the Abbott instruments. If the instrument is down and you are waiting on service, you can go into the maintenance section and comment at the bottom that the instrument is down, waiting on service. This will help us with ensure that we are compliant on this CAP requirement. - Competencies – you need proof of instrument maintenance. When you are doing daily or weekly maintenance you should sign in using your name as the user and print the screen shot of it completed with your name. - ICT Modules are only stabile onboard for 90 days or 20,000 tests (which ever comes first). The module on C5 expired and no one changed the module and continued to use the expired module. - Date and initial your calibrators and QC material. We can’t determine if QC or instruments issues are because the QC is old if it is not dated. - Do not run blood, cloudy, turbid urines on the Urinalysis analyzers. Follow the procedure on how to handle these types of specimens. When you test these samples on the instrument it causes the pipette and tubing to become clogged and causes issues. - When verifying your Urinalysis QC in Cerner, enter the number 2 times to ensure that you are not verifying QC that was left unverified from the previous shift. We have had QC being out of range and patient samples tested and released because QC was not verified. - You MUST perform your Urinalysis QC at the start of your shift. It is not acceptable to state that QC was overlooked, forgotten or I was to busy to run. - Document as much detail on issues in the communication log. It is hard to follow up on issues when we are not given much detail. - Read the communication log. Be familiar with issues that are documented. Read your emails, there are times that information is sent via email that you need to be aware of. - CAP Window opens on December 8th this year. We need to be CAP ready at all times. - 5 minute hand off communication. This is absolutely mandatory. Issues need to be recorded in the communication binder as well as verbally told to the next shift coming in. - Stop scribbling out on calibration logs, and documents. Place a line through the item and your initials. No use of white out or labels over documentation. - Working on validating NGAL on C4. (neutrophil gelatinase-associated lipocalin, lipocalin-2, siderocalin) It will be ordered by select departments only for grant purposes. It is the first diagnostic acute kidney injury biomarker which can be measured in human urine. ETA to go live is beginning of September. - I am updating the laminated QC sheets with the QC Numbers and Tests for all 5 analyzers with the most recent changes to assay locations. Effective 9/1/19 it is mandatory to use these laminated QC check sheets to ensure that all QC has be performed and all reagent is on board. Going forward, missed QC will not be acceptable as these sheets are a last check to ensure all QC is accounted for. - PCT TAT has been drastically decreased for Dayshift for the past 4 months. We broke out the RIA/ UA benches and this has not improved this past month. You need to be checking the Pending logs to ensure that all samples are accounted for in a timely manner. TAT for PCT is 90 minutes. I will be working with CP so these tubes are not spun in the back room and forgotten about. - Documentation on QC logs on the side of the instruments is a must. QC is being repeated but no documentation that it was out the first time. - If you perform QC and accidently mix up levels (ie. 1 for 2) you must repeat the QC that is out regardless of knowing that you mixed up the QC cups. - Overuse of QC material. We are going through QC material way to fast. This is due to the over pouring of the QC material when only a few drops are needed. The QC material is expensive and frequent ordering is causing for new lots and having to setup on the analyzers and in Cerner. - Lot to Lots are still not being completed for the 5 look backs. Documentation needs to be completed and placed in Jennifer’s door immediately. Do not place on the clip boards or in the binders. If there are no five look back samples available, you still need to complete the form for review and place in Jennifer’s door. - We will begin to track specimens from the time they are completed on the instrument to the time a tech releases or rejects the result in Instrument Manager. There is an issue with a delay in handling these results. We always need to have someone in Instrument Manager keeping up to date with the results. We can not have critical values or hemolyzed samples not called for hours. - Pending logs need to be pulled and printed at the beginning and end of the shift and reviewed. Through out the shift the pending log should be pulled at least electronically. Do not pull the pending log and leave it for the next day to review. The purpose of pending logs is to ensure the timely resulting of results. Make sure the logs are placed in the correct bins so when there are issues we can review them for following up on any issues. | Discussed |
| Technical HEMATOLGOY | - Communication between shifts - Pending logs throughout shift - Anti X a - Supplies | - At change of shift peding log should be reviewed by the both shifts. Any pending testing should be accounted for prior to leaving the bench to go home. - Reminder to check the pending logs at least hourly. Checking more frequently will help decrease TAT. You can set an auto refresh for the electronic log and review there frequently. - Anti X A go live on hold. Waiting for clinical education component - When supplies are received it is everyone’s responsibility to put them away. There are times the reagents are not delivered until after 4pm. Please be sure to follow the receiving log procedure. Do not just open a box and take what you need. | Discussed |
| General Hematology/ Chemistry Updates General Hematology/ Chemistry Updates Continued | Cleanliness Lunches and breaks Restocking bench/ Receipt of supplies Vendors | - Please make sure to clean your area after yourself. Shred OR orders for TEG; gloves; pipettes’ in trash not on the counter. Please watch what you are throwing in the sharps container: gloves, paper towels and drink bottles do not belong in these cans. - Lunch times- Dayshift please follow the schedule. Schedule is posted on the board above the Chemistry daily checklist binder. Breaks are to be 15 mins and only if work and staffing allows. Lunches are 45 minutes. Remember missed breaks should be approved by a Supervisor. All employees are now receiving the same lunch deductions and must follow the 45 min lunch schedule as discussed. - Restock benches for the next shifts. If something is out, please restock the shelves. - If you receive reagents, please use the log sheets. Expiration dates, lot numbers and quantity are necessary. Also make sure reagents and supplies are put away. Do not leave anything in the hallways. - If you see we are running low on supplies, please continue to notify Supervisors/ Lead Techs - Vendors need to visit the Kiosk in the entrance of the hospital prior to coming into the lab. This is a new security requirement. |  |
| DDR | Review of DDRs | DDR for September/October: 2 Critical platelets not correctly documented. Reminder to call prelim critical count and enter call back info then perform. When the smear is reviewed add scan to the comment and do not override the critical callback info See posted PSN entered each month that pertain to the lab on the Core Studer board | Reviewed |
| Goals | - Hematology Goals | - All goals have been met/ Exceeded. Keep up the great work. |  |
| EMPLOYEE ISSUES/ Competency | EMCP-employees due for competency Evaluations CON ED and electronic assignments | - Please remember it is your responsibility to provide the supervisor with all necessary documentation for your competency. Chanh, Loretta and Chris will still provide staff with the unknown samples. Anyone who is competent may observe and sign you off on the duties. It does not have to be Ashley, Loretta, or Chris. Must have completed prior to the end of your assigned month. The observer should be the individual initialing under direct observation and not the tech performing the test. - Failure to complete any annual compliance will result in accountability actions. - MediaLab/MedTraining/HealthStream/Competencies: You must complete the assignments on time. These are all a part of the annual eval. We should not have to remind anyone numerous times. Clarification to the Admin Procedure-It is for the employees calendar year Ex: if you are hired in February you can collect paperwork from February to February. | Discussed |
| HOSPITAL/ LAB NEWS HOSPITAL/ LAB NEWS CONTINUED | - Benefits (Healthy Steps) - Tobacco Surcharge - Manager Minutes | - All employees and spouse plan members will need to attest to their tobacco status in Redbrick by 12/31/19. The surcharge will be effective 1/20/20. - Employee and spouse plan members who use tobacco will either pay a $25 surcharge per pay. - Counseling Services Available 24/7 - Everyone faces challenges in their day-to-day lives, and at times those challenges may feel overwhelming. - Remind your staff that confidential counseling services are available 24/7 for all employees and their families, - regardless of whether they have Einstein health benefits. A Carebridge representative can be reached directly at - 800-437-0911, or employees can call LiveWell at 215-456-8484 and press #3 to reach Carebridge. Additionally, - staff can reach out to their human resources representative, or our Chaplaincy at 215-456-6055. - Einstein Part of Team to Acquire Hahnemann’s Residency Program - MossRehab Achieves Milestone on U.S. News & World Report’s Best Rehab Hospitals List - Einstein Represents Community in Hearing Before Government Officials - Einstein Philadelphia Hosts Biggest Blood Drive Since 2014 - Einstein Philadelphia Hosts Biggest Blood Drive Since 2014 - Accessing Domestic Violence Screening and Referral Process - Engaging Employees Through Psychological Safety | Discussed |
| GOALS | FY 2019 Pillar Goals and Tactics | Updated FY19 goals for the Lab require participation from all staff. Please review Studer boards for LEM Success information LEM Scorecard is hanging on the Studer board. Also attached. Areas of improvement are the ED TAT; Employee engagement; and overtime. Lab is trending upward but still hope to improve. | Reviewed |
| HUMAN RESOURCES | Open Positions Vacancies Scheduled Sick | - Open positions are posted monthly in the hallway on the former continuing education board. All three sites are posted. - If an appointment/procedure requires all day give 7.5 hours of scheduled sick. - If an appointment is only 30 min we will factor in travel time and have you come to work at least half a day. If the employee still wants an entire day they can take half ss and half pto. - The employee can also use an entire day and schedule all appointments for that day. The following do not require an entire day of scheduled sick Teeth cleaning Healthy Steps annual checkup All day ss is needed for Colonoscopies Procedures in which the employee maybe sedated | discussed |
| STUDER | SLR Studer | - What tools do you need to do your job? - Hematology waiting on engineering for AC diverter above diff scope. Also waiting on response regarding new Coag Desk. - Rounding – does everyone know what Rounding is? This is where your leaders (lead techs, supervisors, managers) ask you how everything is going, what is going well, what tools or supplies you need to do your job, any improvements you would like to suggest, and if there is anyone you would like to recognize. | Hematology: Sam reviewing the Benchtop availability. Sam will look into the AC diverter. TAT monitor to be hung third ticket entered. |
| EMPLOYEE RECOGNITION | Days to Days | - If you see a coworker deserving of a day to day for going the extra mile, please notify the supervisors the person and what they did that was extraordinary. I have requested more day to day cards. |  |