TITLE:    MEDLINE hCG Combo Pregnancy Test

PRINCIPLE:  The MEDLINE hCG Combo Pregnancy Test Cassette is a rapid chromatographic immunoassay for the qualitative detection of human chorionic gonadotropin (hCG) in serum or urine to aid in the early detection of pregnancy. The test utilizes a combination of antibodies including mouse monoclonal anti-hCG antibodies and goat polyclonal anti-hCG antibodies to selectively detect elevated levels of hCG. The assay is conducted by adding serum or urine specimen to the specimen well of the test cassette and observing the formation of colored lines. The specimen migrates via capillary action along the membrane to react with the colored conjugate. Positive specimens react with the specific colored antibody conjugates and form a colored line at the test line region of the membrane. Absence of this colored line suggests a negative result. To serve as a procedural control, a colored line will always appear at the control line region if the test has been performed properly.

CLINICAL SIGNIFICANCE:  Human chorionic gonadotropin (hCG) is a glycoprotein hormone produced by the developing placenta shortly after fertilization.  In normal pregnancy hCG can be detected in both serum and urine as early as 7 to 10 days after conception.  hCG levels continue to rise very rapidly, frequently exceeding 100mIU/mL by the first missed menstrual period, and peaking in the 100,000-200,000 mIU/mL range about 10-12 weeks into pregnancy.  The appearance of hCG in both serum and urine after conception, and its subsequent rapid rise in concentration during early gestational growth, make it an excellent marker for the early detection of pregnancy.

Personnel:

Medical Technologists/Medical Laboratory Scientists

REAGENTS AND EQUIPMENT:

A. Reagents and Materials Provided

- Test Cassette (Contains mouse anti-beta hCG antibody conjugated to colloidal gold and goat anti-alpha hCG antibody coated on the membrane.)

- Disposable pipettes

- Product Insert

B. Reagents and Materials not Provided

- Specimen collection container

- Timer

- Alta Positive and Negative controls

C. Storage and Stability

Store as packaged in the sealed pouch at 2-30°C. The test cassette is stable through the expiration date printed on the sealed pouch. The test cassette must remain in the sealed pouch until use. DO NOT FREEZE. Do not use beyond the expiration date.

D. Warnings and Precautions

- For professional in vitro diagnostic use only. Do not use after the expiration date.

- The test cassette should remain in the sealed pouch until use.

- All specimens should be considered potentially hazardous and handled in the same manner as an infectious agent.

- The test cassette should be discarded in a proper biohazard container after testing.

- The test cassette should not be reused.

SPECIMEN COLLECTION AND STORAGE:

A. Specimen:

Urine - A first morning urine specimen is preferred since it generally contains the highest concentration of hCG; however, urine specimens collected at any time of the day may be used. Urine specimens exhibiting visible precipitates should be centrifuged, filtered, or allowed to settle to obtain a clear specimen for testing.

Serum - Blood should be collected aseptically into a clean tube without anticoagulants. Separate the serum from blood as soon as possible to avoid hemolysis. Use clear non-hemolyzed specimens when possible

.

B. Transport Container:

Clean, dry container for urine, and a clean, dry tube for serum.

C. Specimen Storage:

Serum or urine specimen may be stored at 2-8°C for up to 48 hours prior to testing. For prolonged storage, specimens may be frozen and stored below -20°C. Frozen specimens should be thawed and mixed before testing.

D. Handling Precautions:

Patient samples, controls, and test devices should be handled as though they could transmit disease. Observe established precautions against microbial hazards.

CALIBRATION:

None Indicated

QUALITY CONTROL:

Internal Quality Control:

Internal procedural controls are included in the test. A red line appearing in the control region (C) is the internal procedural control. It confirms sufficient specimen volume and correct procedural technique. A clear background is an internal negative background control. If the test is working properly, the background in the result area should be white to light pink and not interfere with the ability to read the test result.

External Quality Control (hCG controls by Alere):

Positive and negative controls will be run with each shipment/new lot number and once every 30 days.  If values are not within range, repeat controls.  Contact a senior tech if Quality Control is still out after repeat testing.

- External Positive Control

Process the control as you would a patient sample.  A positive test is indicated by    Two distinct red lines. One line should be in the control region (C) and another line should be in the test region (T).

- External Negative Control

Process the control as you would a patient sample.  A negative test is indicated by one red line that appears in the control region (C). No apparent red or pink line appears in the test region (T).

Record these control results in the Laboratory Information System QC program.  Refer to the LIS procedure Manual for complete instructions.

STEPWISE PROCEDURE:

Allow the test cassette, serum or urine specimen and/or controls to equilibrate to room temperature (15-30°C) prior to testing.

- Remove the test cassette from the sealed pouch and use it as soon as possible.

2.   Place the test cassette on a clean and level surface. Hold the pipette vertically and transfer 3 full drops of serum or urine (approx. 100 μL) to the specimen well of the test cassette, and then start the timer. Avoid trapping air bubbles in the specimen well.

3.   Wait for the red line(s) to appear. Read the result at 3-4 minutes when testing a urine specimen, or at 5-6 minutes when testing a serum specimen. Do not interpret results after the appropriate read time. It is important that the background is clear before the result is read.

Interpretation of Test Results

POSITIVE*: Two distinct red lines appear. One line should be in the control region (C) and another line should be in the test region (T).

Note: A sample hCG concentration below the cut-off level of this test might result in a weak line appearing in the test region (T) after an extended period of time. A line in the test region (T) seen after the read time could be indicative of a low hCG level in the sample. If such results are seen, it is recommended that the test be repeated with a new sample in 48-72 hours or that an alternate confirmation method is used.

NEGATIVE: One red line appears in the control region (C). No apparent red or pink line appears in the test region (T).

INVALID: Control line fails to appear. Insufficient specimen volume or incorrect procedural techniques are the most likely reasons for control line failure. Review the procedure and repeat the test with a new test cassette. If the problem persists, discontinue using the test kit immediately and contact MEDLINE Customer Support at 1-800-MEDLINE.

*Note: The intensity of the red color in the test line region (T) will vary depending on the concentration of hCG present in the specimen. However, neither the quantitative value nor the rate of increase in hCG can be determined by this qualitative test.

Expected Values

Negative results are expected in healthy non-pregnant women and healthy men. Healthy pregnant women have hCG present in their serum and urine specimens. The amount of hCG will vary greatly with gestational age and between individuals. The MEDLINE hCG Combo Pregnancy Test Cassette has a sensitivity of 10 mIU/mL in serum and 20 mIU/mL in urine and is capable of detecting pregnancy as early as 1 day after the first missed menses.

Limitations

1. Very dilute urine specimens, as indicated by a low specific gravity, may not contain representative levels of hCG. If pregnancy is still suspected, a first morning urine specimen should be collected 48 hours later and tested.

2. False negative results may occur when the levels of hCG are below the sensitivity level of the test. When pregnancy is still suspected, a first morning serum or urine specimen should be collected 48 hours later and tested.

3. Very low levels of hCG (less than 50 mIU/mL) are present in serum and urine specimen shortly after implantation. However, because a significant number of first trimester pregnancies terminate for natural reasons, a test result that is weakly positive should be confirmed by retesting with a first morning serum or urine specimen collected 48 hours later.

4. This test reliably detects intact hCG up to 500,000 mIU/mL. It does not reliably detect hCG degradation products, including free-beta hCG and beta core fragments. Quantitative assays used to detect hCG may detect hCG degradation products and therefore may disagree with the results of this rapid test.

5. A number of conditions other than pregnancy, including trophoblastic disease and certain non-trophoblastic neoplasms including testicular tumors, prostate cancer, breast cancer, and lung cancer, cause elevated levels of hCG. Therefore, the presence of hCG in serum or urine specimen should not be used to diagnose pregnancy unless these conditions have been ruled out.

6. As with any assay employing mouse antibodies, the possibility exists for interference by      human anti-mouse antibodies (HAMA) in the specimen. Specimens from patients who have received preparations of monoclonal antibodies for diagnosis or therapy may contain HAMA. Such specimens may cause false positive or false negative results.

7. This test provides a presumptive diagnosis for pregnancy. A confirmed pregnancy diagnosis should only be made by a physician after all clinical and laboratory findings have been evaluated.

INTERFERING SUBSTANCES:

See MEDLINE hCG Pregnancy Test Cassette (Serum/Urine) Product Insert for interfering substances.

REFERENCE:

MEDLINE hCG Pregnancy Test Cassette (Serum/Urine) Product Insert (4/1/15)

Medline Industries, Inc., Mundelein, IL 60060