TITLE: ARK – Methotrexate Assay on Vitros 5600

Principle:

ARK Methotrexate Assay is a homogeneous immunoassay based on competition between

drug in the specimen and methotrexate labeled with the enzyme glucose-6-phosphate

dehydrogenase (G6PDH) for binding to the antibody reagent. As the latter binds antibody,

enzyme activity decreases. In the presence of drug from the specimen, enzyme activity increases

and is directly proportional to the drug concentration. Active enzyme converts the coenzyme

nicotinamide adenine dinucleotide (NAD) to NADH that is measured spectrophotometrically as

a rate of change in absorbance. Endogenous serum G6PDH does not interfere with the results

because the coenyzme NAD functions only with the bacterial enzyme used in the assay

Clinical Significance:

Methotrexate [N-[4[[(2,4-diamino-6-pteridinyl) methyl] methylamino]benzoyl]-L-glutamic

acid], formerly Amethopterin, is an antimetabolite used in the treatment of certain neoplastic

diseases, severe psoriasis, and adult rheumatoid arthritis.1-3 Methotrexate has the potential

for serious toxicity. Patients undergoing methotrexate therapy should be closely monitored so

that toxic effects are detected promptly.

Specimen Collection and Handling:

• Serum  is required.

• The sampling time of methotrexate will be dependent on dose, duration of infusion, and

clinical status of the patient. Consult treatment protocols for sampling time.

• Whole blood cannot be used.

• Blood collection must be performed with collection tubes compatible for use with therapeutic

drug monitoring (TDM).

• Do not induce foaming and avoid repeated freezing and thawing to preserve the integrity of

the specimen from the time it is collected until the time it is assayed.

• Fibrin, red blood cells, and other particulate matter may cause an erroneous result. Ensure

adequate centrifugation.

Specimens may be stored up to two weeks at 2 to 8°C. If testing will be delayed,

specimens may be stored frozen (≤ -10°C) prior to being tested. Care should be taken to limit the number of freeze-thaw cycles.

• Handle all patient specimens as if they were potentially infectious.

Reagents:

ARK Methotrexate Assay

Reagent –R1 - Antibody/Substrate

Rabbit polyclonal antibodies to methotrexate, glucose-6-phosphate, nicotinamide adenine

dinucleotide, bovine serum albumin, preservative,and stabilizers

Reagent –R2 -  Enzyme

Methotrexate labeled with bacterial G6PDH, buffer, bovine serum albumin, preservative, and stabilizers

Reagent Handling and Storage

ARK Methotrexate Assay reagents are provided liquid, ready to use and may be used directly

from the refrigerator. When not in use, reagents must be stored at 2–8°C (36–46°F), upright and

with screw caps tightly closed. If stored as directed, reagents are stable until the expiration date

printed on the label. Do not freeze reagents. Avoid prolonged exposure to temperatures above

32°C (90°F). Improper storage of reagents can affect assay performance.

Reagents  need to be transferred to analyzer-specific reagent containers

prior to use. Avoid cross-contamination of reagents.

Loading Reagents on the Vitros 5600

1  Pour Reagent 1 into the A side of the UD reagent bottle.

2. Pour Reagent 2 into the B side of the UD reagent bottle.

Load the bottle into Load three on the 5600. It will show in inventory as UD-01.

Vitros UD Microtip Container

R2 -  Enzyme                                       R1 - Antibody/Substrate

Calibration:

Perform a full calibration (6-point) procedure using the ARK Methotrexate Calibrators A, B, C,

D, E, and F; test calibrators in duplicate. Calibration is required with each new reagent kit lot

number. Verify the calibration curve with at least two levels of quality controls according to the

established laboratory quality assurance plan..

When to Re-Calibrate

• Whenever a new lot number of reagents is used

• Whenever indicated by quality control results

• Whenever required by standard laboratory protocols

Vitros 5600 Settings

Full Assay name :  METHTREXATE

Short Assay Name : MTX                                                            Fluid Type: Serum

Assay Model Type: 2 Point Rate                                                  Template:  2PT R1-s-R2

Cal Model Type: Logit/Log4                          Calibrator Bottles: 6      Replicates Per Cal: 2

Standard Dilution Factor: 1.0                 Diluent : Saline

Reflex Dilution Information

Reflex Dilution: On                           Dilution Factor 1.0             Reduction Factor: 1.0

Sample Inices Check: Enabled

Threshold Limits

Hemolysis: 1000                               Icterus: 25                               Turbidity: 800

Reporting Type: Quantitative                                               Units: umol/L

Significant Digits: 3                                                              Precision Digits: 2

Slope: 1.00                                                                             Intercept: 0.00

CuveTip Expiration Time: 35                                                Temperature Sensitive: No

Ranges

Reference: 0.00 - 5.00

Supplementary: 0.00 – 900000000

Measuring (Reportable);  0.00 – 1.20

Initial Absorbance Limits:  -0.200 – 2.700              Second Absorbance Limits:  -0.200 – 2.700

Antigen Excess Factor: 9.0000

Monotonicity : Increase

Max Response High: 5.000                        Max Response Low:  -5.000

Min Response High:  5.000                        Min Response Low: -5.000

Cal fit Goodness Limit:  0.990

Measuring (Reportable) Conc.                    Triple Read Limit

Measuring (Reportable) Min:                         0.00                                                 400

Critical Conc.:                                                 1.00                                                 8.00

Measuring (Reportable) Max:                         1.20                                                 8.00

Kit Lot           Bottle Number                 Dilution Factor            Cal. Value            Response Range

7001                     1                                        1.0                           0.00                      0.20000

7001                     1                                        1.0                           0.05                      0.20000

7001                     1                                        1.0                           0.15                      0.20000

7001                     1                                        1.0                           0.25                      0.20000

7001                     1                                        1.0                           0.50                      0.20000

7001                     1                                        1.0                           1.20                      0.20000

Protocol Steps

- Reagent            	Volumes( uL) = 150.0, Pack/Bottle = UD01 / A

- Incubation           	Seconds = 0.00

- Sample               	Volume (uL) = 8.5

- Incubation            	Seconds = 114.0

- Reagent                       Volume (uL) = 75.0, Pack/Bottle= UD01 / B

- Incubation           	Seconds= 76.00

- Read                    	Wavelength= 340 nm

- Incubation            	Seconds= 140.50

- Read                     	Wavelength = 340nm

Quality Control:

Refer to the Chemistry Quality Control Procedure for Specifics.

Good laboratory practice suggests that at least two levels (low and high medical decision

points) of quality control be tested each day patient samples are assayed and each time a

calibration is performed. Monitor the control values for any trends or shifts. If any trends or shifts are detected, or if the control does not recover within the specified range, review all operating parameters according to your clinical laboratory quality procedures.

Procedure:

Materials Provided

ARK Methotrexate Assay –5026-0001-00

ARK Methotrexate Assay, Roche® cobas c pack – 5026-0001-02

Materials Required – Provided Separately

ARK Methotrexate Calibrator – 5026-0002-00

Quality Controls – ARK Methotrexate Control – 5026-0003-00

ARK Methotrexate Dilution Buffer 5026-0004-00

Instruments

Reagents will  need to be transferred to analyzer-specific reagent containers

prior to use. Avoid cross-contamination of R1and R2 . See Illustration under Reagents

Assay Sequence

To run or calibrate the assay, see the Vitros 5600 settings above.

Sample Dilution:

If the sample is above the measuring range (>1.20). Dilute using the Dilution Buffer and program a manual dilution on the 5600.

Results:

Results are reported as umol/L.

Methotrexate serum levels depend on indication for use, dosage, mode of administration,

treatment regimen, individual pharmacokinetics, metabolism and other clinical factors.1,3 While

the serum level may typically reach approximately 10 to 100 μmol/L in treatment of breast

cancer (for example),15 concentrations may exceed 1000 μmol/L16 with high dose therapy for

osteosarcoma, and up to 3100 μmol/L methotrexate was reached following a 4-hour infusion

in pediatric patients with osteosarcoma.17 For treatment of osteosarcoma,16 the methotrexate

decay curve has wide variability: 24 hours, 30 to 300 μmol/L; 48 hours, 3 to 30 μmol/L; and 72

hours, less than 0.3 μmol/L. A dose of 10 mg of leucovorin is usually administered intravenously

24 hours after initiation of the MTX infusion. Subsequent doses are adjusted and administered

according to the MTX levels obtained at 24, 48, and 72 hours. Methotrexate levels in excess

of 50 μmol/L at 24 hours, 10 μmol/L at 48 hours, and 0.5 μmol/L at 72 hours portend potential

toxicity and are usually treated with an increase in the dose of leucovorin in accordance

with algorithms until the MTX level is <0.1 μmol/L. Guidelines for methotrexate therapy with

leucovorin rescue usually recommend continuance of leucovorin until the methotrexate level

falls below 0.05 μmol/L.1, 9 Some centers follow ≤ 0.10 μmol/L.16, 18

From prescribing and other information: Laboratory Indicators of Toxicity Following Leucovorin Rescue Schedules with High Dose Methotrexate.1, 9, 19

METHOTREXATE

<5.00 umol/L at 24 hours

<0.50 umol/L at 48 hours

<0.20 umol/L at 72 hours

For conventional low dose leucovorin rescue.

Renal toxicity is a significant risk and may be exacerbated by coadministration of other

drugs, 14, 19 for example vancomycin.20 Other forms of toxicity can occur, including digestive

disorders (e.g., nausea, vomiting, abdominal pain), cutaneous–mucous disorders (especially

mucositis), haematological abnormalities (e.g., neutropenia and thrombocytopenia), liver

function test disturbances, and neurotoxicity.21-28

Given the profile of the appearance of the 7-hydroxymethotrexate metabolite,15, 27 its molar ratio to methotrexate of up to approximately 100-fold,29 and relative insolubility versus the parent drug,14, 19 possible nephrotoxicity due to precipitation of the metabolite in renal tubules29 may delay elimination of methotrexate itself.                                                Glucarpidase therapy (available for compassionate use) reduces the circulating level of

methotrexate rapidly, not intracellular drug. A rebound effect in the serum level of methotrexate

following glucarpidase therapy has been observed.14 Elimination of DAMPA may take several

days before it no longer interferes with the monitoring of methotrexate by immunoassay.13

Procedure Notes:

LIMITATIONS OF PROCEDURE

This assay is designed for use with serum or plasma only. It is generally good practice to use the same method (as well as matrix) consistently for individual patient care due to the potential for method-tomethod variability. .

As with all analyte determinations, the methotrexate value should be used in conjunction with

information available from clinical evaluation and other diagnostic procedures.

IMPORTANT: Specimens from patients who have received glucarpidase (carboxypeptidase

G2) as a high dose methotrexate rescue therapy should not be tested with the ARK Methotrexate

Assay. These specimens have increased serum levels of 4-[[2,4-diamino-6-(pteridinyl)methyl]-

methylamino]-benzoic acid (DAMPA)10-12 that result from metabolism of methotrexate by

glucarpidase. DAMPA crossreacts with the methotrexate antibody used in this assay, and

may continue to circulate for at least five to seven days before accurate measurements of

serum methotrexate may return.13 Oncologists on the clinical team should notify the laboratory

when glucarpidase is administered to avoid the reporting of falsely elevated methotrexate

concentrations due to interference by DAMPA that would confuse the efforts of glucarpidase

therapy.13 While glucarpidase is well tolerated and rapidly reduces circulating MTX, delayed

renal elimination of MTX can still be a problem for adult and elderly patients.14

Measurement Range

The measurement range of the ARK Methotrexate Assay is 0.04 - 1.20 μmol/L. Specimens

containing methotrexate in higher concentrations are assayed by dilution of the specimen.

Report assayed values exceeding the LoD according to the information provided for LoQ.

Multiply the assayed result by the dilution factor for specimens containing methotrexate above

the measurement range.

Recovery

Accuracy (analytical recovery) was performed by adding concentrated methotrexate

drug into human serum negative for methotrexate. A certified stock concentrate of highly

pure methotrexate was added volumetrically to human serum negative for methotrexate,

representing drug concentrations across the assay calibration range. Six replicates of each

sample were assayed on an automated clinical chemistry analyzer. The results were averaged

and compared to the target concentration and percent recovery calculated. Results are shown

below.

Linearity

Linearity studies were performed as suggested in CLSI/NCCLS Protocol EP6-A. A 1.30μmol/L serum sample was prepared and dilutions were made proportionally with human serum

negative for methotrexate. Linearity at specific dilutions was considered acceptable if the

percent difference was ±10% between the predicted 1st and 2nd order regressed values for

concentrations >0.10 μmol/L or ±0.01 μmol/L at concentrations ≤ 0.10 μmol/L.

Samples containing methotrexate between 2 and 1200 μmol/L were prepared proportionally in

pooled human serum and then diluted into the calibration range with ARK Methotrexate Dilution

Buffer. Regression of assayed methotrexate concentrations was linear throughout the range.

Method Comparison

Correlation studies were performed using CLSI/NCCLS Protocol EP9-A2. Results from

the ARK Methotrexate Assay were compared with results from Fluorescence Polarization

Immunoassay method (monoclonal FPIA)30.

Methotrexate concentrations by FPIA ranged 0.04 to 1440 μmol/L (μM). ARK Methotrexate

values ranged 0.04 to 1500 μmol/L. Results of the Passing-Bablok31 regression analysis for the

study are shown below (with 95% confidence limits) for 102 specimens within the measurement

range as well as for all 147 specimens including those above the measurement range requiring

dilution.

Precision

Precision was determined as described in CLSI/NCCLS Protocol EP5-A2. The six-level ARK

Methotrexate Control and pooled human specimens containing methotrexate were used in the

study. Each level was assayed in quadruplicate twice a day for 20 days. Each of the runs per

day was separated by at least two hours. The within run, between day, total SD, and percent

CVs were calculated. Results are shown below. Acceptance criteria: ≤10% total CV at >0.1

μmol/L, SD ≤0.01 at ≤0.1 μmol/L.

Interfering Substances

Interference studies were conducted using CLSI/NCCLS Protocol EP7-A2 as a guideline.

Clinically high concentrations of the following potentially interfering endogenous substances

in serum with known levels of methotrexate (approximately 0.05 and 0.50 μmol/L) were

evaluated. Each sample was assayed using the ARK Methotrexate Assay, along with a serum

control of methotrexate. Measurement of methotrexate was not substantially affected at the

levels of endogenous substances tested.

Specificity

Methotrexate’s metabolites, folate analogs and other compounds having structural similarity

were tested to determine whether these compounds affect the quantitation of methotrexate

concentrations using the ARK Methotrexate Assay. High levels of these compounds

were spiked into serum pools containing no methotrexate, 0.05 μmol/L or 0.50 μmol/L of

methotrexate. The samples were analyzed and the methotrexate concentrations of samples

containing interferent were compared to a serum control.

Crossreactivity to 7-Hydroxymethotrexate, the major metabolite

After administration of high-dose methotrexate (HDMTX), the serum/plasma concentration of

7-hydroxymethotrexate typically exceeds that of methotrexate at later time points. It has been

reported that 7-hydroxymethotrexate levels exceed those of methotrexate by up to 100-fold 12

to 48 hours after HDMTX administration.15, 27, 29, 32, 34-35

Crossreactivity by 7-hydroxymethotrexate in the measurement of methotrexate was determined

for the ARK Methotrexate Assay by testing paired samples containing (1) both 0.05 μmol/L

methotrexate and 5 μmol/L 7-hydroxymethotrexate and (2) both 0.50 μmol/L methotrexate and

50 μmol/L 7-hydroxymethotrexate in human serum.

The ARK Methotrexate Assay did not crossreact (≤ 0.07%) with the major metabolite

7-hydroxymethotrexate.

Crossreactivity to 2,4-Diamino-N10-methylpteroic acid (DAMPA)

As a minor metabolite of methotrexate, DAMPA is not expected to circulate at concentrations

that would interfere in measurement of methotrexate.33 However, following glucarpidase rescue

therapy, the serum concentration of DAMPA can be substantial.13, 14 The ARK Methotrexate

Assay crossreacts substantially with the minor metabolite DAMPA. Tests were performed

in the absence of the parent drug methotrexate. Crossreactivity to DAMPA ranged 64.3 to

100%. The assay should not be used during possible compassionate therapy with glucarpidase

(carboxypeptidase G2) that rapidly converts circulating methotrexate to DAMPA.

Drugs that crossreact

The ARK Methotrexate Assay crossreacts slightly with triamterene and trimethoprim, however

these drugs may be contraindicated for MTX cancer treatment due to additional adverse effects

if co-administered. The structures of these compounds closely match the pteridine ring moiety

of methotrexate.

Crossreactivity to folate analogs and other compounds

The ARK Methotrexate Assay did not crossreact (≤ 0.01%) with folate analogs or other

compounds at ≥ 1000 μmol/L as tested.

REFERENCES:

1. Prescribing information. 2008. Methotrexate Injection, USP. Hospira, Inc. Lake Forest, IL.

2. Jonsson, O. G. and Kamen, B. A. 1991. Methotrexate and childhood leukemia. Cancer Investigatio n9:53 – 60.

3. Bleyer, W. A. 1978. The clinical pharmacology of methotrexate: New applications of an old drug. Cancer 41:36 – 51.

4. Saeter, G. et al. 1991. Treatment of osteosarcoma of the extremities with the T-10 protocol, with emphasis on the effects of preoperative chemotherapy with single-agent high-dose methotrexate: A scandinavian sarcoma group study. Journal of Clinical Oncology 9:1766 – 1775.

5. Abromowitch, M. et al. 1988. High-dose methotrexate improves clinical outcome in children with acute lymphoblastic leukemia: St. Jude total therapy study X. Medical and Pediatric Oncology 16:297 – 303.

6. Hann, I. M. et al. 1990. ‘MACHO’ chemotherapy for stage IV B cell lymphoma and B cell acute lymphoblastic leukaemia of childhood. British Journal of Haematology 76:359 – 364.

7. Wheeler, C. A. et al. 1991. Cisplatin, continuous infusion,5-fluorouacil, and intermediate dose

methotrexate in the treatment of unresectable non-small cell carcinoma of the lung. Cancer 67:892– 895.

8. Powles, T. J. et al. 1991. A randomized trial comparing combination chemotherapy using mitomycin C, mitozantrone and methotrexate (3M) with vincristine, anthracycline and cyclophosphamide (VAC) in advanced breast cancer. Br J Cancer 64:406 – 410.

9. Leucovorin (Fusilev) Prescribing Information. 2008. Spectrum Pharmaceuticals, Inc. Irvine, CA.

10. Chabner, B. A. et al. 1972. Enzymatic cleavage of methotrexate provides a method for prevention of drug toxicity. Nature 239:395 – 397.

11. Widemann, B. C. et al. 1995. Carboxypeptidase-G2 rescue in a patient with high dose methotrexateinduced nephrotoxicity. Cancer 76:521 – 526.

12. Buchen, S. et al. 2005. Carboxypeptidase G2 rescue in patients with methotrexate intoxication and renal failure. British Journal of Cancer 92:480 – 487.

13. Al-Turkmani, M. R. et al., 2010. Difficulty Measuring Methotrexate in a Patient with High-Dose

Methotrexate–Induced Nephrotoxicity. Clin Chem 56:1792 – 1796.

14. Schwarz, S. et al. 2007. Glucarpidase (Carboxypeptidase G2) intervention in adult and elderly cancer patients with renal dysfunction and delayed methotrexate elimination after high-dose methotrexate therapy. The Oncologist 12:1299 – 1308.

15. Bore, P. et al. 1987. Pharmacokinetics of Methotrexate and 7-Hydroxy-Methotrexate After Methotrexate Infusions. Cancer Drug Delivery 4:177 – 183.

16. Jaffe, N. and Gorlick, R. 2008. High-Dose Methotrexate in Osteosarcoma: Let the Questions Surcease—Time for Final Acceptance. J Clin Oncol 26:4365 – 4366.

17. Colom, H. et al. 2009. Population Pharmacokinetics of High-Dose Methotrexate After Intravenous Administration in Pediatric Patients With Osteosarcoma. Ther Drug Monit 31:76 – 85.

18. Dombrowsky, E. et al. 2011. Evaluating performance of a decision support system to improve methotrexate pharmacotherapy in children and young adults with cancer. Ther Drug Monit 33:99 – 107.

19. Widemann, B. C. and Adamson, P. C. 2006. Understanding and managing methotrexate nephrotoxicity. Oncologist 11:694 – 703.

20. Blum, R. et al. 2002.Significant  impairment of high-dose methotrexate clearance following vancomycin administration in the absence of overt renal impairment. Annals of Oncology 13:327 – 330.

21. Martelli, N. et al. 2011. Methotrexate pharmacokinetics in childhood acute lymphoblastic leukaemia: a prognostic value? J Clin Pharm Ther 36:237 – 245.

22. Mazanec, D. J. and Grisanti, J. M. 1989. Drug-induced osteoporosis. Cleve Clin J of Med 56:297 – 303.

23. Chessells, J. M. et al. 1990. Neurotoxicity in lymphoblastic leukaemia: Comparison of oral and intramuscular methotrexate and two doses of radiation. Archives of Disease in Childhood 65:416 – 422.

24. Allen, J. C. et al. 1980. Leukoencephalopathy following high-dose IV methotrexate chemotherapy with leucovorin rescue. Cancer Treat Rep 64:1261 – 1273.

25. Jacobs, P. et al. 1991. Methotrexate encephalopathy. Eur J Cancer 27:1061 – 1062.

26. Flombaum, C. D. and Meyers, P. A. 1999. High-Dose Leucovorin as Sole Therapy for Methotrexate Toxicity. J Clin Oncol 17:1589 – 1594.

27. Collier, C. P. et al. 1982. Analysis of methotrexate and 7-hydroxymethotrexate by high-performance liquid chromatography and preliminary clinical studies. Ther Drug Monit 4:371 – 380.

28. Widemann, B. C. et al. 2010. Glucarpidase, leucovorin, and thymidine for high-dose methotrexateinduced renal dysfunction: clinical and pharmacologic factors affecting outcome. J Clin Oncol 28:3979– 3986.

29. Erttmann, R. et al. 1985. 7-Hydroxy-Methotrexate and Clinical Toxicity Following High-Dose Methotrexate Therapy. J Cancer Res Clin Oncol 109:86 – 88.

30. Albertioni, F. et al. 1996. Evaluation of clinical assays for measuring high-dose methotrexate in plasma. Clin Chem 42:39 – 44.

31. Bablok, W. et al. 1988. A general regression procedure for method transformation. Application of linear regression procedures for method comparison studies in clinical chemistry. Part III. J. Clin Chem Clin Biochem 26:783 – 790.

32. Jacobs, S. A. et al. 1976. 7-Hydroxymethotrexate as a urinary metabolite in human subjects and rhesus monkeys receiving high dose methotrexate. J Clin Invest 57:534 – 538.

33. Wolfrom, C. et al. 1990. Pharmacokinetic study of methotrexate, folinic acid and their serum metabolites in children treated with high-dose methotrexate and leucovorin rescue. Eur J Clin Pharmacol 39:377– 383.

34. Belz, S. et al. 1994. High-performance liquid chromatographic determination of methotrexate,7-hydroxymethotrexate, 5-methyltetrahydrofolic acid and folinic acid in serum and cerebrospinal fluid. J Chromatogr B Biomed Appl 661:109 – 118.

35. Breithaupt, H. and Kuenzlen, E. 1982. Pharmacokinetics of methotrexate and 7-hydroxy-m

following infusions of high dose methotrexate. Cancer Treat Rep 66:1733 – 1741.

ARKTMis a trademark of ARK Diagnostics, Inc.