TITLE:  Pipette Function and Maintenance Checks

PRINCIPLE:

A variety of equipment is used to support the performance of analytical procedures.  The laboratory is responsible for performance verification and maintenance of all equipment, including pipettes.

PERSONNEL:

Medical Technologists, Lead Medical Technologists

POLICY:

- When pipettes are purchased they come with a calibration/certification document that attests to accuracy of the new device.

- Annually an outside vendor services all pipettes by mechanical cleaning and replacing of  worn or broken parts.

- The outside vendor performs a check of accuracy and precision.

  - Pipette Calibration is performed using a Gravimetric method on a balance calibrated with NIST traceable weights.  These test procedures comply with ISO/IEC 17025/CLSI guidelines.

  - Accuracy:  The closeness of a measured volume to the true volume as specified by the volume setting of the pipette, also called the Mean Error.

  - Precision:  The closeness of agreement among the individual volume samples, also called Standard Deviation or reproducibility.

  - Adjustable pipette performance is calibrated and verified at the lowest volume.

  - Repeater pipettes are confirmed at all settings.

  - Pipettes must meet or exceed ISO/8655-2 specifications, both at the minimal and maximum volume.

  - All failed pipettes are removed from service

.

.

REFERENCES