TITLE: Alere Determine HIV-1/2 Ag/Ab Combo

PRINCIPLE:

Alere Determine™ HIV–1/2 Ag/Ab Combo is an immunochromatographic test for the simultaneous and separate qualitative detection of free HIV-1 p24 antigen and antibodies to HIV-1 and HIV-2. The test device is a laminated strip that consists of a Sample Pad containing monoclonal biotinylated anti-HIV-1 p24 antibody, a Conjugate Pad containing monoclonal anti-HIV-1 p24 antibody-colloidal selenium and HIV-1 and HIV-2 recombinant antigen-colloidal selenium, and a nitrocellulose membrane with an immobilized mixture of recombinant and synthetic peptide HIV-1 and HIV-2 antigens in the Lower Test Area, immobilized streptavidin in the Upper Test Area, and an immobilized mixture of anti-HIV-1 antibody, HIV-1/2 antigens, and HIV-1 p24 recombinant antigen and anti-HIV-1 p24 monoclonal antibody in the Control Area.

A specimen serum sample is applied to the Sample Pad (followed by Chase Buffer for venipuncture or fingerstick whole blood specimens) and migrates by capillary action through the Conjugate Pad and then through the nitrocellulose membrane.

If HIV-1 p24 antigen is present in the specimen, it binds with the monoclonal biotinylated anti-HIV-1 p24 antibody from the Sample Pad and then with monoclonal anti-HIV-1 p24 antibody-colloidal selenium from the Conjugate Pad to form a complex (biotinylated antibody-antigen-colloidal selenium-antibody). This complex migrates through the solid phase by capillary action until it is captured by immobilized streptavidin at the Upper Test Area (labeled “Ag”) where it forms a single pink/red “Ag” line. If HIV-1 p24 antigen is not present in the specimen or is below the limit of detection of the test, no pink/red Ag line is formed. NOTE: The monoclonal biotinylated anti-HIV-1 p24 antibody used in this assay does not cross react with HIV-2 p26 antigen.

If antibodies to HIV-1 and/or HIV-2 are present in the specimen, the antibodies bind to recombinant gp41 (HIV-1) and gp36 (HIV-2) antigen-colloidal selenium conjugates from the Conjugate Pad. The complex migrates through the solid phase by capillary action until it is captured by immobilized HIV-1 and HIV-2 synthetic peptide antigens and recombinant gp41 antigen at the Lower Test Area (labeled “Ab”) and forms a single pink/red “Ab” line. If antibodies to HIV-1 and/or HIV-2 are absent or are below the detection limit of detection of the test, no pink/red Ab line is formed.

To ensure assay validity, a procedural “Control” line containing a mixture of anti-HIV-1 antibody, HIV-1/2 antigens, and HIV-1 p24 recombinant antigen and anti-HIV-1 p24 monoclonal antibody is incorporated in the nitrocellulose membrane. For a test result to be valid there must be a visible pink/red Control line. During the testing procedure the colloidal selenium conjugates released from the Conjugate Pad will be captured by the antibodies and antigens immobilized in the Control Area and form a pink/red Control line for samples that are either positive or negative. NOTE: A pink/red Control line may appear even when a test sample has not been applied to the Test Unit.

CLINICAL SIGNIFICANCE:

Alere Determine™ HIV–1/2 Ag/Ab Combo is an in vitro, visually read, qualitative immunoassay for the simultaneous detection of Human Immunodeficiency Virus Type 1 (HIV-1) p24 antigen (Ag) and antibodies (Ab) to HIV Type1 and Type 2 (HIV-1 and HIV-2) in human serum. It is intended for use as a point-of-care test to aid in the diagnosis of infection with HIV-1 and HIV-2, including an acute HIV-1 infection, and may distinguish acute HIV-1 infection from established HIV-1 infection when the specimen is positive for HIV-1 p24 antigen and negative for anti-HIV-1 and anti-HIV-2 antibodies. The test is suitable for use in multi-test algorithms designed for the statistical validation of rapid HIV test results. When multiple rapid HIV test are available, this test can be used in appropriate multi-test algorithms.

Alere Determine™ HIV-1/2 Ag/Ab Combo is not intended for use in screening blood, plasma, cell, or tissue donors.

SPECIMEN COLLECTION AND HANDLING:

| A. Specimen: | Prior to specimen collection, provide test subjects with the Subject Information Notice. Serum, plasma and whole blood may be used for testing. SST/Yellow top tubes are the preferred specimen for adult testing Serum specimens MUST be from a tube containing no anticoagulant. Purple Microtainer tubes are acceptable for newborn testing |
| --- | --- |
| B. Specimen Transport: | If specimens are to be shipped, they should be packed in compliance with regulations covering the transportation of etiologic agents. Serum specimens should be shipped refrigerated with cold packs or wet ice. |
| C. Specimen Storage: | Serum specimens may be stored at room temperature (15- 30°C) for up to two days before testing. If testing will not be performed within two days of sample collection, serum specimens should be stored at 2- 8°C if the test is to be run within 7 days of collection. If testing is delayed more than 7 days, the specimen should be frozen (-20°C or colder). Mix specimen well by gentle inversion of the tube immediately before testing. - Avoid repeated freeze/thaw cycles. Specimens that have been frozen and thawed more than 3 times cannot be used. - All frozen specimens must be centrifuged at 10,000g for 5 min at room temperature. Carefully remove the 50 μL test sample from the supernatant. If a lipid layer is formed on the surface of the liquid, ensure that the sample is taken from the clear liquid below that layer. |
| D. Handling Precautions: | - Handle the samples, material contacting samples, and kit controls as if capable of transmitting infection. - Wear protective clothing such as laboratory coats, disposable gloves, and eye protection when handling patient samples. - Do not eat, drink, smoke, apply cosmetics, or handle contact lenses in areas where samples and kit reagent materials are handled. Avoid any contact between hands, eyes, or mouth during sample collection and testing |

REAGENTS AND EQUIPMENT:

- Reagents and Materials Provided

- Aluminum ziplock pouch containing Alere Determine™ HIV-1/2 Ag/Ab Combo Cards. Each Card consists of 5 or 10 Test Units which can be separated from each other by tearing along the perforated lines. Each Test Unit has a cover that is to be removed for sample application and visualization of test results.

- Desiccant Package

- Chase Buffer: 1 in the 25 Test Units kit, 2 in the 100 Test Units kit. Containing sodium chloride, disodium hydrogen phosphate, and Nipasept as a preservative.

- Quick Reference Card

- Package Insert

- Subject Information Notices: 25 in the 25 Test Units kit, and 100 in the 100 Test Units kit.

- Customer Letter

- Disposable Workstations: 25 in the 25 Test Units kit, and 100 in the 100 Test Units kit.

- MATERIALS REQUIRED AND AVAILABLE AS AN ACCESSORY TO THE KIT

• Alere Determine™ HIV–1/2 Ag/Ab Combo Controls (Catalog #7D2628). Each package           contains:

-- 	HIV-1 p24 Antigen Control: 1.5 mL, HIV-1 viral lysate in defibrinated pooled normal human plasma; negative for antibodies to HIV-1, HIV-2 and HCV; negative for HBsAg.

-- 	HIV-1 Reactive Control: 1.5 mL, human plasma positive for anti-HIV-1 antibodies, diluted in defibrinated pooled normal human plasma; negative for antibodies to HIV-2 and HCV; negative for HBsAg.

-- 	HIV-2 Reactive Control: 1.5 mL, human plasma positive for anti-HIV-2 antibodies, diluted in defibrinated pooled normal human plasma; negative for antibodies to HIV-1 and HCV; negative for HBsAg and HIV-1 p24.

-- 	Nonreactive Control: 1.5 mL, defibrinated normal human plasma; negative for antibodies to HIV-1, HIV-2, and HCV; negative for HBsAg and HIV-1 p24.

--    40 Disposable Pipettes – for use testing the external controls only. The disposable pipettes are not to be used for testing patient samples.

-- 	Package Insert

- Materials Required, but not Provided

• 	Clock, watch, or other timing device

• 	Precision pipette capable of delivering 50 μL of sample with disposable tips, to be used in lieu of the disposable capillary tubes supplied with the kit.

• 	Disposable gloves

• 	Antiseptic wipes

• 	Biohazard disposal container

- Storage and Stability

Alere Determine™ HIV–1/2 Ag/Ab Combo Test Cards and Chase Buffer must be stored at 2-30°C (36-86°F) until expiration date.

CALIBRATION:

None Required

QUALITY CONTROL:

- Internal Quality Control

To ensure assay validity, a procedural “Control” line containing a mixture of anti-HIV-1 antibody, HIV-1/2 antigens, and HIV-1 p24 recombinant antigen and anti-HIV-1 p24 monoclonal antibody is incorporated in the nitrocellulose membrane. For a test result to be valid there must be a visible pink/red Control line. During the testing procedure the colloidal selenium conjugates released from the Conjugate Pad will be captured by the antibodies and antigens immobilized in the Control Area and form a pink/red Control line for samples that are either positive or negative. NOTE: A pink/red Control line may appear even when a test sample has not been applied to the Test Unit.

- External Quality Control

Alere Determine™ HIV-1/2 Ag/Ab Combo Controls should be tested under the following circumstances:

• Each new untrained operator prior to performing tests on patient specimens

• When opening a new test kit lot

• Whenever a new shipment of test kits is received

• If the temperature of the test storage area falls outside of 2 to 30°C (36 to 86°F)

• If the temperature of the testing area falls outside of 15 to 30°C (59 to 86°F)

• At periodic intervals indicated by the testing facility

Controls should be tested in the same manner as serum samples in the following Test Procedure.

Good Laboratory Practices (GLP) necessitates testing external control material along with the test samples to ensure proper performance of the test kit. Alere Determine™ Combo HIV-1, HIV-2, p24 Reactive and Nonreactive Controls are available separately for use with Alere Determine™ HIV-1/2 Ag/Ab Combo. The HIV Controls are used to verify proper functioning of the test and the operator’s ability to properly perform the test and to interpret the results.

The HIV-1 and HIV-2 Reactive Controls will produce a REACTIVE test result and have been manufactured to produce a visible Test “Ab” line. The HIV-1 p24 Antigen Control will produce a REACTIVE test result and has been manufactured to produce a visible Test “Ag” line. The Nonreactive Control will produce a NONREACTVE Test Result. Run the Controls as per the TEST PROCEDURE for serum/plasma samples (the use of Chase Buffer is not required) and interpret results as described in INTERPRETATION OF TEST RESULTS sections of the Product Insert. It is the responsibility of each facility using Alere Determine™ HIV–1/2 Ag/Ab Combo to establish an adequate quality assurance program to ensure the performance of the device under specific locations and conditions of use.

Please refer to the Alere Determine™ HIV-1/2 Ag/Ab Combo Package Insert for pictorial examples of REACTIVE, NONREACTIVE and INVALID Test Results.

Record these control results in the Laboratory Information System QC program. Refer to the LIS Procedure Manual for complete instructions.

PRECAUTIONS:

- Safety Precautions

- Handle the samples, material contacting samples, and kit controls as if capable of transmitting infection.

- Wear protective clothing such as laboratory coats, disposable gloves, and eye protection when handling patient samples.

- Do not eat, drink, smoke, apply cosmetics, or handle contact lenses in areas where samples and kit reagent materials are handled. Avoid any contact between hands, eyes, or mouth during sample collection and testing.

- Do not touch the Sample Pad.

- Decontaminate and dispose of all specimens, reagents, disposable workstations, and other potentially contaminated materials in a biohazard waste container in accordance with local regulations.  The recommended method of disposal of biohazard waste is autoclaving for a minimum of 1 hour at 121°C. Disposable materials may be incinerated. Liquid wastes may be mixed with appropriate chemical disinfectants. A freshly prepared solution of 10% bleach (0.5% solution of sodium hypochlorite) is recommended. Allow 60 minutes for effective decontamination. NOTE: Do not autoclave solutions that contain bleach. The workstations are for single use only. The used workstation and Test Unit should be regarded as potentially infectious material. They should be disposed of together, without trying to remove the Test Unit from the workstation, in a biohazard waste container as indicated above.

- Clean and disinfect all spills of specimens or reagents using 10% bleach or other appropriate disinfectant. The bleach solution should be made fresh every day.

- For additional information refer to: Centers for Disease Control and Prevention: Updated U.S. Public Health Service Guidelines for the Management of Occupational Exposures to HBV, HCV, and HIV and Recommendation for Post-exposure Prophylaxis (see list of references below).

- Handling Precautions

  - If Desiccant Packet is missing, DO NOT USE.  Discard Test Cards (all Test Units) and use a new Test Card.

  - Do not use any Test Units from Test Cards if the pouch has been perforated.

  - Do not use kit components beyond the expiration date printed on the label. Always check expiration date prior to testing.

  - Adequate lighting is required to read a test result.

PROCEDURE:

Test Procedure

NOTE:  Alere Determine™ HIV-1/2 Ag/Ab Combo Controls should be tested prior to testing patient specimens when a new untrained operator performs testing, a new test kit lot is to be used, a new shipment of test kits is received, if the temperature of the test storage area falls outside of 2 to 30°C (36 to 86°F), if the temperature of the testing area falls outside of 15 to 30°C (59 to 86°F), and at periodic intervals indicated by the testing facility. Controls should be tested in the same manner as serum or plasma samples in the following Test Procedure.

In instances where a mother has had no prenatal care and is refusing HIV testing; her newborn will be drawn and tested for exposure status only.  (A disclaimer will appear on the results report regarding kit is not FDA approved for newborns and children under 12)

Kit Component Preparation

1. 	Open the aluminum pouch containing the Alere Determine™ HIV-1/2 Ag/Ab Combo Cards.

2. 	Remove the desired numbers of test units from the 5 or 10-Test Unit Card by bending and tearing at the perforation.

NOTE: Removal of the test units should start from the right side of the Card to preserve the lot number which appears on the left side of the Card.

3. 	Return the unused test units to the aluminum pouch and close the pouch with the ziplock.

NOTE: Store the unused cards and test units only in the aluminum pouch containing the desiccant package. Carefully close the ziplock, so that the cards are not exposed to ambient humidity during storage.

4. 	Remove the protective foil cover from each Test Unit. Lay the Test Unit flat in the workstation. The test should be initiated within 2 hours after removing the protective foil cover from each Test Unit.  Do NOT touch the Sample Pad with your fingers.

NOTE: Use of the workstation is optional. If the workstation is not used, place the Test Unit on a flat surface.

1. 	Apply 50 μL of sample (precision pipette) by touching the tip of the pipette to the Sample Pad (marked by the arrow symbol).

2. 	Read the test result between 20 and 30 minutes after the addition of the Sample. Do not read Test Results after 30 minutes.

INTERPRETAION OF TEST RESULTS/REPORTING:

ANTIBODY REACTIVE (Two Lines - Control and Ab Line)

A PINK/RED Control line appears in the Control Area AND a PINK/RED Ab line must appear in the Lower Test Area of the Test Unit. The intensity of the Ab and Control lines may vary. Any visible PINK/RED color in both the Control and Lower Test Areas, regardless of intensity, is considered REACTIVE. A Reactive test result means that HIV-1 and/or HIV-2 antibodies have been detected in the specimen. The test result is interpreted as PRELIMINARY POSITIVE for HIV-1 and/or HIV-2 antibodies.

ANTIGEN (HIV-1 p24) REACTIVE (Two Lines - Control and Ag Line)

A PINK/RED Control line appears in the Control Area AND a PINK/RED Ag line must appear in the Upper Test Area of the Test Unit. The intensity of the Ag and Control lines may vary. Any visible PINK/RED color in both the Control and Upper Test Areas, regardless of intensity, is considered REACTIVE. A Reactive test result means that HIV-1 p24 antigen has been detected in the specimen. The test result is interpreted as PRELIMINARY POSITIVE for HIV-1 p24 antigen.

NOTE: A test result that is PRELIMINARY POSITIVE for HIV-1 p24 antigen in the absence of reactivity for HIV-1 or HIV-2 antibodies may indicate an acute HIV-1 infection in the test subject. In this case the acute HIV-1 infection is distinguished from an established HIV-1 infection in which antibodies to HIV-1 are present.

ANTIBODY REACTIVE AND ANTIGEN (HIV-1 p24) REACTIVE (Three Lines - Control, Ab and Ag Lines)

A PINK/RED Control line appears in the Control Area AND a PINK/RED Ab line must appear in the Lower Test Area AND a PINK/RED Ag line appears in the Upper Test Area of the Test Unit. The intensity of the Ab, Ag and Control lines may vary. Any visible PINK/RED color in the Control Area, the Lower Test Area and the Upper Test Area, regardless of intensity, is considered REACTIVE. The test result is interpreted as PRELIMINARY POSITIVE for HIV-1 and/or HIV-2 antibodies and HIV-1 p24 antigen.

NONREACTIVE (One Line – Control Line)

A PINK/RED Control line appears in the Control Area of the Test Unit, and no PINK/RED Ab or Ag line appears in the Lower Test Area and the Upper Test Area of the Test Unit, respectively. A NONREACTIVE test result means that HIV-1 or HIV-2 antibodies and HIV-1 p24 antigen were not detected in the specimen.

INVALID (No Control Line)

If there is no PINK/RED Control line in the Control Area of the Test Unit, even if a PINK/RED line appears in the Lower Test Area or the Upper Test Area of the Test Unit, the result is INVALID and the test should be repeated. If the problem persists, contact Alere™ Technical Support.

RESULTING

The test code is HIVRC and will be resulted in the immunology worklist.  Internal QC will automatically verify.  Make sure you put in results for both the Antigen (HIVAG) and the Antibody (HIVAD) for each test.  Positives will reflex to send outs for confirmation.  Please print label and bring specimen to send outs.

If the testing was performed on a newborn to determine HIV exposure status, a comment will accompany the result.  Comment will say: Result indicate HIV exposure status only.  HIV cannot be diagnosed from this result.

NOTES/LIMITATIONS:

- Alere Determine™ HIV–1/2 Ag/Ab Combo must ONLY be used with venous whole blood, serum or EDTA plasma.  Serum samples should be collected without anticoagulant.

- Alere Determine™ HIV–1/2 Ag/Ab Combo must be used in accordance with the instructions in the Package Insert to obtain accurate results.

- This assay does not detect or has not been validated to detect HIV-2 antigen.

- A Reactive result using Alere Determine™ HIV-1/2 Ag/Ab Combo suggests the presence of HIV-1 p24 antigen and/or antibodies to HIV-1 and/or HIV-2 in the sample. The Reactive result is interpreted as Preliminary Positive for HIV-1 p24 antigen and/or antibodies to HIV-1 and/or HIV-2. Alere Determine™ HIV-1/2 Ag/Ab Combo is intended as aid in the diagnosis of infection with HIV-1/2. AIDS-related conditions are clinical syndromes, and their diagnosis can only be established clinically.

- For a Reactive result, the intensity of the test line does not necessarily correlate with the titer of antigen or antibody in the sample.

- Reactive test results should be confirmed by additional testing using other tests.

- A Nonreactive result does not preclude the possibility of exposure to HIV or infection with HIV.

- A person who has HIV-1 p24 antigen or antibodies to HIV-1 or HIV-2 is presumed to be infected with the virus. However, a person who has participated in an HIV vaccine study may develop antibodies to the vaccine and may or may not be infected with HIV.

- This assay has not been FDA approved for newborn screening, cord blood specimens, or individuals less than 12 years of age.    This test will be used to establish newborn exposure to HIV.   It is not used to diagnose HIV in newborns but will test for maternal antibodies in the baby’s blood.  The baby’s specimen will be referred for HIV 1 and 2 if results of the rapid test are positive or indeterminant.

- Specimens from individuals infected with HIV-1 and/or HIV-2 who are receiving antiretroviral (ART) therapy may produce false negative test results.

- Specimens from individuals with Toxoplasma IgG, human anti-mouse antibodies, rheumatoid factor, elevated triglycerides (above 600 mg/dL), herpes simplex virus infection, hospitalized and cancer patients may give false positive test results.

Performance Characteristics

The clinical performance of Alere DetermineTM HIV-1/2 Ag/Ab Combo was established in a prospective clinical study conducted at 11 clinical trial sites located in the United States from 2010 – 2011. In this study, venous whole blood, serum and plasma specimens were evaluated from individuals either known to be infected with HIV-1 as confirmed by FDA-licensed confirmatory assays and/or FDA-approved NAT assays, at low risk for HIV infection, or at high risk for HIV infection. A subset of individuals also provided capillary fingerstick samples for evaluation.  In addition, a CLIA waiver study was done where the performance of Alere Determine™ HIV-1/2 Ag/Ab Combo was evaluated in a prospective study conducted at 9 geographically diverse sites located in the United States. At each site, Alere Determine™ HIV-1/2 Ag/Ab Combo was conducted by operators who had no laboratory experience and were representative of users at CLIA waived testing sites (intended use). Refer to the Package Insert for detailed performance characteristics.

REFERENCES:

1. 	Julian W Tang, Paul KS Chan (2007) The Basics of HIV Medicine. http://www.info.gov.hk/aids/pdf/g190htm/i_index.htm

2.   	Pilcher C, Eron JJ, Galvin S, Gay S and Cohen MS (2004) Acute HIV revisited: new opportunities for treatment and prevention. The Journal of Clinical Investigations 113(7): 937-945.

3.   	Respess RA, Rayfield MA and Dondero TJ (2001) Laboratory testing and rapid HIV assays: applications for HIV surveillance in hard- to-reach populations. AIDS 15 Supplement 3: S49-S59.

4.   	Constantine N. HIV Viral Antigen Assays (2001). University of Maryland School of Medicine. http://hivinsite.ucsf.edu/InSite?page=kb-02-02-02-02

5.   	Hoffmann C, Rockstroh JK, and Kamps BC, HIV Medicine 2007, Flying Publisher, Paris, France.

6.   	Lyons MS, Lindsell CJ, Hawkins DA, Raab DL, Trott AT and Fichtenbaum CJ (2008) Contributions to early HIV diagnosis among patients linked to care vary by testing venue. BMC Public Health 8:220.

7.   	Centers for Disease Control and Prevention (CDC) Universal Precautions for Prevention of Transmission of Human Immunodeficiency Virus, Hepatitis B Virus, and other Blood borne Pathogens in Health-Care Settings. MMWR 1988; 37(24):377-388.

8.   	Centers for Disease Control and Prevention (CDC). Updated U.S. Public Health Service Guidelines for the Management of Occupational Exposures to HBV, HCV, and HIV Recommendations for Post exposure Prophylaxis. MMWR 2001; 50(RR-11): 1-42.

9.   	National Committee for Clinical Standards Clinical Waste Management: Approved Guideline. NCCLS Document GPS-A. Villanova, PA: NCCLS, 1993; 13(22):1-18, 29-42.

10. 	US Environmental Protection Agency EPA Guide for Infections Waste Management: Publication No. EPA /530-SW-86-014. Washington, DC: US Environmental Protection Agency, 1986:1-1-5, R1-R3, A1-A24.

11. 	Clinical and Laboratory Standards Institute. Procedures and Devices for the Collection of Diagnostic Capillary Blood Specimens; Approved Standard-5th H4-A5 Vol.24 No.21.

12. 	Centers for Disease Control and Prevention (CDC). Revised Case Definitions for HIV Infection Among Adults, Adolescents, and Children Aged <18 Months and for HIV Infection Among Children Aged 18 Months to <13 Years –United States, 2008 MMWR 2008; 57(RR-10): 1-8 http://www.cdc.gov/osels/ph_surveillance/nndss/casedef/aids2008.htm