TITLE: Xpert® MTB/RIF Assay

Principle:

The Xpert® MTB/RIF Assay, performed on the GeneXpert® Instrument Systems, is a qualitative, nested real-time polymerase

chain reaction (PCR) in vitro diagnostic test for the detection of Mycobacterium tuberculosis complex DNA in raw sputum or

concentrated sputum sediment prepared from induced or expectorated sputum. In specimens where Mycobacterium tuberculosis

complex (MTB-complex) is detected, the Xpert MTB/RIF Assay also detects the rifampin-resistance associated mutations of the

rpoB gene.

The Xpert MTB/RIF Assay is intended for use with specimens from patients for whom there is clinical suspicion of tuberculosis

(TB) and who have received no antituberculosis therapy, or less than three days of therapy. This test is intended as an aid in the

diagnosis of pulmonary tuberculosis when used in conjunction with clinical and other laboratory findings.

An Xpert MTB/RIF Assay result of “MTB NOT DETECTED” from either one or two sputum specimens is highly predictive of

the absence of M. tuberculosis complex bacilli on serial fluorescent acid-fast sputum smears from patients with suspected active

pulmonary tuberculosis and can be used as an aid in the decision of whether continued airborne infection isolation (AII) is

warranted in patients with suspected pulmonary tuberculosis. The determination of whether testing of either one or two sputum

specimens is appropriate for decisions regarding removal from AII should be based on specific clinical circumstances and

institutional guidelines. Clinical decisions regarding the need for continued AII should always occur in conjunction with other

clinical and laboratory evaluations and Xpert MTB/RIF Assay results should not be the sole basis for infection control practices.

The Xpert MTB/RIF Assay must always be used in conjunction with mycobacterial culture to address the risk of false negative

results and to recover organisms when MTB-complex is present for further characterization and drug susceptibility testing.

However, decisions regarding the removal of patients from AII need not wait for culture results. Sputum specimens for TB

culture, AFB smear microscopy, and Xpert MTB/RIF Assay testing should follow CDC recommendations with regard to

collection methods and time frame between specimen collection.

The Xpert MTB/RIF Assay does not provide confirmation of rifampin susceptibility since mechanisms of rifampin resistance

other than those detected by this device may exist that may be associated with a lack of clinical response to treatment.

Specimens that have both MTB-complex DNA and rifampin-resistance associated mutations of the rpoB gene detected by the

Xpert MTB/RIF Assay must have results confirmed by a reference laboratory. If the presence of rifampin-resistance associated

mutations of the rpoB gene is confirmed, specimens should also be tested for the presence of genetic mutations associated with

resistance to other drugs.

The Xpert MTB/RIF Assay should only be performed in laboratories that follow safety practices in accordance with the

CDC/NIH Biosafety in Microbiological and Biomedical Laboratories publication as a BSL Level 2 Laboratory.

Globally, about 2 billion people are infected with MTB.1 In 2010, 8.8 million people developed active disease and 1.4 million

people lost their lives to the illness.2 There were 9,951 new cases of tuberculosis reported in the United States in 2012 (a rate of

3.2 cases per 100,000). In 2011, 536 U.S. deaths were attributed to tuberculosis infections.3,4

Standard treatment regimens for tuberculosis involve prolonged administration of multiple drugs and are usually highly

effective. However, MTB-complex strains that are resistant to one or more of first line drugs require individualized treatment.

Rifampin resistance is often an indication of multidrug resistance to tuberculosis (MDR TB), which is defined as resistance to at

least rifampin (RIF) and isoniazid (INH). In the United States overall RIF resistance is approximately 1.8%, with approximately

90% of these strains resistant to at least RIF and INH.5

Active pulmonary TB is a highly infectious airborne disease. All patients in healthcare facilities with suspected TB should be

maintained in airborne infection isolation (AII) according to recommended infection control guidelines.6 The testing of either

one or two sputum specimens by the Xpert MTB/RIF Assay may serve as an alternative to serial acid-fast stained sputum

smears as an aid in the decision of whether continued infection control precautions are warranted in patients with suspected

pulmonary tuberculosis.

Sputum specimens that are AFB smear-negative but subsequently TB culture positive have lower MTB-complex organism loads

than specimens that are AFB smear-positive. Because of the greater sensitivity of the Xpert MTB/RIF Assay for the detection of

MTB-complex than that of acid-fast microscopy, MTB-complex may be detected by the Xpert MTB/RIF Assay in AFB smear negative samples.

Patients with HIV infection and pulmonary TB are known to have lower organism loads of MTB-complex in their sputum

specimens relative to HIV-uninfected patients, despite more rapid disease progression if untreated. As a consequence, sputum

specimens from HIV-infected patients with pulmonary TB tend to be AFB smear-negative more frequently than HIV-uninfected

patients. Overall rates of detection of MTB-complex with the Xpert MTB/RIF Assay may be lower in settings with a high

percentage of HIV-infected patients because these patients are more likely to produce AFB smear-negative specimens with low

organism loads.

The Xpert MTB/RIF Assay is an automated in vitro diagnostic test using nested real-time PCR for the qualitative detection of

MTB-complex and RIF resistance. The primers in this test amplify a portion of the rpoB gene containing the 81 base pair core

region. The probes are designed to differentiate between the conserved wild-type sequence and mutations in the core region that

are associated with RIF resistance. This assay can be performed on Cepheid GeneXpert® Instrument Systems.

GeneXpert Instrument Systems automate and integrate sample purification, nucleic acid amplification, and detection of the

target sequence using real- time reverse transcriptase PCR (RT-PCR) and real-time PCR assays. The systems consist of an

instrument, personal computer, and preloaded software for running tests and viewing the results. The systems require the use of

single-use disposable GeneXpert cartridges that hold the RT-PCR and PCR reagents and host the RT-PCR and PCR processes.

Because the cartridges are self-contained, cross-contamination between samples is minimized. For a full description of the

systems, refer to the appropriate GeneXpert Dx System Operator Manual or GeneXpert Infinity System Operator Manual.

The Xpert MTB/RIF Assay includes reagents for the detection of MTB-complex and RIF resistance from raw sputum samples

and in concentrated sputum sediments. A Sample Processing Control (SPC) and a Probe Check Control (PCC) are also included

in the cartridge. The SPC is present to control for adequate processing of the target bacteria and to monitor the presence of

inhibitors in the PCR reaction. The Probe Check Control (PCC) verifies reagent rehydration, PCR tube filling in the cartridge,

probe integrity, and dye stability.

The Xpert MTB/RIF Assay simultaneously detects MTB-complex and RIF resistance by amplifying a MTB-complex specific

sequence of the rpoB gene, which is probed with five molecular beacons (Probes A – E) for mutations within the rifampinresistance determining region (RRDR). Each molecular beacon is labeled with a different fluorophore. The valid maximum cycle threshold (Ct) of 39.0 for Probes A, B and C and 36.0 for Probes D and E are set for MTB/RIF data analysis.

• MTB DETECTED is reported when at least two probes result in Ct values within the valid range and a delta Ct min (the

smallest Ct difference between any pair of probes) of less than 2.0.

• Rif Resistance NOT DETECTED is reported if the delta Ct max (the Ct difference between the earliest and latest probe) is

≤4.0.

• Rif Resistance DETECTED is reported if the delta Ct max is >4.0.

• Rif Resistance INDETERMINATE is reported when the following two conditions are met:

1. the Ct value of any probe exceeds the valid maximum Ct (or is zero, i.e. no threshold crossing); and

2. the earliest rpoB Ct value is greater than:

[(Valid maximum Ct of probe in condition 1) - (delta Ct max cut-off of 4.0)].

• MTB NOT DETECTED is reported when there is only one or no positive probe.

All assay settings are included as automatic calculations in the Xpert MTB/RIF protocol and cannot be modified by the user.

Reagents and Instruments

Materials Provided

The Xpert MTB/RIF Assay kit contains sufficient reagents to process 10 specimens or quality control samples. The kit contains the following:

Xpert MTB/RIF Assay Cartridges with Integrated Reaction Tubes 10

• Bead 1 (freeze-dried)			 2 of each per cartridge

• Polymerase

• dNTPs (deoxynucleoside triphosphates)

• Probe

• BSA (Bovine serum albumin)

• Bead 2 (freeze-dried)			 2 of each per cartridge

• Primers

• Probes

• BSA (Bovine serum albumin)

• Bead 3 (freeze-dried) 1 per cartridge

• Sample Processing Control (SPC) ~6,000 non-infectious

B. globigii spores

• Reagent 1				 4 mL per cartridge

• Tris Buffer

• Surfactants

• EDTA (ethylenediaminetetraacetic acid)

• Reagent 2 				4 mL per cartridge

• Tris Buffer

• Surfactants

• EDTA (ethylenediaminetetraacetic acid)

Sample Reagent 			8 mL per bottle

• Sodium Hydroxide

• Isopropanol

Disposable Transfer Pipettes 		12

CD 					1 per kit

• Assay Definition File (ADF) – ADF for use with both the GeneXpert Dx and Infinity Systems

• Instructions to import ADF into GX software

• Package Insert

Note Sample Reagent (SR) can be colorless to yellow to amber. Color may intensify with time, but color has no effect on performance.

Note Safety Data Sheets (SDS) are available at www.cepheid.com or www.cepheidinternational.com under the SUPPORT tab.

Note The bovine serum albumin (BSA) in the beads within this product was produced and manufactured exclusively from bovine plasma sourced in the United States. No ruminant protein or other animal protein was fed to the animals; the animals passed ante- and postmortem testing. During processing, there was no commingling of the material with other animal materials.

Note The transfer pipettes have a single mark representing the minimum volume of sample necessary to transfer the GeneXpert cartridge. Use only for this purpose. All other pipettes must be provided by the laboratory.

Storage and Handling

• Store the Xpert MTB/RIF Assay cartridges and reagents at 2 – 28 °C.

• Do not use reagents or cartridges that have passed the expiration date.

• The cartridge is stable up to 6 weeks at 2-45°C after opening the pouch. Do not open a cartridge until ready to test.

• If using a GeneXpert Dx instrument, start the test within four hours of adding the Sample Reagent-treated sample to

the cartridge.

• If using a GeneXpert Infinity system, be sure to start the test and put the cartridge on the conveyor within 30 minutes

of adding the Sample Reagent-treated sample to the cartridge. Remaining shelf-life is tracked by the system by the

Xpertise Software so that tests are run prior to the four hour on-board expiration.

• Do not use any reagents that have become cloudy or discolored.

Materials Required but Not Provided

• GeneXpert Instrument System (catalog number varies by configuration): 6-color GeneXpert instrument, computer

with proprietary GeneXpert Software Version 4.3 or higher, barcode scanner, and operator manual.

• Printer: If a printer is required, contact Cepheid Technical support to arrange for the purchase of a recommended

printer.

• Leak-proof, sterile, screw-capped specimen collection containers

• Disposable gloves, eye protection, laboratory coats, and labels or permanent marking pen

• Sterile, single use pipettes with dry-end barrier plugs for sample processing

• Timer

Materials Available but Not Provided

MMQCI (Maine Molecular Quality Controls, Inc.) INTROL™ External Run Control (catalog # TBNEG-04) as negative control,

and MMQCI INTROL™ External Run Controls (catalog # TBWT-04 and catalog #TBMDR1-04) as RIF susceptible and RIF

resistant positive controls.

Warnings and Precautions

• Treat all biological specimens, including used cartridges, as if capable of transmitting infectious agents. Because it is often

impossible to know which specimens might be infectious, all biological specimens should be treated with standard

precautions. Guidelines for specimen handling are available from the U.S. Centers for Disease Control and Prevention7 and

the Clinical and Laboratory Standards Institute (formerly National Committee for Clinical Laboratory Standards).8

• Wear protective disposable gloves, laboratory coats, and eye protection when handling specimens and reagents. Wash hands

thoroughly after handling specimens and test reagents.

• Follow your institution’s safety procedures for working with chemicals and handling biological samples.

• Preparation of digested, decontaminated and concentrated sputum sediments, and Xpert MTB/RIF procedures should be

done using Biosafety Level 2 practices.9

• Use only for the detection of members of the M. tuberculosis complex using sediments prepared following the NALC-NaOH

or NaOH procedures recommended by the Centers for Disease Control and Prevention (CDC).10 This test may only be used

with raw sputum samples or concentrated sediments prepared from induced or expectorated sputa.

• When processing more than one sample at a time, open only one cartridge; add the Sample Reagent-treated sample and close

the cartridge before processing the next sample. Change gloves between samples.

• Do not substitute Xpert MTB/RIF Assay reagents with other reagents.

• Do not open the Xpert MTB/RIF Assay cartridge lid except when adding the Sample Reagent-treated sample.

• Do not use a cartridge if it appears wet or if the lid seal appears to have been broken.

• Do not use a cartridge that has been dropped or shaken.

• Do not use a cartridge that has a damaged reaction tube.

• Each single-use Xpert MTB/RIF Assay cartridge is used to process one test. Do not reuse spent cartridges.

• Consult your institution’s environmental waste personnel regarding proper disposal of used cartridges and unused reagents.

This material may contain federal EPA Resource Conservation and Recovery Act (RCRA) hazardous waste requiring

specific disposal. Check state and local regulations as they may differ from federal disposal regulations. Institutions should

check their country hazardous waste disposal requirements.

• Sample Reagent contains sodium hydroxide (pH > 12.5); (international chemical safety hazard codes H303, H315, H318),

which may be irritating to eyes and skin and requires eye and skin protection. Sample reagent also contains isopropanol

(international chemical safety hazard code H225), which is a flammable liquid.

• The HIV/MTB co-infected population may contain an increased percentage of patients with smear-negative samples with

MTB-complex levels below the assay’s level of detection.

• Local, state, and federal regulations for notification of reportable diseases are continually updated and include a number of

organisms for surveillance and outbreak investigations.11,12 Additionally the Centers for Disease Control and Prevention

(CDC) recommends that when pathogens from reportable diseases are detected by a culture independent diagnostic test

(CIDT), the laboratory should facilitate obtaining the isolate or clinical materials for submission to the appropriate public

health laboratory to aid in outbreak and epidemiological investigations. Laboratories are responsible for following their state

and local regulations and should consult their state and local public health laboratories for isolate and clinical sample

submission guidelines.

Specimen Collection, Transport and Storage

Collection

Collect raw or induced sputum following standard procedures. At least 1mL is required for testing. Additionally, 5mL of sputum is required for AFB Culture/Smear which must be performed on all specimens tested.

Storage and Transport

Raw sputum: Transport and store specimens at 2–8 °C before processing whenever possible. If necessary, sputum specimens can be stored at a maximum of 35 °C for up to three days and then at 2–8 °C for an additional seven days.

Assay Procedure

Table 1. Required Specimen Volume

Specimen Type        Minimum Volume for One Test      Minimum Total Volume for Test and Retest – See Section 11.2, Retest Procedure

Raw sputum                               1 mL 					 2 mL

Note Reject samples with obvious food particles or other solid particulate.

Note Write on the sides of the cartridge or affix an ID label. Do not put the label on the lid of the cartridge or cover the existing barcode on

the cartridge.

Note Store resuspended sediments at 2–8 °C for up to seven days.

Raw Sputum Specimens

1. Wear protective disposable gloves.

2. Label each Xpert MTB/RIF Assay cartridge with the sample ID.

3. Carefully open the lid of the sputum collection container and transfer 1-4mL of sputum to a 15mL conical tube.

**Save an adequate amount of sputum to send out the remaining sample for AFB Culture and Smear. Add Culture AFB/Smear if not already ordered.

4. Pour or pipette (pipette not provided) approximately 2 times the volume of the Sample Reagent into the sputum (2:1 dilution, Sample Reagent: sputum). Be careful not to allow the Sample Reagent to touch bare skin.

Example 1

8 ml Sample Reagent;

4 mL sputum

Example 2

2 ml Sample Reagent;

1 mL sputum

Note: Discard the leftover Sample Reagent and the bottle in a hard plastic biohazard  waste container.

5. Replace and secure the lid.

6. Shake vigorously 10 to 20 times or vortex for at least 10 seconds.

7. Incubate sample for a total of 15 minutes at 20–30 °C.

8. Between 5 and 10 minutes into the incubation period, shake vigorously 10 to 20 times or vortex for at least 10 seconds.

Preparing the Cartridge

To add the sample and reagents into the cartridge:

1. Open the cartridge lid, and then open the sample container.

2. Using the provided transfer pipette, aspirate the liquefied sample close to the line on the pipette.  If there is insufficient sample volume, do not proceed with testing.

Important

If using a GeneXpert Dx instrument, start the test within four hours of adding the Sample Reagent-treated

sample to the cartridge.

Note When processing more than one sample at a time, open only one cartridge, add the Sample Reagent-treated sample and close the cartridge before moving to the next sample. Change gloves between samples.

To add the sample and reagents into the cartridge:

1. Open the cartridge lid, and then open the sample container.

2. Using the provided transfer pipette, aspirate the liquefied sample close to the line on the pipette.  If there is

insufficient sample volume, do not proceed with testing.

Dispensing the sample slowly to minimize the risk of aerosol formation, transfer the Sample Reagent-treated sample into the sample chamber of the Xpert MTB/RIF cartridge. See Figure 5.

Figure 5. Dispensing Decontaminated Liquefied Sample into the Sample Chamber of the Cartridge

4. Close the cartridge lid firmly. Remaining Sample Reagent-treated sample can be kept for up to four hours at 2–8 ºC in case retesting is required.

Starting the Test

This section lists the basic steps for running the test. For detailed instructions, see the GeneXpert Dx System Operator Manual or the GeneXpert Infinity System Operator Manual.

1. Turn on the GeneXpert Instrument System.

• If using the GeneXpert Dx instrument, first turn on the GeneXpert Dx instrument and then turn on the computer. The

GeneXpert software will launch automatically or may require double-clicking on the GeneXpert Dx software shortcut

icon on the Windows® desktop or

2. Log on to the GeneXpert Instrument System software using your user name and password.

3. In the GeneXpert System window, click Create Test (GeneXpert Dx)

4. Scan or type in the Patient Order #. If typing the Patient Order #, make sure the Patient ID is typed correctly. The Patient ID is associated with the test results and is shown in the View Results window.

5. Scan the barcode on the Xpert MTB/RIF Assay cartridge. The Create Test window appears. Using the barcode information, the software automatically fills the boxes for the following fields: Select Assay, Reagent Lot ID, Cartridge SN, and Expiration Date.

6.Click the Notes box and type in the Patient Name.

7. Click Start Test (GeneXpert Dx)

For the GeneXpert Dx Instrument:

A. Open the instrument module door with the blinking green light and load the cartridge.

B. Close the door. The test starts and the blinking green light changes to a solid green light. When the test is finished, the light turns off.

C. Wait until the system releases the door lock before opening the module door and removing the cartridge.

D. The used cartridges should be disposed of in the appropriate biohazard specimen waste containers.

Viewing and Printing Results

This section lists the basic steps for viewing and printing results. For more detailed instructions on how to view and print the results, see the GeneXpert Dx System Operator Manual.

1. Click the View Results icon to view results.

2. Upon completion of the test in approximately two hours, click the Report button of the View Results screen to view and/or generate a pdf report file.

Quality Control

Each test includes a Sample Processing Control (SPC) and Probe Check Control (PCC).

Sample Processing Control (SPC) – Ensures the sample was correctly processed. The SPC contains non-infectious spores in the form of a dry spore cake that is included in each cartridge to verify adequate processing of MTB. The SPC verifies that conditions for lysis of MTB have occurred if the organisms are present and verifies that specimen processing is adequate. Additionally, this control detects specimen-associated inhibition of the real-time PCR reactions and acts as an internal positivecontrol.

The SPC should be positive in a MTB-negative sample and can be negative or positive in a positive sample. The SPC passes if it meets the validated acceptance criteria. The test result will be Invalid if the SPC is not detected in an MTB-negative sample.

Probe Check Control (PCC, QC1, QC2) — Before the start of the first and second reactions of the nested PCR assay, the GeneXpert Instrument System measures the fluorescence signal from the QC1 and QC2 probes (reaction 1) and the rpoB and SPC probes (reaction 2) to monitor bead rehydration, reaction-tube filling, probe integrity and dye stability. The PCC passes if it meets the assigned acceptance criteria.

External Controls

MMQCI (Maine Molecular Quality Controls, Inc.) INTROL™ External Run Controls are performed with each lot lot/shipment as well as 30 day QC.

- (catalog # TBNEG) as negative control  MTB-NEGATIVE,  RIFAMPIN RESISTANCE-NEGATIVE

- (catalog # TBWT-04) as positive ctrl 1   MTB-POSITIVE,  RIFAMPIN RESISTANCE-NEGATIVE

- (catalog#TBMDR1-04) as positive ctrl 2 MTB-POSITIVE,  RIFAMPIN RESISTANCE-POSITIVE

For each external control add 200uL of control to 2mL of sample reagent in a 15mL conical tube and proceed from step 5 of Assay Procedure for raw sputum specimens.

Interpretation of Results

The GeneXpert Instrument System generates the results from measured fluorescent signals and embedded calculation algorithms. The results can be seen in the View Results window. See Figure 6, Figure 7, Figure 8, and Figure 9 for specific examples, and see Table 2 for a list of all possible results.

Result Interpretation

MTB DETECTED;

Rif Resistance DETECTED

(Figure 6)

The MTB target is detected within the sample:

• A mutation in the rpoB gene has been detected.

• SPC: NA (not applicable). An SPC signal is not required because MTB amplification can compete with this

control.

• Probe Check (QC1 and QC2): PASS. All probe check results pass.

MTB DETECTED;

Rif Resistance NOT DETECTED

(Figure 7)

The MTB target is detected within the sample:

• A mutation in the rpoB gene has not been detected.

• SPC: NA (not applicable). An SPC signal is not required because MTB amplification can compete with this control.

• Probe Check (QC1 and QC2): PASS. All probe check

results pass.

MTB DETECTED;

Rif Resistance INDETERMINATE

(Figure 8)

The MTB target is detected within the sample:

• A mutation in the rpoB gene could not be determined due to insufficient signal detection.

• SPC: NA (not applicable). An SPC signal is not required because MTB amplification can compete with this control.

• Probe Check (QC1 and QC2): PASS. All probe checkresults pass.

MTB NOT DETECTED

(Figure 9)

The MTB target is not detected within the sample.

• SPC: PASS. The SPC met the acceptance criteria.

• Probe Check (QC1 and QC2): PASS. All probe check results pass.

INVALID (Figure 10) The presence or absence of MTB cannot be determined.

The SPC does not meet the acceptance criteria, the sample was not properly processed, or PCR was inhibited. Repeat the test. See Section 11.2, Retest Procedure.

• MTB INVALID: The presence or absence of MTB DNA cannot be determined.

• SPC: FAIL. The MTB target result is negative, and the SPC Ct is not within valid range.

• Probe Check (QC1 and QC2): PASS. All probe check results pass.

ERROR The presence or absence of MTB cannot be determined.

Repeat the test. See Section 11.2, Retest Procedure.

• MTB: NO RESULT

• SPC: NO RESULT

• Probe Check (QC1 and QC2): PASS/FAIL.

Probe check failure can be the source of error but other errors, such as system component failure, can occur

even if probe check passes.

NO RESULT The presence or absence of MTB cannot be determined.

Repeat the test. See Section 11.2, Retest Procedure. A

NO RESULT indicates that insufficient data were collected. For example, the operator stopped a test that was in

progress.

• MTB: NO RESULT

• SPC: NO RESULT

• Probe Check (QC1 and QC2): NA (not applicable).

Results will be manually entered into Soft using the result keypad.  Use the Genexpert resulting worklist for result entry/finalization. See Soft LIS Micro Procedure Manual for further details.

The results are interpolated by GeneXpert® Dx System from measured fluorescent signals and embedded calculation algorithms and will be shown in the View Results window.  Possible results are:

Figure 6. An Example of a MTB DETECTED; Rif Resistance DETECTED Result

Choose following results from the keypad:

- }PMTB ?Pos MTB – “Mycobacterium tuberculosis complex DNA DETECTED”

- }PRIF ?Pos Rif Resistance – “Rifampin Resistance DETECTED”

?Internal QC acceptable.

Document critical call. Caregiver and Infection Prevention must be immediately notified. Order AFB Culture/Smear on remaining sample.

Figure 7. An Example of a MTB DETECTED; Rif Resistance NOT DETECTED Result

Choose following results from the keypad:

1.   }PMTB ?Pos MTB – “Mycobacterium tuberculosis complex DNA DETECTED”

2.   }NRIF ?Pos Rif Resistance – “Rifampin Resistance NOT DETECTED”

?Internal QC acceptable.

Document critical call. Caregiver and Infection Prevention must be immediately notified. Order AFB Culture/Smear on remaining sample.

Figure 8. An Example of a MTB DETECTED; Rif Resistance INDETERMINATE Result

Choose following results from the keypad:

1.   }PMTB ?Pos MTB – “Mycobacterium tuberculosis complex DNA DETECTED”

2.   }IRIF ?Indeterm Rif Resist – “Rifampin Resistance INDETERMINATE”

?Internal QC acceptable.

Document critical call. Caregiver and Infection Prevention must be immediately notified. Order AFB Culture/Smear on remaining sample.

Figure 9. An Example of a MTB NOT DETECTED Result

Choose following result from the keypad:

- }NMTB ?Neg MTB – “Mycobacterium tuberculosis complex DNA NOT DETECTED”

?Internal QC acceptable.

Order AFB Culture/Smear on remaining sample.

Figure 10. An Example of an INVALID Result

If the test has been repeated with the same result:

Choose following result from the keypad:

- }QCFR ?QC fail. Recollect – “Internal QC failure. Recollection required.”

Document call to caregiver.

Order AFB Culture/Smear on remaining sample.

Reasons to Repeat the Test

Repeat the test using a new cartridge if one of the following test results occurs:

• An INVALID result indicates that the SPC failed. The sample was not properly processed, or PCR was inhibited.

• An ERROR result indicates that the PCC (QC1 or QC2) failed or a system failure occurred, and the assay was aborted. The cause of the errors are possibly due to the reaction tube being filled improperly, a reagent probe integrity problem was detected, the maximum pressure limits were exceeded, or a GeneXpert module failed.

• A NO RESULT indicates that insufficient data were collected. For example, the operator stopped a test that was in progress.

Retest Procedure

If you have leftover raw sputum always use new Sample Reagent to treat the sputum before running the assay. If you have sufficient leftover Sample Reagent-treated sample and are within four hours of sample preparation, you can use the leftover sample to prepare and process a new cartridge immediately on the GeneXpert instrument. When retesting, always use a new cartridge. See Preparing the Cartridge.

Limitations

1. The performance of the Xpert MTB/RIF Assay was evaluated using induced or expectorated sputa. Testing of other clinical samples (e.g., blood, CSF, gastric aspirate, stool, tissue, urine) has not been evaluated and may alter test performance.

2. Concentrated sputum sediments used during the performance evaluation of the Xpert MTB/RIF Assay were prepared following the NALC-NaOH or NaOH procedures recommended by the Centers for Disease Control and Prevention (CDC).8 Use of other methods of sediment preparation may alter the performance of the test.

3. The Xpert MTB/RIF Assay is not indicated for use with sputum samples from patients being treated with antituberculosis drugs either to determine bacteriologic cure or to monitor response to therapy.

4. A negative test does not exclude the possibility of isolating MTB-complex from the sputum sample. The Xpert MTB/RIF Assay must be used in conjunction with mycobacterial culture to address the risk of false negative results and to recover the organism for further characterization and susceptibility testing.

5. A positive test does not necessarily indicate the presence of viable organisms.

6. The Xpert MTB/RIF Assay does not differentiate between the species of the MTB-complex (i.e., M. tuberculosis, M. bovis, M. africanum, M. canettii, M. microti, M. caprae, M. pinnipedi, M. mungi, and M. orygis). In addition, culture must also be performed to determine if mycobacteria other than tuberculosis complex (MOTT) is present in addition to MTB-complex.

7. Due to the low prevalence of rifampin resistant TB in the United States and the implications of rifampin resistance for treatment, all MTB-complex strains determined to be rifampin resistant by the Xpert MTB/RIF Assay must have the presence of rifampin resistance associated mutations of the rpoB gene confirmed by a reference laboratory. Additional testing for the presence of mutations associated with resistance to other drugs for the treatment of TB should also be performed.

8. Because the detection of MTB-complex is dependent on the number of organisms present in the sample, accurate results are dependent on proper specimen collection, handling, and storage. Erroneous test results might occur from improper specimen collection, failure to follow the recommended sample collection procedure, handling or storage problems, technical error, sample mix-up, or an insufficient concentration of starting material. Careful compliance to the instructions in this insert is necessary to avoid erroneous results.

9. The performance of the Xpert MTB/RIF Assay has not been evaluated with samples from pediatric patients.

10. The performance of the Xpert MTB/RIF test is dependent on operator proficiency and adherence to assay procedures. Assay procedural errors may cause false positive or false negative results. All device operators should have appropriate device training.

11. A trained health care professional should interpret assay results in conjunction with the patient’s medical history, clinical signs and symptoms, and the results of other diagnostic tests.

12. Assay interference may be observed in the presence of Lidocaine (>20% v/v), mucin (>1.5% w/v), Ethambutol (>5 μg/mL), Guaifenesin (>2.5 mg/mL), Phenylephrine (>25% v/v), or tea tree oil (>0.008% v/v).13. Studies demonstrated that M. scrofulaceum when tested at a concentration of 108 CFU/mL produced a false positive Xpert MTB/RIF Assay result.

References

1. WHO report 2008. Tuberculosis (TB). http://www.who.int/tb/publications/global_report/2008/en/index.html

2. WHO report 2011. http://www.who.int/tb/publications/global_report/en/.

3. Center for Disease Control and Prevention. Trends in Tuberculosis, 2012. http://www.cdc.gov/tb/publications/factsheets/ statistics/Trends.pdf.

4. Hoyert DL, Xu J. Deaths: Preliminary data for 2011. National Vital Statistics Reports. 2012;61(6);p. 16.

5. Reported Tuberculosis in the United States, 2011. Atlanta, GA: U.S. Department of Health and Human Services, CDC, October 2012.

6. Centers for Disease Control and Prevention. Guidelines for preventing the transmission of Mycobacterium tuberculosis in health-care settings, 2005. MMWR 2005;54(No. RR-17).

7. Centers for Disease Control and Prevention. Biosafety in Microbiological and Biomedical Laboratories. Richmond JY and McKinney RW (eds) (1993). HHS Publication number (CDC) 93-8395.

8. Clinical and Laboratory Standards Institute. Protection of Laboratory Workers from Occupationally Acquired Infections; Approved Guideline. Document M29 (refer to latest edition).

9. U.S. Department of Health and Human Services, Public Health Service Centers for Disease Control and Prevention, National Institutes of Health. 2009. Biosafety in Microbiological and Biomedical Laboratories, 5th Edition HHS Publication No. (CDC) 21-1112. Available from: http://www.cdc.gov/biosafety/publications/bmbl5/BMBL.pdf. Accessed December 14, 2012.

10. Kent, PT, Kubica, GP, 1985. Public health mycobacteriology: a guide for the level III laboratory. U.S. Department of Public Health and Human Services, Atlanta, GA.

11. Centers for Disease Control and Prevention. Summary of Notifiable Diseases. MMWR.

http://www.cdc.gov/mmwr/PDF/wk/mm6153.pdf

12. Council to Improve Foodborne Outbreak Response. Analysis of State Legal Authorities.

http://www.cifor.us/documents/CIFORAnalysisStateLegalAuthorities.pdf.