## TITLE:  Osmolality on Advanced Micro-Osmometer

# PRINCIPLE

Advanced Osmometers are devices for the determination of the concentration of solutions by means of freezing point measurement.

Advanced Osmometers utilize high-precision electronic thermometers to sense the sample temperature, to control the degree of supercooling and freeze induction, and to measure the freezing point of the sample.  They can routinely determine differences

Of 2 mOsm/kg H2O.

The quickest and most precise way to measure the freezing point of a solution is to supercool it several degrees below its freezing point and then mechanically induce the sample to freeze.  The heat of fusion suddenly liberated causes the sample temperature to rise toward a temporary plateau wherein a liquid/solid equilibrium is maintained.  The equilibrium temperature is, by definition, the freezing point of the solution.

The time over which liquid/solid equilibrium develops and is maintained is a function of the temperature differential between the sample and its environment and the ability of the surrounding materials to conduct heat.  The cooling of the sample should be rapid, but not so fast that supercooling is uncontrolled.  If the sample container and/or probe are allowed to become much colder (or much warmer) than the sample during the plateau, heat flow will distort the freezing curve.

Sensitive thermistor probes monitor the sample temperature, control the degree of supercooling and freeze induction, and measure the freezing point of the sample.  Automatic probe centering minimizes gradients and ensure uniform, precise sample temperature measurement.  Fully automatic operation minimizes imprecision due to operator technique.

Figure 1. Standard Freezing Curve

Figure 1 illustrates the temperature of a sample as it progresses through the freezing cycle and shows the action of the instrument at each stage.

## DEFINITIONS

Solution:  A homogeneous mixture of solute and solvent in which the solvent is the major component and the solute is the minor component.

Concentration:  The amount of solute in a given amount of solvent.  The amount of solute is usually expressed in terms of moles, i.e., gram molecular weight.  One mole = 6.02 x 10 23 molecules (Avogadro’s number).  One mole of glucose (180.2g) and one mole of sodium chloride (58.4 g) each contain Avogadro’s number of molecules.  Common units of concentration are:

Molality:  moles of solute per kilogram of pure solvent.

Osmolality:  osmoles of solute particles per kilogram of pure solvent.   Most ionic solutes do not completely dissociate.  Osmolality is a unit of concentration which takes into account the dissociative effect.  Osmolality is usually expressed in mOsm/kg.  One milliosmole (mOsm) is 10 3  osmoles.  Osmolality = osmol/kg Hx0 = onC, where:

o = the osmotic coefficient, which accounts for the degree of molecular dissociation.

n= the number of particles into which a molecule can dissociate.

C= the molal concentration of the solution.

Molarity: moles of solute per liter of solution.

Osmolarity:  osmoles of solute particles per liter of solution.

Although molarity and osmolarity are common units of measurement in other branches of chemistry, they are not used in osmometry because they are temperature dependent (water changes volume as the temperature varies) while molality and osmolality are not.  However, in the case of simple, fairly dilute solutions, it is possible to convert between units if necessary.

Freezing Point/Melting Point:  The temperature at which the liquid and crystalline phases of a substance will remain together in equilibrium.

Freezing Point Depression: When a solute is added to a solvent, the freezing point of the solvent is lowered.  In aqueous solutions, one milliosmole of solute per kilogram of water depresses the freezing point by 1.86 millidegrees Celsius (mC).

Supercooling:   The tendency of a substance to remain in the liquid state when cooled below its freezing point.

Crystallization Temperature: Aqueous solutions can be induced to freeze (i.e., crystallize) most reliably when supercooled.  When supercooled, crystal formation is induced by agitating

the solution (freeze pulse).  The crystallization temperature is the temperature at which crystallization is induced.  During crystallization, the heat of fusion raises the temperature of the sample to an ice/water freezing-point plateau.

Heat of Fusion: The heat released when the mobile molecules of a liquid are frozen into rigid crystals.

Freezing-Point Plateau:  The constant temperature maintained during the time that ice and liquid exist in equilibrium after crystallization is initiated.

CLINICAL SIGNIFICANCE:

Serum Osmolality can be elevated in: Water depletion, hyperosmolar nonketotic diabetic coma, diabetic ketoacidosis, hypernatremic ethanol intoxication, diabetes insipidus, hypercalcemia, cerebral lesions; often with tube feeding.

Serum Osmolality can be decreased in: Adrenocortical insufficiency, panhypopituitarism, water intoxication, postoperative states (especially with excessive water replacement therapy), syndrome of inappropriate ADH secretion.

Urine Osmolality can be decreased in: Diabetes insipidus and primary polydipsice.

READOUT-  As soon as the plateau is constant, the sample osmolality is ready.

# SPECIMENS

- No special sample or patient preparation required.

- Serum and urine are the preferred specimens.

Whole blood and plasma may be used.

- Minimum required specimen  500 u/L (Each test requires 250 uL sample run in duplicate.

If the duplicate results are not within 3 (example 299 and 297), you may need to run the specimen again until you get two results that are with in 3.

# EQUIPMENT

- The Advanced Micro-Osmometer Model 3300

- 3MO828 20-Microliter Sampler

# REAGENTS

- 290 Standard, 50 STD, 850 STD

- 2 levels of current lot Quality Control material

# CALIBRATION

Calibration of the Model 3300 is a simple procedure that requires no adjustment of the instrument by the user.  Simply run the menu-driven calibration program requiring the testing of standards at each of two calibration points.  If the repeatability is acceptable on these calibration points, the instrument automatically performs internal calibration.

The calibration procedure is as follows.  Note that maintaining and loading accurate, uncontaminated standard samples is extremely important.

- At “Osmometer Ready,” press the left soft key, labeled (CALIB) to initiate the calibration process.  The calibration program will prompt you to run a sample of the 50 mOsm calibration standard.

- Run the 50 m)sm calibration standard as you would an actual sample (see Recommended Procedure).  When the instrument completes each test and reports the results (not necessarily the exact standard value), raise the operating cradle, remove the sampler, and clean the chamber as recommended.

- The calibration program will prompt you for a second 50 mOsm calibration standard.  Run the second  50 m)sm calibration standard.  When the test finishes, your instrument will display a result and prompt you for an 850 m)sm calibration standard.

- Run your first 850 mOsm calibration standard as described above.  The program will report a result and request another 850 mOsm calibration standard.  Continue in this manner until the instrument display reads “Calibration complete”.  The program can require anywhere between three and six 850 standard samples, depending on the repeatability of the results.

- Verify the calibration by running a Clinitrol 290 reference solution before running unknown samples.

- Calibration is to be performed:  a) After major maintenance or troubleshooting is performed.  b) If QC is outside +/- 2SD.  c) every 6 months.

If the instrument display shows “Calibration Not Complete” after six samples at either calibration level, the calibration procedure has failed and you should carefully repeat the procedure.  If you experience “Calibration Not Complete” more then once, you have a repeatability problem and should consult Chapter 4 for troubleshooting and service information.

# CALIBRATION NOTES

- The Model 3300 will retain its previous calibration data until it completes a re-calibration and the display reads “Calibration Complete”.

- If the instrument has calibration information in memory and you have not changed any probes, the first test result displayed should be close to the nominal value of the standard loaded.  If the instrument has no calibration information in memory or a probe has been changed, the test result displayed may be far from the nominal value of the standard loaded.  If the displayed values repeat consistently, the calibration will automatically adjust when the calibration test sequence is complete.

- You may terminate the calibration procedure at any time by raising the operating cradle and pressing (CANCEL).  The Model 3300 will display “Test Canceled”, beep once, then display “Calib. Canceled” and beep again.  The existing calibration will not be affected.

- If you accidentally press a function soft key during a series of regular tests, the same procedure will terminate the Calibrate mode without affecting the calibration and normal testing may be continued.

# QUALITY CONTROL

- 290 Standard.

- Multiqual Normal (Level 1) and Abnormal (Level III) for serum and Urine Chemistry  Controls I & II for urine.  Refer to Quality Control Procedure in the General Chemistry book for further details.

- The 290 Standard and 2 levels of Control material (determined by patient specimen being tested) are to be run once every 24-hours when testing is ordered, and after every calibration verify the calibration.

- Enter all control values into the appropriate LIS QC program.

- Repeat and/or calibrate if any control values exceed acceptable limits.

### STEPWISE  PROCEDURE

Use the following procedure to run 290 Reference, QC,  and patient samples.

- Snap a sample tip into place on the sampler.  The sample cell must be straight and firmly seated.  Be careful not to crack the sample tip.

- Depress the sampler’s plunger and insert the sample tip at least ¼ inch (6 mm) below the surface of the fluid to be tested.  Gently release the plunger to load a 20 uL sample.

Figure 2:  Sample Cell Tips and Sample Levels

- Visually inspect the sample.  If there are any large voids or bubbles in the

sample, expel the sample and load a bubble-free sample.

- Blot the sides of the loaded sample cell with a soft, no-lint, non-ionic paper tissue to remove any clinging droplets.  Then blot the end of the cell tip to remove any fluid protruding beyond the tip.  Be careful not to wick out any of the sample.  The meniscus remaining may be slightly concave, but the sample must be slightly longer than it is wide.  (See figure 2).

- Remove the chamber cleaner from the sample port and discard.

- Holding the sampler by the barrel and letting the filled sample tip follow the guide groove into the sample port, rest the sampler within the operating cradle and beneath the cradle top.

- To start the test, push the entire operating cradle down until it reaches a positive stop.  Your instrument will run the test for approximately one minute and display the result in the format “Osmolality xxx mOsm.

- Record the results and raise the operating cradle to a positive stop.

- Remove the sampler from the operating cradle.

- Insert a clean ,dry chamber cleaner into the sample port and rotate it four or five times in both a clockwise and counterclockwise direction.  Withdraw chamber cleaner and insert the opposite end.  Rotate the chamber cleaner in the same manner and leave it in the sample port until your next test.

- Remove the used sample tip from the sampler by pressing firmly enough on the sampler plunger to dislodge the tip.  Discard the used sample tip.

- Blot the Teflon plunger tip with a soft, no-lint, non-ionic paper tissue.  Be careful not to dislodge the tip.

- To run additional tests, repeat this procedure beginning with step 1.

If the duplicate results are not within 3 (example 299 and 297), you may need to run the specimen again until you get two results that are with in 3.

# REPEATABILITY TIPS

- Treat all samples, as well as standards and reference solutions, uniformly before the test.

- Microsamples are more susceptible to contamination and evaporation then larger samples.  Avoid leaving sample containers open.  Cold samples are susceptible to condensation: warmer samples are susceptible to evaporation.

- Keep the cooling chamber clean between tests.  Never inject anything into the cooling chamber.

# REPORTING RESULTS

- Results are reported directly from the Osmometer display in m0sm/kg.

- Enter m0sm patent results into the appropriate LIS program and verify.

- Document all patient and control results on the Chemistry MISC Worksheet.

- Normal values: Serum  0-3 Months- 275-295 mOsm/kg.

3 Months – 150 Years – 280-300 mOsm/Kg

Random or spot urine – 300- 1300 mOsm/kg

24 hr urine – 300-900 Osm/kg.

### NOTES

- For best repeatability the sampler plunger should be replaced every 500 tests and calibrated.

- See the Advanced Micro-Osmometer Model 3300 User’s Guide for sample plunger replacement and verification procedure.

- Refer to User’s Guide - Troubleshooting and Service Section for any problems.

# REFERENCES

- The Advanced Micro-Osmometer Model 3300 User’s Guide.

- Clinical Guide to Laboratory Tests, Third Edition. W.B. Saunders Company, 1995.