## TITLE:  Coagulation Quality Control

PURPOSE

Quality Control is run in Coagulation to show that instrument and reagent systems are functioning properly prior to reporting patient results, and that patient results are accurate.  All Coag testing is performed at 37oC.

### PERSONNEL

Medical Technologists and Technicians

### SPECIMEN

Manufacturer’s Reagent Controls

### REAGENT PREPARATION & EQUIPMENT

1.	CA660 and instrument supplies.

2.	Dade Citrol controls: Levels 1 and 3 and D-Dimer controls: Levels 1 and 2.

3.	All reagents Coagulation are properly labeled, as applicable and

appropriate, with: the content and quantity, concentration or titer; storage requirements;

dated prepared or reconstituted/opened by the laboratory and the expiration date on the

container or the box housing the individual smaller bottles.  All reagents are stored

according to manufactures instructions and used within their expiration dates.  Reagents

with no stated manufactures expiration are used until visable or performance changes

occur, then discarded.

Citrol Levels 1 and 3:

3.	Preparation of Controls:

- Allow to warm up to room temperature slightly: 10 - 15 minutes at room

temperature.

B.	Add 1 ml of sterile deionized water using the 1000 ul MLA

pipette.  Restopper vial and gently tilt the vial to mix.

C.	Allow to stand for 15 minutes until dissolved after reconstitution.

D.	Invert gently to mix. -- DO NOT SHAKE.  Gently mix once more before use.

E.	Store at 2 to 8 degrees C when not in use.

F.	Date bottle or control sample cup with reconstitution date,

time and initials.  Expiration dates/times need to be on all reagent bottles.

- Any other control in coagulation -- prepare using manufacturer’s protocol, see

individual test policy.

### CALIBRATION

Calibration for D-Dimers is performed as necessary and at 6 month intervals.

### STEPWISE PROCEDURE

I.	Dade Citrol Quality Control - PT & PTT

A.     Level 1 and 3 (Citrol) - QC01 and QC03

Run daily each 8 hours or with the change of any consumables on the 				instrument.  Two levels of Q.C. must be run each 8 hours or with any change 			reagents, tubing etc.  (All consumables).

- Controls are good to 16 hours once reconstituted.

- Quality Control results for the CA660 will cross to the LIS.  QC from CA660 is reviewed and certified in LIS on an ongoing basis.  Monthly QC reports are printed and data sent in to the  IQAP program.

II.	D-Dimer Quality Control:

- Level 1 and 2 (Innovance) – QC02 and QC06

Run Level 1 and Level 2 with the reconstitution of new reagents, when

you have patient samples, and if the reagents have been removed from the analyzer.

Every 8 hours, run 2 levels QC if patient testing is performed.

- Controls are stable 7 days  once reconstituted.

- Quality Control results for the CA660 will cross to the LIS.  .  QC from CA 660 is

reviewed and certified in LIS on an ongoing basis.  Monthly QC reports are printed

and data sent in to the IQAP program.

III. 	Quality Control Decisions:

A.       PT & PTT

- Level 1and 3 should be within 2 SD.  Patient results will not autoverify if QC is out of 2 SD.  Rerun if not within 3 SD.  If still

not within 2 - 3 SD, prepare new bottle of Quality Control material and

rerun. Document your actions in the LIS.  If problem still persists, prepare

new reagents and rerun Any further problems - contact a Sr. Tech.

- D-Dimer

- Both levels of QC must be within tolerance limits.  Patient results will not autoverify if QC is out of 2 SD.  Rerun if not within 3 SD.    If still not within 2 - 3 SD, prepare new bottle of Quality Control material and rerun.  Document your actions in the LIS.   If problem still persists, prepare new reagents and rerun.

Any further problems - contact a Sr. Tech.

- Document action taken in the LIS for all outliers. ALL QC problems must be           documented in the LIS, it is a required action.  Patient results are not reported until QC results are acceptable.

- New Lots:  All new lots of reagents and Quality Control materials are verified

prior to being put into use by performing patient comparison studies with the

new lots and establishing quality control ranges prior to patient reporting.

- Performance Verification after service/maintenance: the FSI performs all function tests on the analyzer and runs all QC material to verify instrument function, QC is entered in the LIS, and service records for that analyzer are saved in a binder.

### PROCEDURAL NOTES

All QC results and QC outliers are reviewed on an ongoing basis in the LIS, and monthly by designated person in Heme/Coag or Sr. Tech.

### REPORTING RESULTS

QC results are entered in the LIS

REFERENCES

1.  Quality Control:  American Dade Publications by David Plant and Janis Silberman.

2.   Dade -Ci-trol Coagulation Control Product inserts.