## TITLE:   Coagulation Result Acceptance

- General Acceptance

- Tests run on Sysmex CA660

- Controls must be within acceptable range of 2 SD for patient

autoverification to work.

- No error comments on the print out or workstation will autoverify.

- Normal Range Patient Results

- Controls must be within acceptable limits prior to reporting patient results.

- Results will autoverify from the Instrument Menu.

- Abnormal, not Critical Value Patient Results

- Controls must be within acceptable limits.

- Results will autoverify if there are no error codes or flags from the analyzer.

- Critical Value Patient Results

- Controls must be within acceptable limits.

- Repeat PT with INR’s greater than or equal to 4.0 and PTT’s greater than 75 seconds.

- “Review Results” in Instrument Menu in the LIS to verify trending or first time critical.  All critical results are called and documented per the RBTO (read back) protocol of the hospital.