TITLE:  Activated Partial Thromboplastin Time – Sysmex CA-660

PRINCIPLE

The activated partial thromboplastin time (PTT) is a screening procedure for deficiencies in coagulation factors of the intrinsic coagulation pathway, factors VIII, IX, XI, XII, Fletcher factor and Fitzgerald factor. Severe deficiencies of fibrinogen and factors II, V and X can also be detected by the PTT. It is a useful and effective method for screening patients with a bleeding tendency, for evaluating the effect of therapy in procoagulant disorders and as the basis for several specific coagulation factor assay procedures. The PTT is widely advocated as a test for monitoring and regulating heparin therapy. The presence of non-specific inhibitors, such as the lupus anticoagulant, may prolong the PTT; this effect is variable and is generally recognized as being related to the nature and level of the inhibitor present. The PTT consists of recalcifying plasma in the presence of a standardized amount of platelet-like phosphatides and an activator of the contact factors of the intrinsic coagulation pathway measured at 37degrees. The PTT is the time in seconds, required for a fibrin clot to form.

#### SPECIMEN TYPE

Mix nine parts of freshly collected blood with one part of 0.11mol/L (3.2%) sodium citrate

anticoagulant.  Properly Collected, see policy 7180-CO-0120 Proper Collection & Preparation of Patient Samples.

Handling Conditions:

The specimen should be transported at room temperature. The whole blood specimen is checked for clot formation by gentle inversion and rimmed with sticks prior to centrifugation.. Centrifuge the capped specimen tube for a minimum of 10 minutes at 4500 RPM  to consistently produce platelet poor plasma (platelet count <10 x 109 /L ). To separate plasma, use a plastic transfer pipette; remove the plasma to a plastic tube without disturbing the buffy coat, respin and transfer plasma to a second plastic tube. If testing is not complete within 4 hours, specimens can be frozen at -20°C or below for short term storage (up to 2 weeks) or -70°C or below for long term storage. Frozen plasma samples must be rapidly thawed at 37C while gently mixing and tested immediately.  Mixing is critical before testing.

EQUIPMENT AND MATERIALS

Equipment:

Sysmex CA-660 System

4.0 mL Sample cups

Reaction tubes

Waste and rinse bottles

Materials:

Dade Actin FSL Activated PTT Reagent

Calcium Chloride, 0.025M

Dade Ci-Trol Coagulation Controls, Level 1, Level 3

CA Clean I

Preservative -free distilled or deionized water

Preparation:

1. Dade Actin FSL Activated aPTT Reagent -purified soy and rabbit brain     phosphatides in 1 x 10-4 M ellagic acid with added buffer, stabilizers and preservative.

Dade® Actin FSL Reagent must be mixed gently by inversion (5 to 8 times) before use. The reagent is liquid and ready for use.

Store at 2-8C. Do not freeze.

Stability:

On board		48 hours (cooled position)

24 hours (non-cooled position)

+ 2 to 8C 		used by the labeled expiry date (closed container)

+ 2 to + 15C 		7 days

2. Calcium Chloride (0.025M)

The reagent is liquid and ready to use.

Store at 2-25C when not in use.

Stability:

On board	 48 hours

+ 2 to 25C    	 Stable until the date indicated on the vial label date.

3. Dade Ci-Trol Coagulation Control, Level 1 (Normal) - Lyophilized preparation of human plasma, stabilizers and buffers. Used for Quality Control.

4. Dade Ci-Trol Coagulation Control Levels  3 (Abnormal) - Lyophilized preparations of human plasma, stabilizers and buffers. Used for Quality Control.

Stability after reconstitution:

| Condition | Ci-Trol 1, 3 |
| --- | --- |
| +2 to +8°C | 16 hours (closed vial) |
| +15 to +25°C | 8 hours (closed vial) |
| -20°C | -------------- |
| *Frozen and thawed once | NO |

*	Note:  Do not expose thawed control to 15 to 25°C for longer than 2 hours

5. CA Clean I is liquid and ready for use.
Store at 2-8 °C. Do not freeze.

Store unopened at 2-8 °C until expiration date on bottle.

Opened bottle stable for 30 days at 2-8 °C.

These products are for in vitro diagnostic use only.

Reagent Integrity:

Indication of deterioration: No evidence of vacuum in vial upon opening, difficulty in reconstituting reagent, control values outside of determined range.

WARNING: Human source material. Each donor or donor unit was tested and found to be negative for human immunodeficiency virus (HIV) 1 and 2, hepatitis B virus (HBV) and hepatitis C 	virus (HCV) using either tests found to be in conformance with the In Vitro Diagnostic Directive in 	the EU or FDA approved tests.  Because no known test can offer complete assurance of the absence of infectious agents, all human derived products should be handled with appropriate caution.

Performance Parameters:

Indication of deterioration: Normal plasma or controls will show deviations in results    from the established laboratory range.

QUALITY CONTROL:

Dade Ci-Trol Control Level 1 Normal Controls

Dade Ci-Trol Control Levels 3 Abnormal Controls

- Controls should be tested at the initiation of testing each day, upon reagent changes, and at least once each 8 hour shift.

- 2.  Controls should be run in the same manner as the test samples.

- Control tolerance limits—the range is calculated based on +/-2.0 to +/-2.5 standard deviations (SD) from the mean control value.

- Corrective action when tolerance limits are exceeded:

- Rerun out-of-range control material.

- Verify reagent performance

- Check instrument performance

- Document actions taken to identify and correct the problem before reorting any patient data

PROCEDURES FOR SETUP:

Loading Reagents: 
    (when the instrument is not operating)

- 1.  Open Light Shield.

- 2.  Place Actin FSL Reagent in Reagent holder position 2. Place CaCl2 in Reagent holder

- position 7.

- 3.  Place CA Clean I in rinse position 1, Place CA Clean II in position II.

Loading Consumables and Discarding Waste Material: 
(do the following as necessary)

- Replenish Reaction tubes. 
Reaction tubes should be replenished as needed. The instrument will keep track of the last reaction tube position used.

- Replenish Rinse Solution.

  - Confirm that the instrument is not operating or has been interrupted due to       insufficient rinse (error message “Rinse Bottle Empty”).

  - Remove the cap with the float switch from the Rinse solution container by turning the cap counter-clockwise. 
Caution: Do not touch the float switch. Rinse the float switch thoroughly with deionized or distilled water if contact is made.

  - Replace the Rinse solution with distilled or deionized water.

  - Replace the float switch by tightening the cap clockwise to close.

  - Verify that the tubing is securely connected and is not kinked.

  - If the analysis was interrupted due to insufficient Rinse, press [Start] key.

- Dispose of Liquid Waste.

  - Confirm that the instrument is not operating or has been interrupted due to waste is full.

  - Remove the cap with the float switch from the Waste container by turning the cap counter-clockwise.
Caution: Biohazard

  - Dispose of the Biohazardous liquid waste.

  - Replace the float switch by tight the cap clockwise to close.

  - Verify that the tubing is securely connected and is not kinked.

- Dispose of Used Reaction Tubes from the Trash Box.

  - Confirm that the instrument is not operating.

  - On the right side of the instrument, remove the reaction tube trash box.

  - Discard the used reaction tubes.
Caution: Biohazard

  - Clean the reaction tube trash box with tap water, and thoroughly wipe off moisture from the reaction tube trash box.

  - Replace the reaction tube trash box.

QUALITY CONTROL PROCESSING

- Controls are run from the sample rack.

- 1.  From the Main Menu, press [ID NO. ENTRY].  Use the assigned QC numbers.

- Press the [APTT] key.

- Press [ENTER].

- Place the control material in a sample rack.

- If necessary, continue to program routine patient or other levels of quality control material. When completed, press [Start].

SAMPLE PROCESSING:

Select test group to be used. Go to [Main Menu], [Settings], [Analysis Settings], [Test Group]. Arrow to the appropriate group. Select [Return], [Fix]. Return to Main Menu.

Samples can be programmed and loaded on the analyzer using one of four options depending whether the sample is barcoded and/or the host is connected:

- Manually enter sample ID (non-barcoded) / 
No automatic inquiry of tests (no host connection or the host is down)

- Manually enter sample ID (non-barcoded) / 
Automatic inquiry of tests (host connection operational)

- Sample ID number (barcoded sample) read by barcode reader / 
No automatic inquiry of tests (no host connection or the host is down)

- Sample ID number (barcoded sample) read by barcode reader / 
Automatic inquiry of tests (host connection operational)

I.	Manually enter sample ID (non-barcoded) / 
No automatic inquiry of tests (no host connection or the host is down)

- 1.   From the Main Menu, press [ID No. Entry]. If ID No. Entry does not appear, press [Next] for    next available rack/position.

- 2.   Use the up and down arrow keys to select the desired sample position.

- 3.   The numeric keypad is displayed. Enter the sample ID number.
       4.   Press the [APTT] key.

- 5.   Continue to select test to be performed on the sample if necessary.

- 6.   Press [Enter]. Each test selected should show the symbol "".

- 7.   Load the sample in the respective sample rack position.

- Repeat until all sample ID numbers and tests are entered for the respective

- sample rack positions.

- When all samples are programmed and loaded in the respective sample rack position, press [Quit] on the numeric keypad.

- With all reagents and consumables on board, place the sample rack into position.

- Press [Start].

- Select [First Tube] or [Continue]: 
[Continue] will pick up cuvettes from the last position of cuvettes used.
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message “Analyzing” will be displayed. When all samples have been aspirated, the “Replace Rack? Yes” message will appear.

II.	Manually enter sample ID (non-barcoded) / 
Automatic inquiry of tests (host connection operational)

- From the Main Menu, press [ID No. Entry]. If ID No. Entry does not appear, press [Next] for next available rack/position.

- Use the up and down arrow keys to select the desired sample position.

- The numeric keypad is displayed. Enter the sample ID number.

- Press [Enter].

- Load samples in the respective rack position.

- Repeat until all sample ID numbers are entered for the respective sample rack positions.

- When all samples are programmed and loaded in the respective sample rack positions, press [Quit] on the numeric keypad

- With all reagents and consumables on board, place the sample rack into position.

- Press [Start].

- Select [First Tube] or [Continue]. 
[Continue] will pick up cuvettes from the last position of cuvettes used.
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message “Analyzing” will be displayed. When all samples have been aspirated, the “Replace Rack? Yes” message will appear.

III. Sample ID number (barcoded sample) read by barcode reader / 
No automatic inquiry of tests (no host connection or the host is down)

- Press [Next] for next available rack/position.

- Use the up and down arrow keys to select each tube position on the rack. Press the [APTT] key.

- Continue to select test to be performed on the sample if necessary.

- Press [Enter]. Each test selected should show the symbol "".

- Load the sample in the respective sample rack position.

- With all reagents and consumables on board, place the sample rack into position. The bar code labels will face the user.

- Press [Start].

- Select [First Tube] or [Continue]. [Continue] will pick up cuvettes from the last position of cuvettes used. [First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message "Analyzing" will be displayed. When all samples have been aspirated, the "Replace Rack? Yes" message will appear.

IV. Sample ID number (barcoded sample) read by barcode reader / 
Automatic inquiry of tests (host connection operational)

- When the host computer is connected using bi-directional communication, host inquiry takes place when the sample ID is read and the analysis parameters are automatically registered.

- Load the sample in the respective sample rack position.

- Press [Start] to begin processing.

- Select [First Tube] or [Continue]. 
[Continue] will pick up cuvettes from the last position of cuvettes used. 
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message "Analyzing" will be displayed. When all samples have been aspirated, the "Replace Rack? Yes" message will appear.

NOTE: The cause of any printed error code must be investigated and the appropriate corrective action taken prior to reporting results.

REPORTING RESULTS:

Results of the PTT testing should be reported as the PTT in seconds. These results should be related to the reference interval for the PTT. Verify results in the computer system. Record quality control values in the LIS. Hemolyzed, lipemic, or icteric samples must be noted with the result.

Reference Interval:

Reference interval values determined for the current Actin FSL lot number  (+/- 2.0 SD) are:        22.4 – 35.8 seconds.

Sysmex CA-660 Reportable Range limits are: 20.0 – 205.0 seconds, report as <20 or >205.

If results are <20, before you result check sample for clot. Also, rerun sample to verify <20 or >205,  Document in comment field that sample was rerun and checked for a clot. If clot is found request redraw and cancel test.

Procedures for Critical Results:

PTT values greater than 75, and >205 seconds, and No Clot Detected must be phoned to the patient’s nurse or physician as a Critical or Physician Alert Value.  Document with (RBTO) comment; RN’s full name, time/date called and your initials.

PROCEDURE NOTES:

Overall performance of aPTT testing is dependent on reagent and instrument performance.

Acceptable variability (imprecision) should be such that the total day-to-day coefficient of variation (CV) of the analytic system is less than 5% on the same lot of control plasma.

LIMITATIONS OF THE PROCEDURE:

Plasma collection and storage conditions must be carefully controlled. Avoid small plasma volumes prior to testing which may cause pH changes in the plasma leading to inactivation of specific components of the coagulation system. Testing may be affected by a number of commonly administered drugs. Short PTT values during conjugated estrogen therapy in males and oral contraceptive administration in females may occur. Diphenylhydantoin, heparin, warfarin, naloxone and radiographic agents may increase the PTT. Hemolyzed, lipemic, chromogenic specimens may also affect results. Blood clotting factor deficiencies, which should produce prolonged PTT values may be compensated and appear normal by elevated levels of one or more different clotting factors. The presence of active intermediates, which would normally shorten the clotting time may also mask conditions that would lead to prolongation of the PTT. Mild or minor deficiencies in several factors may have an additive effect on increasing the PTT. Unexpected abnormal PTT results should be followed by additional coagulation studies to determine the cause of abnormal results.

REFERENCES:

- Dade Actin FSL Activated PTT Reagent package insert, Siemens Healthcare Diagnostics Inc., Newark, DE, June 2008.

- Dade Ci-Trol Coagulation Controls, Levels 1, 3 package inserts, Siemens Healthcare Diagnostics Inc., Newark, DE, May 2008.

- Clinical Laboratory Standards Institute. Collection, Transport and Processing of Blood Specimens for Testing Plasma-Based Coagulation Assays:  Approved Guideline-Fifth Edition. CLSI Publication H21-A5. Wayne, PA, January, 2008.

S:\Laboratory Policies and Procedures\YORKVILLE\Coagulation  12/17/2018 TML