Title:  D-DIMER (FEU) Method by Innovance for the Sysmex CA 660

PRINCIPLE:

Coagulation activation results in the cleavage of fibrinogen to fibrin. The fibrin monomers spontaneously aggregate to fibrin and are cross-linked by Factor XIII; this produces a fibrin clot. In response to the coagulation process the fibrinolytic system is activated resulting in the conversion of plasminogen into plasmin, which cleaves fibrin (and fibrinogen) into the fragments D and E. Due to the cross-linkages between the D-domains in the fibrin clot, the action of plasmin releases fibrin degradation products with cross-linked D-domains. The smallest unit is the D-Dimer. Detection of D-Dimers, which specifies cross-linked fibrin degradation products generated by reactive fibrinolysis, is an indicator of coagulation activity.

CLINICAL SIGNIFICANCE:

Elevated D-Dimer levels are observed in all diseases and conditions with increased coagulation activation, e.g. thromboembolic disease, disseminated intravascular coagulopathy (DIC), acute aortic dissection, myocardial infarction, malignant diseases, obstetrical complications, third trimester of pregnancy, surgery or polytrauma.1-6

The relevance of the D-Dimer assay is as an aid in the diagnosis of thromboembolic events. Elevated concentrations of D-Dimer are indicative of the presence of a clot and have been reported in deep vein thrombosis, pulmonary embolism and disseminated intravascular coagulation.  Innovance D-Dimer assays have been FDA cleared to exclude both PE and DVT as of January 2011.

Polystyrene particles covalently coated with monoclonal antibody (8D3)8 are aggregated when mixed with samples containing D-Dimer. D-Dimer cross-linkage region has a stereosymmetrical structure, i.e. the epitope for the monoclonal occurs twice. Consequently, one antibody suffices in order to trigger and aggregation reaction, which is then detected turbimetrically via the increase in turbidity.

SPECIMEN COLLECTION:

Type:

Mix nine parts of freshly collected blood with one part of 0.11mol/L (3.2%) sodium citrate anticoagulant. Properly Collected, see policy 7180-CO-0120 Proper Collection & Preparation of Patient Samples.

Handling Conditions:

The specimen should be transported at room temperature.  The whole blood specimen is checked for clot formation by gentle inversion and checked with applicator sticks, then centrifuged and tested within 4 hours from the time of specimen collection at room temperature, 24 hours at 2-8, or plasma frozen for 4 weeks. Centrifuge the capped specimen tube for a minimum of 10 minutes at 4500 RPM to consistently produce platelet poor plasma (platelet count <10 x 109/L).  If immediate testing is to be done, the plasma may remain on the packed cells or separated. To separate plasma, use a plastic transfer pipette; remove the plasma to a plastic tube. If testing is not completed within 4 hours or if shipment is required, the plasma may be stored frozen at -18 C or below for up to 4 weeks if frozen within 4 hours of blood collection. Frozen plasma samples must be rapidly thawed at 37 C while gently mixing and tested within 2 hours. Do not refreeze.

EQUIPMENT AND MATERAILS:

Equipment:

Sysmex CA-660 Instrument

4.0 mL Sample Cups

Reaction Tubes

Sample Plates

Waste and Rinse Bottles

Materials:

Innovance D-Dimer Kit containing:

Innovance D-Dimer Reagent
Innovance D-Dimer Buffer 
Innovance D-Dimer Supplement

Innovance D-Dimer Diluent

Innovance D-Dimer Calibrator

Control Material:

Innovance D-Dimer Control 1
Innovance D-Dimer Control 2

CA Clean I & Clean II

Preservative-free distilled or deionized water

Preparation:

All components of a kit are lot-specific.  The combination of lots other than those specified for the particular kit lot may lead to incorrect results.

Follow the preparation instructions prior to use according to Table 1.  Storage instructions are detailed in section “Storage and Stability”.

Table 1:  Instructions for the preparation of the kit components

| Instructions | Innovance D-Dimer Reagent | Innovance D-Dimer Buffer/Supplement/Diluent | Innovance D-Dimer Calibrator |
| --- | --- | --- | --- |
| Reconstitution | - Dissolve with 4.0 mL distilled water - Invert 3 times - Leave the vial for at least 15 minutes at 15 -25 º C | Ready to use | - Dissolve with 1.0 mL distilled water - Mix carefully without foam formation - Leave the vial for at least 15 minutes at 15 -25 º C |
| Prior to placing on the system | - Mix well (again) by inverting 3 times - Avoid foam formation - Remove bubbles | - Mix carefully - Avoid foam formation - Buffer only: resuspend potential precipitates by gently swirling. Any residual precipitates after resuspension do not impact test results - Remove bubbles | - Mix (again) careful - Do not use if vial contains visible clot. 3. Label the vial with the open date/time it was reconstituted and the 4 hour expiration time. |
| Aliquoting | 1. Mix well (again) by inverting 3 times 2. Aliquot into an empty vial provided with the same kit 3. Discard empty vials if unused until complete consumption of the kit. | 1. 1. Aliquot into an empty vial provided with the same kit. 2. Discard empty vials if unused until complete consumption of the kit. | N/A |
| Freeze and thaw | - Use the original container or the empty vial provided with the same kit - Follow storage instructions in the section “Storage and Stability: - Thaw at 37 °C within 10 minutes. Thereafter the vial may no longer be stored at 2 – 8 °C. - Do not freeze again after thawing. |  | Refer to section “Storage and Stability” |
| Placing on the system | Place Innovance D- Dimer Reagent in Reagent Holder position 6. | - Place Innovance D-Dimer Buffer in Reagent Holder positions 4. - Place Innovance D-Dimer Supplement in Reagent Holder positions 8. - Place Innovance D-Dimer Diluent in Reagent Holder positions 10. | Place Innovance D-Dimer Calibrator in Reagent Holder positions D1 Sample Rack position 1. |
| Note | The reconstitution, opening, or freezing date may be noted on the vial label using the framed free space. |  |  |

Storage and Stability:

The kit may be used up to the expiry date indicated on the label if stored unopened at 2 – 8 °C.

Table 2: Stability after reconstitution or first opening (closed vial)

| Temperature | Innovance D-Dimer Reagent | Innovance D-Dimer Buffer | Innovance D-Dimer Supplement | Innovance D-Dimer Diluent | Innovance D-Dimer Calibrator | Innovance D-Dimer Controls |
| --- | --- | --- | --- | --- | --- | --- |
| 2 -8 ° C | 4 weeks | 4 weeks | 4 weeks | 4 weeks | - | 1 week |
| ≤ -18 °C | 4 weeks | 4 weeks | 4 weeks | 4 weeks | - | 4 weeks |
| 15 – 25 °C | - | - | - | - | 4 hours | 8 hours |
| Onboard stability | *24 Hours | *24 Hours | *24 Hours | *24 Hours |  | **4 hours |

Do not refreeze and thaw.  Follow freeze and thaw instructions in section “Preparation of the Reagents”.

*Reagents can be stored on and off the instrument for the intervals of 1hr x 7 over a maximum of 14 days.

**After each use, store Controls back into refrigeration for optimum use of 1 week.

Innovance D-Dimer Control 1: 
Innovance D-Dimer Control 2:
 Lyophilized human plasma based products containing D-Dimer containing

- `		5-chloro-2-methyl-4-isothiazole-3-one

- 2-methyl-4-isothiazole-3-one (< 1 mg/L)

- sodium azide (< 1 g/L)

- Assayed controls used for the assessment of precision and analytical bias in the normal and pathological range for the determination of D-Dimer in the Innovance D-Dimer Kit.

- Reconstitute the Innovance D-Dimer Control 1and 2 with 1.0 mL of sterile water.

- Re-stopper vial and invert gently to dissolve (without foam formation).

- Let stand at +15 to +25 °C for at least 15 minutes before use

- Mix carefully once before using

Stability after reconstitution:

-18 C:	4 weeks (do not refreeze)**
2- 8 C: 	7 days (closed vial) 
15-25 C: 	8 hours (closed vial)

**Must be frozen in the original containers.  Do not refreeze after thawing.

Innovance D-Dimer Controls 1 and 2 may be frozen in the original contained and thawed once after reconstitution.  The previous storage time at 15-25 C must not have exceeded 4 hours. The plasma must be well sealed and frozen as quickly as possible. Thawing must be completed at 37 C and within a maximum of 10 minutes.  Innovance D-Dimer Control 1 and Innovance D-Dimer Control 2 should not stand for more than 4 hours at 15-25 C after thawing.   Innovance D-Dimer Control 1 and Innovance D-Dimer Control 2 should not be used if they contain visible clots.

- CA Clean I is liquid and ready for use

Store at 2-8 °C. Do Not Freeze.

Stable unopened at 2-8 °C until expiration date on bottle.

Opened bottle stable for 30 days at 2-8 °C.

These products are for in vitro diagnostic use only.

WARNING: Each donor or donor unit was tested and found to be negative for human immunodeficiency virus (HIV) 1 and 2, hepatitis B virus (HBV) and hepatitis C virus (HCV) using either tests found to be in conformance with the In Vitro Diagnostic Directive in the EU or FDA approved tests.  Because no known test can offer complete assurance of the absence of infectious agents, all human derived products should be handled with appropriate caution.

WARNING: Sodium azide may react with lead and copper plumbing to form highly explosive metal azides. If discarded into a sink, flush with a large volume of water to prevent azide build up.

Reagent Integrity:

Indication of deterioration: No evidence of vacuum in vial upon opening, difficulty in reconstituting reagent, control values outside of determined range

QUALITY CONTROL:

Innovance D-Dimer Control 1

Innovance D-Dimer Control 2

- Innovance D-Dimer Controls must be tested at least every eight hours of testing and for each vial of reagent for the respective measurement range to ensure that the system is functioning correctly.  Control of the lower measurement range is performed with Innovance D-Dimer Control 1, and for the upper range with Innovance D-Dimer Control 2.

- The measured values obtained must be within the acceptance ranges given in the respective Table of Assigned Values.

- If values are obtained outside the acceptance ranges, the measurement must be repeated. If the deviations are confirmed, a new calibration must be performed.

- Do not report patient results unless the cause of the deviating control results has been identified and corrected!

PROCEDURES FOR SETUP:

Loading Reagents 
(when the instrument is not performing a test)

- Verify the appropriate group (Group 1 for D-Dimer) is displayed on the 'Consumable' screen. The appropriate group will have the DDi reagents displayed.

Note: If the appropriate Group is not displayed, press [Test Group] to display the change Group screen. Press the correct group key, then [Main Menu] and [Fix] . Verify reagent position displayed in pink on the Consumable Screen contains the correct reagent.

- Add the Innovance D-Dimer Reagents into their perspective locations listed above in Table 1.

- Repeat the above procedure to add CA Clean I reagent holder position 11 and OVB buffer into holder position 12.

- Close the light shield lid.

Loading Consumables and Discarding Waste Material: 
(do the following as necessary)

- Replenish Reaction tubes

  - Reaction tubes can be replenished only when the instrument is ready or testing is interrupted.  Load reaction tubes manually into the tray carefully avoiding fingerprints on the bottom portion of the tube.

- Replenish Rinse Solution (Di-Water)

  - Confirm that the instrument is not operating or has been interrupted due to insufficient rinse (error message "Replenish Rinse Fluid").

  - Remove the cap with the float switch from the rinse solution container by turning the cap counter-clockwise.
Caution: Do not touch the float switch. 
Rinse the float switch thoroughly with deionized or distilled water if contact is made.

  - Replace the rinse solution with distilled or deionized water.

  - Replace the float switch by tightening the cap clockwise to close.

  - Verify that the tubing is securely connected and is not kinked.

  - If the analysis was interrupted due to insufficient rinse, press [Resume] key.

- Dispose of Liquid Waste

  - Confirm that the instrument is not operating or has been interrupted due to waste is full.

  - Remove the cap with the float switch from the waste container by turning the cap counter-clockwise.
Caution: Biohazard

  - Dispose of the biohazardous liquid waste.

  - Replace the float switch by tightening the cap clockwise to close.

  - Verify that the tubing is securely connected and is not kinked.

- Dispose of Used Reaction Tubes from the Trash Box

  - Confirm that the instrument is not operating.

  - On the left side of the instrument, pull out the trash box.

  - Discard the used reaction tubes.
Caution: Biohazard

  - Clean the reaction tube trash box with tap water, and thoroughly wipe off moisture from the reaction tube trash box.

  - Restore the reaction tube trash box.

CALIBRATION

Calibration criteria include: at changes of reagent lot numbers, failure to recover QC tolerance limits, after major service or maintenance on the CA 660 or at the recommendation of the manufacturer and at approximately every 6 months.

The reference curve is valid for the respective lot of the reagent employed.  A new curve should be prepared with a new lot of Innovance D-Dimer using Innovance D-Dimer Calibrator provided in the same kit and if indicated by any change in analytical conditions. The calibration curve can be used as long as the assay-dependent assigned values (i.e. Innovance D-Dimer Control 1 and Innovance D-Dimer Control 2) are within the corresponding acceptance ranges.

In order to calibrate the Innovance D-Dimer assay, a calibrator must be loaded onto the CA 660. The calibrator (Innovance D-Dimer Calibrator = DDi.CAL) can be placed in a predefined reagent holder position (procedure A) or from position 1 of a sample rack (procedure B). Follow procedure A or B given below depending on the desired placement of the calibrator. Follow procedure C to accept/update the curve and for standard curve documentation.  The package insert in the kit has the assigned calibrator value listed in concentration mg/L FEU.

A. Calibrating using a pre-defined Standard Reagent Holder position:

Software Set-up for calibration from a Reagent Holder position

- From the 'Main Menu', press [Standard Curve] to display the 'Standard Curve' screen.

- [Select test] choosing the Innovance D-Dimer assay: Press [DDi].

- Press [Lot no. Entry] and enter all the reagents for D-Dimer including expiration dates. Enter and Fix.

- Press [Standard Analysis] to display the 'Standard Curve Analysis Setting' screen.

- Enter the assay value of the calibrator from the Innovance D-Dimer Calibrator lot number package insert.

- Press [Enter] on the key pad.

- Verify the replication for each dilution ratio is set to 2.

  - Use the arrow keys to highlight the incorrect replication field.

  - Enter in the keypad "2".

  - Press [Enter] on the keypad

- Pour the reconstituted Innovance D-Dimer Calibrator into a 4 mL analyzer cup.

- Place the analyzer cup into the first position of the sample rack.

- Press [Start] on the right hand corner of screen.

- Upon completion of the calibration see procedure "B”. Accept/Update the Standard Curve" below.

B. Accept/Update/Print  the Standard Curve

Update the standard curve

- When the standard curve analysis is complete, press [Standard Curve] at the 'Main Menu' to display the' Standard Curve' screen.

- Press select test,  Press [DDi].

- Check the cal date, if it displays the current day that it completed the calibration then it updated.

- Print  the standard curve

- Run both levels of QC material to validate new curve.

QUALITY CONTROL PROCESSING :

- Innovance D-Dimer Control Plasma 1 and 2

:

- Select [ID No. Entry], and type the QC code appropriate for the level of D-Dimer you are running, select test [DDi] and hit enter.  When finished entering select [Quit].

- Pour the control material into a labeled 4 mL analyzer cup. Place the cup in the appropriate sample position.

- Verify that the in use reagent lot number and reagent expiration date are valid.

- 4.   Press [Start] to begin processing.

- 5.   Select [First Tube] or [Continue]. 
[Continue] will pick up cuvettes from the last position of cuvettes used. 
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message "Analyzing" will be displayed. When all samples have been aspirated, the "Replace Rack? Yes" message will appear.

- Results cross into LIS, verify they are in range according to laboratory procedure

SAMPLE PROCESSING:

Samples can be programmed and loaded on the CA 660 using one of four different options depending whether the sample is barcoded and/or the host is connected:

- Manually enter sample ID (non-barcoded) / 
No automatic inquiry of tests (no host connection or the host is down)

- Manually enter sample ID (non-barcoded) / 
Automatic inquiry of tests (host connection is operational)

- Sample ID number (barcoded sample) read by bar-code reader / 
No automatic inquiry of tests (no host connection or the host is down)

- Sample ID number (barcoded sample) read by bar-code reader / 
Automatic inquiry of tests (host connection is operational)

- Manually enter sample ID (non-barcoded) / 
No automatic inquiry of tests (no host connection or the host is down)

- From the Main Menu, press [ID NO. Entry]. If ID No. Entry does not appear, press [Next] for next available rack/position.

- Use the up and down arrow keys to select the desired sample position.

- The numeric keypad is displayed. Enter the sample ID number.
Press the [DDi] key.

- Continue to select test to be performed on the sample if necessary.

- Press [Enter]. Each test selected should show the symbol "".

- Load the sample in the respective sample rack position.

- Repeat until all sample ID numbers and tests are entered for the respective sample rack positions.

- When all samples are programmed and loaded in the respective sample rack position, press [Quit] on the numeric keypad.

- With all reagents and consumables on board, place the sample rack into position.

- Press [Start].

- Select [First Tube] or [Continue]: 
[Continue] will pick up cuvettes from the last position of cuvettes used.
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message “Analyzing” will be displayed. When all samples have been aspirated, the “Replace Rack? Yes” message will appear.

- Manually enter sample ID (non-barcoded) / 
Automatic inquiry of tests (host connection is operational)

- 1.    From the Main Menu, press [ID NO. Entry]. If ID No. Entry does not appear, press [Next] for next available rack/position.

- Use the up and down arrow keys to select the desired sample position.

- The numeric keypad is displayed. Enter the sample ID number.

- Press [Enter].

- Load samples in the respective rack position.

- Repeat until all sample ID numbers are entered for the respective sample rack positions.

- When all samples are programmed and loaded in the respective sample rack positions, press [Quit] on the numeric keypad

- With all reagents and consumables on board, place the sample rack into position.

- Press [Start].

- Select [First Tube] or [Continue]. 
[Continue] will pick up cuvettes from the last position of cuvettes used.
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message “Analyzing” will be displayed. When all samples have been aspirated, the “Replace Rack? Yes” message will appear.

- Sample ID number (barcoded sample) read by barcode reader/ 
No automatic inquiry of tests (no host connection or the host is down)

- 1.   From the Main Menu, press [ID NO. Entry]. If ID No. Entry does not appear, press [Next] for next available rack/position.

- Use the up and down arrow keys to select each tube position on the rack. Press the [PT-INN] key.

- Continue to select test to be performed on the sample if necessary.

- Press [Enter]. Each test selected should show the symbol "".

- Load the sample in the respective sample rack position.

- With all reagents and consumables on board, place the sample rack into position. The bar code labels will face the user.

- Press [Start].

- Select [First Tube] or [Continue]. [Continue] will pick up cuvettes from the last position of cuvettes used. [First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message "Analyzing" will be displayed. When all samples have been aspirated, the "Replace Rack? Yes" message will appear.

- Sample ID number (barcoded sample) read by barcode reader / 
Automatic Inquiry of tests  (host connection is operational)

- 1.   When the host computer is connected using bi-directional communication, host inquiry takes place when the sample ID is read and the analysis parameters are automatically registered

- 2.  Load the sample in the respective sample rack position.

- 3.  Press [Start] to begin processing.

- 4.Select [First Tube] or [Continue]. 
[Continue] will pick up cuvettes from the last position of cuvettes used. 
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message "Analyzing" will be displayed. When all samples have been aspirated, the "Replace Rack? Yes" message will appear.

REPORTING RESULTS:

The D-Dimer concentration in mg/L is calculated automatically by the analyzer based on the reference curve. The D-Dimer level is expressed as initial fibrinogen equivalent unit (FEU). An FEU is the quantity of fibrinogen initially present that leads to the observed D-Dimer level.  Increases in D-Dimer concentration observed with thromboembolic events can be variable due to localization, extension and age of the thrombus. Therefore, a thromboembolic event cannot be excluded with certainty solely on the basis of a D-Dimer concentration being within the reference range of ostensibly healthy persons.

Record and report patient and quality control values according to laboratory procedure.

Hemolyzed, lipemic, or icteric samples must be noted with the result.

Reference Range:

Reference Range values determined for D-Dimer are:

Less than or equal to 0.59mg/L FEU for non ED Patients

Less than 0.50 mg/L FEU for ED Patients

Cut-off value for D-Dimer is: 0.50mg/L FEU

PROCEDURE NOTES:

Overall performance of D-Dimer testing is dependent on reagent and instrument performance. Acceptable variability (imprecision) should be such, that the total coefficient of variation (CV) of the analytic system is less than 20% on the same lot of control plasma.

A new reference curve should be established with each change of reagent lot and change of instrument or with any deviation from control or proficiency testing limits. Recalibration occurs if there are any problems with the curve, or outlier points.  One point can be repeated or a new set of reagents/calibrator made up and run.  Troubleshoot with the hot line for further issues.

D-Dimer concentrations below 3.65 mg/L FEU obtained by redilution measurement have to be confirmed without redilution.

The Innovance D-Dimer total measuring range is defined by the concentration of the calibrator used and is approximately 0.19 to 4.40 mg/L FEU.  Report a result less than .19 as <0.19.  The measuring range can be extended to approximately 33.50 mg/L FEU by automatic redilution of samples above 4.40 mg/L FEU.  We report results out as an alpha result of: >34.00mg/L FEU.

All components of the Innovance D-Dimer Kit are lot dependent.  The combination of lots other than those specified may lead to incorrect results.

LIMITATIONS OF THE PROCEDURE:

Turbidity and particles in plasma may interfere with the determination.  Therefore plasmas containing particles must be centrifuged for 10 minutes at approximately 4500 RPM again prior to testing.

Lipemic samples or samples containing particles that cannot be clarified by centrifuging must be excluded from testing.

Higher levels of lipids or turbid samples can lead to falsely elevated or decreased values.

Patient samples may contain heterophilic antibodies (e.g. human anti-mouse antibodies (HAMA) and rheumatoid factors) that could react with immunoassays to give a falsely elevated or depressed result.  This assay has been designed to minimize interference from heterophilic antibodies. Nevertheless, complete elimination from this interference cannot be guaranteed

Results of the D-Dimer test should always be interpreted in conjunction with the patient’s medical history, clinical presentation, and other findings. Clinical diagnosis should not be based on he results of INNOVANCE D-Dimer alone.

Patients with a distal DVT may have a normal INNOVANCE D-Dimer result.

Interfering Substances

Levels of the following do not interfere with the Innovance D-Dimer assay:

| Analyte | Interference Up to |
| --- | --- |
| Bilirubin | 15 mg/dL |
| Hemoglobin | 100 mg/dL |
| Triglycerides | 600 mg/dL |

General Reagent Specificity

There is no significant effect on the determination of D-Dimer due to levels of the following:

| Analyte | Interference Up to |
| --- | --- |
| Creatinine | 30 mg/dL |
| Heparin, sodium | 3.0 U/mL |
| Rheumatoid Factors | 1330 IU/mL |
| Albumin | 6 g/L |
| Fibrinogen | 1000 mg/dL |
| Urea | 500 mg/dL |
| Uric Acid | 20 mg/dL |
| Immunoglobulin G (IgG) | 5 g/dL |

REFERENCES

- Innovance D-Dimer package insert. Dade Behring. Marburg, Germany, April 2007

- Innovance D-Dimer Control package insert. Dade Behring. Marburg, Germany, April 2007.

- Dade® CA System Buffer package insert. Dade Behring. Marburg, Germany, May 2008

- Dade® Owrens Veronal Buffer package insert. Dade Behring. Marburg, Germany, May, 2008

- Clinical Laboratory Standards Institute. Collection, Transport and Processing of Blood Specimens for Testing Plasma-Based Coagulation Assays: Approved Guideline-Fourth Edition. 
CLSI Publication H21-A5. Wayne, PA,  January, 2008

- Sysmex® Operator’s Manual CA-500 Automated Coagulation Analyzer. No. 461-2662-6. Kobe, Japan, October 2003

- Application Sheet for Innovance D-Dimer on CA-660 Automated Coagulation Analyzer