## TITLE:   Alere i Strep A 2 Assay

## PRINCIPLE:

The Alere i Strep A 2 assay performed on the Alere™ i Instrument is a rapid molecular in vitro

diagnostic test utilizing an isothermal nucleic acid amplification technology for the qualitative

detection of Streptococcus pyogenes, Group A Streptococcus bacterial nucleic acid in throat

swab specimens obtained from patients with signs and symptoms of pharyngitis. It is intended

to aid in the rapid diagnosis of Group A Streptococcus bacterial infections.

Alere™ i Strep A 2 utilizes isothermal nucleic acid amplification technology for the qualitative

detection of Group A Strep bacterial nucleic acids. It is comprised of a Sample Receiver,

containing elution buffer, a Test Base, comprising two sealed reaction tubes, each containing a

lyophilized pellet, a Transfer Cartridge for transfer of the eluted sample to the Test Base, and

the Alere™ i Instrument.

The reaction tubes in the Test Base contain the reagents required for Group A Strep bacterial

lysis and the subsequent amplification of the target nucleic acid and an internal control. Alere™ i

Strep A 2 utilizes templates (similar to primers) for the specific amplification of DNA from Group

A Strep and a fluorescently-labelled molecular beacon designed to specifically identify the

amplified nucleic acid target.

To perform the assay, the Sample Receiver and Test Base are inserted into the Alere™ i

Instrument. The sample is added to the Sample Receiver and transferred via the Transfer

Cartridge to the Test Base, initiating bacterial lysis and target amplification. Heating, mixing and

detection are provided by the instrument, with results automatically reported.

CLINICAL SIGNIFICANCE:

Streptococcus pyogenes, Group A Strep is the most significant pathogen causing pharyngitis.

Accurate diagnosis of the etiological agent is necessary to properly treat the disease. In the

case of Group A Strep, antibiotic therapy is the treatment of choice. If left untreated, serious

complications such as rheumatic fever may occur.

SPECIMEN TYPE:

Use freshly collected specimens for optimal test performance. Inadequate specimen collection

or improper sample handling/storage/ transport may yield erroneous results.

For optimal performance, use the swabs provided in the test kit. Alternatively, foam, polyester,

HydraFlock® and nylon flocked throat swabs can be used to collect throat swab samples. The

BBL™ CultureSwab™ Liquid Amies transport media system has been tested and is also

acceptable. The dry transport sterile foam applicator in the plastic tube is the same applicator included in the kit and is the preferred collection device if any specimen transport will take place.

Rayon swabs and the BBL™ CultureSwab™ Liquid Stuart transport media (red capped) system are not suitable for use in this assay.

Collect patient specimen by swabbing the posterior pharynx, tonsils and other inflamed areas.

Avoid touching the tongue, cheeks and teeth with the swab.

Soft test order code: STRPA – Rapid Strep A PCR

Specimen Transport and Storage

Swab specimens should be tested as soon as possible after collection. If immediate testing is

not possible, the throat swab can be held in its original package or a clean, dry plastic tube or

sleeve at room temperature (approximately 22°C) or refrigerated at 2-8°C for up to seventy-two

(72) hours prior to testing.

REAGENTS AND EQUIPMENT:

A. Materials

1. Alere™ i Strep A 2 assay test kit:

2. Alere™ i Instrument

3. Alere™ universal printer

4. Thermal printer 59MM Labels (manufacturer part number 26333)

5. Barcode Scanner

6. Gloves / Personal Protective Equipment (PPE)

B. Storage and Stability:

Store kit at 2-30°C. The Alere i Strep 2 kit is stable until the expiration date marked on the

outer packaging and containers.

Ensure all test components are at room temperature before use.

C. Warnings and Precautions:

- Treat all specimens as potentially infectious. Follow universal precautions when

handling samples, this kit and its contents. Dispose of used reagents in biohazard containers

- Proper sample collection, storage and transport are essential for correct results.

Leave test pieces sealed in their foil pouches until just before use.

3.   Do not tamper with test pieces prior to use.

4.   Do not use kit past its expiration date.

5.   Do not mix components from different kit lots.

6.   Solutions used to make the control swabs are inactivated using standard methods.

However, patient samples, controls, and test pieces should be handled as though

they could transmit disease. Observe established precautions against microbial

hazards during use and disposal.

7.   If any assay components are dropped cracked, found to be damaged or opened

when received, DO NOT USE and discard. Do not use scissors or sharp objects to

open foil pouches as damage to test pieces can occur. Do not touch the test tubes

contained within the Test Base.

8.   Do not use a Transfer Cartridge if it is dropped after aspiration of the sample. If the

Transfer Cartridge is dropped, discard the component and continue the test by

transferring the sample with a new Transfer Cartridge.

9.   Do not open the Sample Receiver before placing in the instrument. It will prohibit the

Elution Buffer from reaching temperature and may impact test performance.

10. If the Sample Receiver is spilled while opening, clean the instrument per instructions

provided in the instrument User Manual and cancel test. Repeat test with a new

Sample Receiver.

11. All test pieces must be removed from the instrument according to removal

instructions displayed on the instrument, and disposed of according to country and

local requirements. Pieces must not be separated once they are assembled.

12.  All test pieces are single use items. Do not use with multiple specimens.

13. Once reacted, the Test Base contains large amounts of amplified target (Amplicon).

Do not disassemble the Test Base and Transfer Cartridge. In the case of a positive

sample, this could lead to amplicon leakage and potential Alere i Strep A 2 false

positive test results.

14. At a low frequency, clinical samples can contain inhibitors that may generate invalid

results.

15. Due to the high sensitivity of the assays run on the instrument, contamination of the

work area with previous positive samples may cause false positive results. Handle

samples according to standard laboratory practices. Clean instruments and

surrounding surfaces according to instructions provided in the cleaning section of the

instrument User Manual. Refer to Section 1.6, Maintenance & Cleaning, for further

information.

16. Do not touch the heads of the Control Swabs. Cross contamination with the Positive

Control Swabs may occur due to their high target level and the sensitivity of the

assays run on the instrument.

17. Test results should be interpreted in conjunction with other laboratory and clinical

data.

18. The performance of Alere™ i Strep A 2 was evaluated using the procedures

provided in this package insert only. Modifications to these procedures may alter the

performance of the test.

19. False negative results may occur if a specimen is improperly collected, transported

or handled; or if inadequate quantities of target DNA are present in the system.

20. To avoid contamination, do not move the Alere™ i Instrument during a run or until all

assay components have been removed from the instrument.

21. The orange indicator should rise when the Transfer Cartridge is pressed into the

Sample Receiver until a click is heard. The indicator should descend fully when

correctly connected to the Test Base. Failure to follow this procedure may lead to

false negative or invalid results.

22. Dispose of PPE following hospital protocol.

D. Cleaning and Maintenance

1.   The Alere™ i instrument is maintenance-free and has no serviceable parts. In the

case of instrument failure or damage, contact Point of Care Department.

2.   Clean using 70% ethanol or hospital approved disinfectant wipe (10% bleach

solution). Do not spray or pour solution directly onto instrument when cleaning.

Ensure no excess liquid is used when cleaning as it may damage the instrument.

3.   Do not disassemble the instrument for cleaning.

4.   Do not immerse in water or cleaning solutions.

5.   Do not clean with soap or other solutions.

CALIBRATION:

The Alere™ i Instrument is factory calibrated and does not require any further calibration and

verification. However, if the instrument was transported or moved, a performance check using

Alere™ I positive and negative controls is recommended to ensure proper functionality.

QUALITY CONTROL:

Alere i Strep A 2 has built-in procedural controls. The result of the Procedural Control is

displayed on the screen and is automatically stored in the instrument with each test result.

A. Internal Control:

1. Alere i Strep A 2 contains an internal control that has been designed to control for

sample inhibition, amplification and assay reagent function. In positive samples

where target amplification is strong, the internal control is ignored and the target

amplification serves as the ‘control’ to confirm that the clinical sample was not

inhibitory and that assay reagent performance was robust. At a very low frequency,

clinical samples can contain inhibitors that may generate invalid results.

2. Procedural Control Valid displayed on the instrument screen indicates that the assay

reagents maintained their functional integrity and the sample did not significantly

inhibit assay performance. The Procedural Control is automatically stored in the

instrument with each test result. This can be reviewed later by selecting Review

Memory on the instrument.

B. External Positive and Negative Controls:

1. Good laboratory practice suggests the use of positive and negative controls to

ensure that test reagents are working and that the test is correctly performed.

2. Alere i Strep A 2 kits contain Positive and Negative Control Swabs. These swabs

may be used to demonstrate the ability to generate appropriate positive and negative

results by following the assay process.

3. Test Positive and Negative Control Swabs once with each new shipment received.

4. Test Positive and Negative Control Swabs with every new lot number of reagent.

5. Test Positive and Negative Control Swabs once for each untrained operator.

6. Further controls may be tested in order to conform to local, state and/or federal

regulations, accrediting groups, or your lab’s standard Quality Control procedures.

Note: Run both Positive and Negative Controls on the same Alere i instrument.

C. To Perform a QC Test:

4. Confirm Test

5. QC Sampe ID

The user has the option to enter an ID for the QC Sample being run.

Enter or Scan QC Sample ID. Press ‘OK ’ to confirm.

The QC Sample ID entered will be displayed. Press OK to confirm or cancel to reenter.

If it is not desired to enter or scan in the QC sample ID leave blank and press ‘OK ’ to

continue

Note: The QC test is run in the same manner as a Patient Test. See the To Perform a

Test section above for step by step

### STEPWISE PROCEDURE:

Before testing with Alere™ i Strep A 2:

 Allow all samples to reach room temperature

 Allow all test pieces to reach room temperature

 Check that a reagent pellet is visible at the bottom of each of the reaction tubes prior to

inserting the Test Base in the Alere™ i Instrument. Do not use the Test Base if a pellet is

not visible at the bottom of each reaction tube.

To Perform Test:

1. Follow the step-by-step instructions shown on the instrument screen.

Note: If testing a swab transported in media, the collection swab is to be tested following the

step-by-step instructions shown on the instrument screen.

Note: The optimum environmental operating conditions for Alere™ i Strep A 2 are: 15-30°C and

10-80% relative humidity.

Step 1

Turn on the Alere™ i Instrument - press the power button on the side of the instrument.

Note: If the unit is unattended for one hour, the instrument will go to

a black screen power save mode. Touch the screen to return the unit

to active display operation.

Caution: Once the Test Base has been placed in the holder, the user will have 10 minutes to

confirm the test. If the test is not confirmed within 10 minutes, the instrument will time out and

the Test Base must be removed and discarded.

Note: If the incorrect Test Base has been inserted, remove and dispose of the incorrect Test

Base. Close the lid. The instrument will then run a self-test before proceeding to the Home

screen. Press Run Test and restart the test using the correct Test Base.

Caution: Do not apply excessive force. Excessive force could damage the instrument.

Caution: Confirm that the Sample Receiver foil pouch indicates that it is for use with Alere™ i

Strep A 2 (not another Alere™ i assay). Confirm that the foil seal on the Sample Receiver is for

the Strep A assay. If not, then remove the Sample Receiver and replace it with a new Sample

Receiver for Alere™ i Strep A 2.

Caution: Once the Sample Receiver has been placed in the holder, the user will have 10

minutes to start the test (Steps 3 through 5). If the test is not started within 10 minutes, the

instrument will time out and all test pieces (Test Base and Sample Receiver) must be removed

and discarded. The instrument will proceed to the Home screen. Press Run Test and restart the

test using a new Test Base and Sample Receiver.

Caution: To ensure the Sample Receiver remains in the instrument while removing the foil seal,

place two fingers along the outer edge of the Sample Receiver to hold it in place. If the Sample

Receiver spills after warm up, cancel the test by pressing the Home button. Remove and

discard the test pieces (Sample Receiver and Test Base) and clean the instrument. Press Run

Test to start a new test using a new Test Base and Sample Receiver.

DO NOT OPEN THE LID until the Test Complete message appears on the screen.

Note: The test will be cancelled if the lid is opened.

Caution: This screen will be displayed for up to 30 seconds once the Transfer Cartridge is

detected. If the instrument does not detect that the lid has been closed by then, it will time out

and all test pieces (Sample Receiver, Test Base, and Transfer Cartridge) must be removed and

discarded. The instrument will proceed to the Home screen. Collect a new sample from the

patient. Press Run Test and restart the test using a new Test Base and Sample Receiver.

Caution: DO NOT OPEN THE LID. The test will be cancelled and all test pieces (Sample

Receiver, Test Base, and Transfer Cartridge) must be removed and discarded. A test result will

not be reported or saved in the instrument

When amplification and detection is complete, the instrument will automatically save the

data before advancing to the results screen.

Caution: The test is not saved until the completed result is displayed. Do not open the lid until

the results are displayed.

The Test Results screen displays either a Negative or Positive result for a successfully

completed test. If a test error occurs, the display will read ‘Invalid’. Refer to the Result

Interpretation Section for Interpretation of Results.

Press Print to print test results, press New Test to run another test, Press Home to return

to the Home screen

After printing, or if New Test or Home are selected, the instrument will prompt to open

the lid and discard the used test pieces.

Remove test pieces by lifting the Transfer Cartridge attached to the Test Base, and clicking it

into the Sample Receiver, by pressing into the Sample Receiver.

Caution: Do not try to remove the Sample Receiver by any other method as there is a risk of

spilling the patient sample.

All test pieces will be connected and can now be removed from the instrument and disposed of

according to federal, state and local regulations.

Caution: DO NOT disassemble the Transfer Cartridge and the Test Base before disposal.

Close the lid. The instrument will then run a Self-Test before showing the Home screen or

Enter Patient ID screen, depending on the previous selection.

CALCULATIONS:

N/A

REPORTING AND INTERPRETING OF RESULTS:

When the test is complete, the results are clearly displayed on the instrument screen.

Choose the appropriate result from the resulting keypad.

| Instrument Display | Interpretation of Results and Followup Actions |
| --- | --- |
|  | Positive for Strep A nucleic acid. Note: Control Valid is the internal QC result. |
|  | Negative for Strep A nucleic acid. Note: Control Valid is the internal QC result. |
|  | Invalid. Immediately repeat testing of the sample using new test components. If repeated Invalid results are obtained, repeat test with a new patient sample and new test components. |

NOTE: If an Invalid result is received, one additional test may be run using the same Sample

Receiver. The instructions below should be followed:

a. Remove the connected Test Base and Transfer Cartridge from the instrument and

connect the Test Base portion to an UNUSED Sample Receiver. The connected

Test Base and Transfer Cartridge MUST be attached to a Sample Receiver prior to

disposal. The Sample Receiver from a new Transfer Cartridge package may be

used for this.

b. Remove the blue Sample Receiver separately and carefully from the instrument.

The Sample Receiver should be retained and kept upright, to avoid spilling the

liquid contents.

c. From the Home Screen, start a new test. Follow the screen prompts, however when

asked to insert the Sample Receiver, reuse the Sample Receiver and DO NOT reelute

the swab.

d. If the test result is invalid 2 times call the provider and request reorder and recollection. Choose QC fail x2 from the resulting keypad.

PROCEDURAL NOTES:

1. The performance of the Alere i Strep A 2 was evaluated using the procedures provided

in this package insert only. Modifications to these procedures may alter the performance

of the test.

2. For optimal performance, use the specimen collection swabs provided in the test kit.

Analytical studies have demonstrated that rayon swabs and BBL CultureSwab Liquid

Stuart Medium are not suitable for use with this assay and may produce false negative

results.

3. Alere™ i Strep A 2 does not distinguish between viable and nonviable organisms

4. Performance of Alere™ i Strep A 2 has not been established for monitoring treatment of

pharyngitis caused by Group A Strep.

5. Alere™ i Strep A 2 will not differentiate asymptomatic carriers of Group A Strep from

those exhibiting streptococcal infection.

6. False results may occur if a Sample Receiver for an assay other than Alere™ i Strep A 2

is used.

7. Additional follow-up testing using the culture method is required if the result is negative

and clinical symptoms persist, or in the event of an acute rheumatic fever (ARF)

outbreak.

EXPECTED VALUES:

Overall incidence of Group A Strep in patients tested during the 2017 clinical study was

20.2% (198/981) as determined by the reference culture procedure and 25.2% (247/981) as

determined by Alere™ i Strep A 2.

INTERFEREING SUBSTANCE:

The following substances, naturally present in throat swab specimens or that may be artificially

introduced into the throat, were evaluated with Alere™ i Strep A 2 at the concentrations listed

below and were found not to affect test performance.

REFERENCES:

1. Shulman et al., Clinical Practice Guideline for the Diagnosis and Management of Group

A Streptococcal Pharyngitis: 2012 Update by the Infectious Diseases Society of America

(2012).

2. American Academy of Pediatrics. Red Book: 2012 Report of the Committee on

Infectious Diseases. Pickering LK, ed. 29th ed. Elk Grove Village, IL: American Academy

of Pediatrics; 2012.US Department of Health and Human Services PHS/CDC/NIH.

Biosafety in microbiology and biomedical laboratories, Washington DC: US Government

Printing Office, 2007.

3. Manual of Clinical Microbiology, 11th Edition, Vol. 1, ASM. (2015) pg. 279

4. AlereTM i Strep A 2 Package Insert, Alere Scarborough, Inc., 10 Southgate Road,

Scarborough, Maine 04074, USA, www.alere.com, IN425024 Rev. 1 2018/04

5. Technical Support 1-855-731-2288

6. or ts.scr@alere.com