## TITLE:  Calculations Performed by LIS

## PRINCIPLE:  Calculated values that are reported with patient results are to be reviewed every two years or when a system change is made that may affect calculation.

## SPECIMEN COLLECTION/TREATMENT:

- Serum or plasma – separated from cells by a gel or aliquoted to a pour off tube.

## QUALITY CONTROL:

Follow Chemistry Analyzer procedures for running tests required for calculations.

## STEPWISE PROCEDURE:

Utilizing the Chemistry calculation spreadsheet (Attatchement S:\Yorkville\Architect\QC\Calculation Verification Template), check each calculation performed by the computer for accuracy

- Order TESTX patients in LIS and perform testing as a normal patient

- Compare the computer calculation to a manual or “calculator” calculation

- Attach analyzer printouts, instrument menu screen shots, instant reports and EPIC reports to the calculation verification spreadsheet

- Indicate if the comparison is acceptable

- If follow up is needed, document on the form what actions were taken

- Forward the completed spreadsheet to Lead.

- Serum Calculations:

- Globulin and A/G Ratio

Globulin = Total Protein – Albumin

A/G Ratio =     Albumin

Globulin

B.	Bun/Creat Ratio	=  BUN

###### Creat

C.	LDL/HDL Ratio         = LDL

HDL

D.	CHOL/HDL  Ratio     = CHOL

###### HDL

- Anion Gap  (AGAP) = Sodium – (Chloride + CO2)

- Indirect Bilirubin = Total Bilirubin – Direct Bilirubin

- Non HDL = Chol - HDL

## REPORTING RESULTS:

- All manually performed calculations should match the computer deteremed calculations

- Any discrepancies ashould be resolved and documented

- Calculation Verification should be reviewed ans signed by the Medical Director or designee.

- Serum Calculations

- Globulin and A/G Ratio are report to one decimal place.

- BUN/Creat ratio is reported as a whole number.

- LDL/HDL is reported to one decimal place.

- CHOL/HDL is reported to one decimal place.

- Calculated LDL is reported as a whole number.

- Anion Gap is reported as a whole number.

- Indirect Bilirubin is reported to one decimal place.

- Non HDL is reported as a whole number.

References: CAP checklist item GEN. 43450