## TITLE:  Result Entry through an Instrument Interface - LIS

## PRINCIPLE:

Most of the lab’s analyzers are interfaced with the Laboratory Information System (LIS), allowing bi-directional communication between the analyzer and LIS. A specimen barcode is printed from a Zebra printer and that accession information is downloaded from the LIS to the analyzer upon receipt of the specimen in the LIS. The analyzers with barcode readers scan the barcode and receive the necessary specimen information from the LIS. They then run the requested test(s) on that specimen. The accession number can also be manually entered into the analyzer. The results are then sent back to the LIS and are accessed in Instrument Menu.

PERSONNEL:  Medical Technologists and Technicians

## STEPWISE PROCEDURE:

Interfaced Analyzers:

| Main Lab | Yorkville Lab | Respiratory |
| --- | --- | --- |
| BCSXP1 | ABL 80 YK | ABL 825 1 |
| Sysmex CA 1500 | Sysmex 560 YK | ABL 825 2 |
| Coulter LH 780 | Coulter ACT2 | ABL 90 |
| Coulter DxH 800 | Coulter ACT5 |  |
| Echo Lumena | Roche U411 |  |
| Roche U411 Abbott Architect 1 Abbott Architect 2 EpiCenter Gene Expert | Abbott Architect YK Alere Triage |  |

- Login to SoftLab, if not already logged in.

- Run the specimen on the analyzer.

  - For the ABLs, scan or manually enter the accession number into the analyzer. The ABLs will receive the patient information from the LIS.

  - For all other interfaced instrument, place the barcoded specimen on the analyzer. The LIS will send the patient and order information to the analyzer.

- When the results are complete, they will be sent to the LIS.

- In Soft, click on the Instrument Menu icon

- Select the Instrument you want from the list.

- It defaults to the current date, choose “Not Posted” for “What to Display” and leave “How to Display” at “Sequence”.  Click OK.

- Your patient ‘Order Numbers’, Barcode Numbers and Patient Names are displayed as a list on the left side of your screen, followed by any flags, time and the ordered tests.

- Click on the line to review the patient information/demographics and the results.   Make any necessary repeats, checks etc. if there are flags or errors on the result.  If the result was “normal” or “abnormal”, but not a panic, absurd or flagged with an error, the result will automatically post as long as QC has been run and verified and is within SD limits.

- Only results that require review or further action will be on this list. There will be an error or flag notation on the Instrument Menu that will stop the posting of results.  See department specific policies for further instructions.