## TITLE:   ARCHITECT CALIBRATION CONFIGURATION

## PRINCIPLE:

The purpose of this procedure is to give detailed instructions for calibration configuration on the Architect ci8200.

CLINICAL SIGNIFICANCE:

Perform this procedure to add a new lot number to a configured calibrator set and change the default lot.

REAGENTS:

See individual assay specific procedures

EQUIPMENT:

Architect ci8200

QUALITY CONTROL:

See individual procedures for more information and 4840-CH-255 Chemistry Quality Control.

### STEPWISE PROCEDURE:

i System Reagents

These analytes have the calibrator factors pre built into the Architect.  To change to a new lot of calibrator follow these steps.

- Go to Orders -- Calibration Orders

- Enter cassette and position

- Select assay that needs new calibrator lot entered

- Select Assay Options…(F5) at bottom of screen

- Enter lot number and expriation date and Done

- Add order to save the lot number (if not currently running the calibration, go to order status, select the calibrator and delete. The new calibrator information will be saved)

c System Reagents

Perform this procedure to enter a new calibrator lot number and/or change the default lot.

- All the modules need to be Stopped or Ready

- Log in under System Administrator.  Login: ADMIN  Password: ADM

- Go to System, then Configuration, then QC – Cal settings

- Select Calibrator Set from the QC – Cal categories list on the Configuration screen.

- Select the desired calibrator set, and then select F6 – Configure.  The Configure calibrator set window displays.

- Select the Lot number list button, and then select one of the following:

- The desired lot number if changing the default lot #

- New Lot if entering a lot that has not been previous built

- If the new lot has already been configured check the Default box of the correct lot #.

- If a new lot has not been configured, select New Lot and enter the box Lot number and date in the Lot number data entry box.

- Enter a value in the Expiration date data entry box.

- Highlight the assay(s) from the table, and then select Define data.  The Define calibrator data window displays.

- Enter a value in the Concentration data entry box for each calibrator.

- Use the previous/next buttons to display each assay if you selected more than one, repeat step 11 for each.

- Select Done to return to the Configure calibrator set window.  The assay-specific data displays in the table. Review entered values in Details.

- Select Done to save your changes.

- Return Modules to Running Status.

Import Calibrator Data

(MCC Calibrator)

Perform this procedure to import calibrator set data from a USB flash drive to the SCC.

- Select Calibrator set from the QC – Cal categories list on the Configuration screen.

- Select the desired calibrator set from the Calibrator sets list an then select F6 – Configure.  The configure calibrator set window displays.

- Select the Lot number list button and then select Import.

- Insert USB flash drive (in the MCC calibrator box) and select OK.  USB port is located on the CPU in the lower right hand side of Architect.

- Select the desired data file and then select Done.

- Select Done to return to the Configure calibrator set window.

- Select Done to save the imported data.

- Select OK if the data was imported from a USB flash drive then remove the drive from the USB port.

ASSAYS:

ICT Serum Calibrator

- Prebuilt into Architect so update calibrator lot number like i side reagents.

- Na, K, Cl

ACET

- Acetaminophen (ug/mL)

- 1 calibrator level

- Calibrator in the reagent box.  Value on the top of box.

AmmCA

- Ammonia (umol/L)

- 1 calibrator level

- Calibrator in the reagent box.  Value on reagent box and bottle.

Bil

- BiliD and BiliT (mg/dL)

- 2 calibrator levels

- Calibrator values on paper in cal box

C02C

- C02 (mmol/L)

- 2 calibrator levels

- Calibrator values on the bottles

CRP32

- CRP (mg/L)

- 6 calibrator levels

- Cal 1 5.00 (White)

- Cal 2 10.00 (Lt. Yellow)

- Cal 3 20.00 (Lt. Green)

- Cal 4 40.00 (Lt. Blue)

- Cal 5 160.00 (Magenta)

- Cal 6 320.00 (Brown)

DOAQ

- AmpQ, BarbQ, BenzQ, CocQ, OpiQ, and PCPQ

- 2 calibrator levels: DOA Neg Cal and DOA MC1 Cal 2

- Use DOA MC1 Cal 2 calibrator lot number/expiration for the calibrator set information

- DOA Neg Calibrator enter value of 0 for Cal 1

- Enter DOA MC1 Cal 2 values in Cal 2

- Values on the outside of Calibrator boxes and package insert

THCQ

- THC (ng/mL)

- 2 calibrator levels: DOA Neg Cal and THC50 Cal

- Use THC50 Cal calibrator lot number/expiration for the calibrator set information

- DOA Neg Calibrator enter value of 0 for Cal 1

- Enter THC50 Cal value in Cal 2 (50 ng/mL)

- Values on the outside of Calibrator boxes and package insert

ETH

- Ethanol (mg/dL)

- 2 calibrator levels:  Ethanol Neg Cal and Ethanol 100 Cal

- Use Ethanol 100 calibrator lot number/expiration for the calibrator set information

- Ethanol Neg Calibrator value 0.0 mg/dL for Cal 1

- Ethanol 100 calibrator value 100.00 mg/dL for Cal 2

- Values on the outside of Calibrator boxes and package insert

Lipase

- Lipase (U/L)

- 1 calibrator

- Calibrator value is on the bottle

LIPIDMCC

- DLDL and UHDL (mg/dL)

- Lipid Multiconstituent Calibrator

- 1 calibrator

- Calibrator values on paper in cal box

MCC

- AlbG, CaC, Chol, CreaC, Crea-U, GluC, LactA, Mag, TP, Trig,Fe, Phos,Urea, Urea-U

- Multiconstituent Calibrator

- 2 calibrators

- Use USB card to import calibrator data

SALI

- Salicylate (mg/dL)

- 1 calibrator

- Calibrator value on the box and bottle

UPRO

- Urine/CSF Protein (mg/dL)

- 5 calibrators

- Calibrator values on package insert and bottles

Ha1C

- Glycoslated Hemaglobin (Ha1c)

- 2 calibrator levels

- Calibrator values on package insert in cal box

Clinical Chem Calibrator

- Lithium

- 1 calibrator level

- Calibrator value on package insert in cal box

MicroAlbumin Calibrator

- Urine MicroAlbumin

- 5 calibrator levels

- Calibrator values listed on bottles

Specific Proteins Multiconstituent Calibrator

- IgG, IgM, IgA, C3, C4, Transferrin

- 5 calibrator levels

- Calibrator values listed on package insert in cal box

TDM Multiconstituent Calibrator

- Vancomycin and Valproic Acid

- 6 calibraor values

- Calibrator values listed on package insert in box

Tobramycin Calibrator

- Tobramycin

- 6 calibrator values

- Calibrator values lised on package insert in box

| Ref # | Calibrator | Analytes | Open Expiration |
| --- | --- | --- | --- |
| 1E64-02 | Carbon Dioxide Calibrator (2) | CO2C | 30 days @ 2-8oC |
| 6K26-10 6K26-12 | CRP Calibrator Set (use CRP05, CRP10, CRP20, CRP40, CRP160, CRP320) (Discard Pink bottle CRP180) | CRP32 | 3 months @ 2-8oC |
| 3L43-05 3L41-02 | DOA Neg Calibrator THC50 Cal calibrator | THC5Q | Until kit Expiration @ 2-8oC |
| 3L36-01 3L36-02 | Ethanol Neg Calibrator Ethanol 100 Calibrator | EtohM (Alc) | Until kit Expiration @ 2-8oC |
| 1E69-03 | Fe/Mg - Iron/Magnesium Calibrator (2) | Mg, Fe | 30 days @ 2-8oC |
| 1E46-03 | ICT Serum Calibrator (2) | Cl-C, K-C, Na-C | 7 days @ 2-8oC |
| 3E16-02 | Lipase Calibrator (1) | Lip | 28 days @ 2-8oC |
| 1E66-04 | Bilirubin Calibrator (2) | BiliD, BiliT | 7 days @ 2-8oC |
| 5P56-01 | Lipid Multiconstituent Calibrator (1) | DLDL, UHDL | 7 days @ 2-8oC |
| 1E65-05 | MCC - Multiconstituent Calibrator (2) | AlbG, BUN(Urea), CaC, Chol, CreaC, GluC, LactA, TP, Trig, Phos,UA, Fe | Until kit Expiration @ 2-8oC |
|  | None (Water Blank only) | ALT, AST, AlkP, Amy, CK, LDH, GGT | on board analyzer |
| 3K01-01 | Salicylate Calibrator (1) | Sali | Until kit Expiration @ 2-8oC |
| 2K42-01 | STAT CK-MB Calibrators (6) | CK-MB | Until kit Expiration @ 2-8oC |
| 2K43-01 | STAT Myoglobin Calibrators (6) | MYOGLOBIN | Until kit Expiration @ 2-8oC |
| 2K41-01 | STAT Troponin-I Calibrators (6) | Tnl II (Troponin) | Until kit Expiration @ 2-8oC |
| 6K89-30 | Ammonia Ultra (1) in kit | Amm | 120 days @ 2-8oC |
| 7K78-01 | Total B-hCG Calibrator set (6) | STAT-B-hCG | Until kit Expiration @ 2-8oC |
| 7K62-01 | TSH Calibrators (2) | TSH | Until kit Expiration @ 2-8oC |
| 1E71-02 | Urine/CSF Protein Calibrator (5) | Upro/CSF | Until kit Expiration @ 2-8oC |
| 8K28-02 | BNP Calibrators (6) | BNPStat | Until kit Expiration @ 2-8oC |
| 3L43-05 3L43-02 | DOA Neg Calibrator DOA MC1 Cal 2 calibrator | AmpQ, BarbQ, BenzQ, CocQ, OpiQ, PCPQ | Until kit Expiration @ 2-8oC |
| 506-30 | Acetaminophen L3K Calibrator (1) in kit | Acet | Until kit Expiration @ 2-8oC |
| 1L79-01 | Anti-HCV calibrator | Anti-HCV | Until kit Expiration @ 2-8oC, store upright |
| 7K61-02 | B12 calibrator | B12 | Stable 1 month @ 2-8°C after opening |
| 2K45-01 | CA 125 calibrator | CA125 | Until kit Expiration @ 2-8oC |
| 2K44-01 | CA 15-3 Calibrator | CA 15-3 | Until kit Expiration @ 2-8oC |
| 2K91-01 | CA 19-9 Calibrator | CA19-9 | Until kit Expiration @ 2-8oC |
| 1p36-01 | Carbamazepine Calibrator | Carbamazepine | Until kit Expiration @ 2-8oC, Do not warm to RT |
| 7K68-02 | CEA Calibrator | CEA | Until kit Expiration @ 2-8oC, reconstituted 2 days @ 2-8°C or 14 days @ -10° |
| K630-10 | Multigent Clinical Chemistry Cal | Lithium | Until kit Expiration @ 2-8oC |
| 8D15-02 | Cortisol Calibrator | Coritsol | Store @ -10° - after thaw store @ 2-8°C for 90 days |
| IP32-01 | Digoxin Calibrator | Digoxin | Until kit Expiration @ 2-8oC |
| 7K72-01 | Estradiol Calibrator | Estradiol | Until kit Expiration @ 2-8oC |
| 7K59-01 | Ferritin Calibrator | Ferritin | Until kit Expiration @ 2-8oC |
| 7K63-01 | FT3 Calibraor | FT3 | Until kit Expiration @ 2-8oC |
| 7K65-01 | FT4 Calibrator | FT4 | Until kit Expiration @ 2-8oC |
| 1P31-01 | Gentamicin Calibrator | Gentamicin | Until kit Expiration @ 2-8oC- Do not warm to RT |
| 4P52-02 | Hemoglobin A1c Calibrator | Ha1c | Until kit Expiration @ 2-8oC, reconstituted 10 days at 2-8 C |
| 6L21-01 | HAVAB-M Calibrator | HAVAB-M | Until kit Expiration @ 2-8oC, store upright |
| 4P53-01 | HBsAG Calibrator | HBsAG | Until kit Expiration @ 2-8oC, store upright |

| 2P36-01 | HIV Calibrator | HIV | Until kit Expiration @ 2-8oC, store upright |
| --- | --- | --- | --- |
| 7F93-20 | Multigent Tobramycin Cal | Tobra | Until kit Expiration @ 2-8oC |
| 6P22-01 | PCT Calibrator | PCT | Store @-10°C Discard after 3 thaw cycles |
| IE78-02 | Specific Proteins Multiconstituent Calibrator | C3,C4, IgA, IgG, IgM, Transferrin | Unopened 2-8°C until kit exp. Open 30 days @ 2-8°C |
| 6L23-01 | Core-M Calibrator | Core-M | Until kit Expiration @ 2-8oC |
| 1P74-02 | Folate Calibrator | Folate | Protect from light, Store @-10°C Discard after 3 thaw cycles |
| 8K25-03 | Intact PTH Calibrator | IPTH | Store at -20°C until 1st use. After thaw stable at 2-8°C for 30 days |
| ARK Diag. | Methotrexate Calibrator | Methotrexate | stable for 12 months @ 2-8°C or kit expiration date |
| 2K98-03 | MicroAlbumin Calibrator | MicroAlbumin | 2-8oC for 6 months |
| 1P33-01 | Phenobarb Calibrator | Phenobarbital | Until kit Expiration @ 2-8oC |
| 1P34-01 | Phenytoin Calibrator | Phenytoin | Until kit Expiration @ 2-8oC |
| 7K77-01 | Progesterone Calibrator | Progesterone | store -10°C until 1st use- after opening store @ 2-8°C for 3 weeks |
| 8D06-04 | Syphilis Calibrator | Syphilis | Until kit Expiration @ 2-8oC, store upright, do not warm to RT |
| 1P29-01 | Theophylline Calibrator | Theophylline | Until kit Expiration @ 2-8oC |
| 6C06-01 | Total PSA Calibrator | PSA | Until kit Expiration @ 2-8oC |
| 5P04-01 | Multiconstituent Calibrator (TDM MCC) | Vanco, Valproic Acid | opened 60 days @ 2-8°C or 24 hrs at 15-30°C |
| 6K30-10 | Clinical Chemisty Calibrator | Lithium | 2 days @ 2-8°C or 14 days @ -20°C thaw 1x only |
| 5P02-02 | Vitamin D Calibrator | Vitamin D | Until kit Expiration @ 2-8oC |

REPORTING AND INTERPRETING OF RESULTS:

Calibration records are reviewed for acceptability and QC is verified before running patient specimens.

REFERENCES:

Architect System Operations Manual,  September 2012