## TITLE:  Coronavirus (2019-nCoV) Testing to IDPH

## Principle:    	To properly prepare and ship specimens for Coronavirus (2019-nCoV) testing at Illinois Department of Public Health laboratories and to correctly process results when received.

Personnel:	Phlebotomists and Medical Technologists

Referred Laboratory:		Illinois Department of Public Health (IDPH)

Specimen Requirements:	Nasopharyngeal Swab in 2mL VTM Media or

Oropharyngeal Swab in 2 mL VTM Media or

Sputum (not induced), 2 mL placed in VTM Media

Specimen Stability:	Specimens should be stored at 2-8° C, shipped on ice packs, and tested within 72 hours.

Specimens received on Fridays after 12pm or over a weekend must be frozen at -70° C until shipment and must be sent on dry ice. Write “Order Dry Ice for Monday” on blood bank dry erase board along with the date and your initials.

Required Information:  	Upon receipt of sample, page Infection Prevention (630) 218-3425.

Infection Prevention will obtain an authorization from the county health department and provide the lab with the authorization code.

2019-nCoV coronavirus testing sent to IDPH (Illinois Department of Public Health) requires authorization from a County Health Department.  Do not send specimens without authorization.

IDPH “Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form” is required and must be filled out by collecting unit.

LIS Order Code:	MISC – comment will state “COVID 19”

Stepwise Procedure:

Specimen Processing at Main Laboratory:

- Upon receipt of specimen and completed IDPH “Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form”, page Infection Prevention at (630) 218-3425. Notify Infection Prevention RN that lab has received a specimen for 2019 novel coronavirus testing.

- The Infection Prevention RN will obtain the IDPH authorization code so that the specimen may be sent for testing and will provide the lab with the authorization code.

- While waiting for the authorization code, receive specimen(s) into the LIS using computer LAB1212. Print 3 extra labels.

- Verify that the IDPH “Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form” is completely filled out.

- If a nasopharyngeal or oropharyngeal specimen is received, place the specimen in the appropriate storage location (See Specimen Stability section) until packaging for shipment.

- If sputum is received, transfer the sputum into VTM media under the hood in microbiology, then place the specimen in the appropriate storage location (See Specimen Stability section) until packaging for shipment.

- Completely fill out the IDPH Communicable Diseases Laboratory Requisition.

- Make copies of all forms to keep in Miscellaneous Laboratory Tests file box.

- The IDPH Laboratory is open Monday through Friday from 8a to 4:30pm, and closed on federal holidays. Specimens are only accepted during open hours.

- Specimens will be sent via MedSpeed Courier Monday through Friday only and must be picked up before 2pm.

- Package the specimens in a biohazard specimen bag with absorbent material and include the IDPH Communicable Diseases Laboratory Requisition, the Health Department Authorization, and the Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form.

- Specimens are to be shipped in a white shipping box with cooler. These shipping boxes are stored in the in the store room.

  - For refrigerated specimens:  Place a frozen ice pack in the bottom of the cooler and cover it with a paper towel. Then place the packaged specimen(s) on top of the paper towel. Replace the cooler lid, tape the shipping box shut, and place a green shipping label with the IDPH address on the box.

  - For frozen specimens: Place the packaged specimen(s) in the bottom of the cooler and cover with a paper towel. Place a bag of dry ice on top of the paper towel. Replace the cooler lid, close and tape the box shut, and place a green shipping label with the IDPH address on the box.

- Delivery Address:

ATTN: Viral Diagnostics Laboratory

Illinois Department of Public Health

2121 West Taylor Street

Chicago, IL 60612

- Call MedSpeed (630) 617-5050 and inform them we have a specimen to be sent to IDPH in Chicago.

- Write a note on the white board that there is a specimen going to IDPH and place the box on the counter above the freezer.

- Place a LIS label on the yellow Reference Lab Testing Log and write the name of the lab and test being sent next to it.

- Place the copies of the requisition with health department authorization, Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form, and extra LIS label in the Miscellaneous Laboratory Tests file box under the first letter of the patient’s last name.

Specimen Processing at Yorkville Laboratory:

- Upon receipt of specimen and completed IDPH “Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form”, page Infection Prevention at (630) 218-3425. Notify Infection Prevention RN that lab has received a specimen for 2019 novel coronavirus testing.

- The Infection Prevention RN will obtain the IDPH authorization code so that the specimen may be sent for testing and will provide the lab with the authorization code.

- Verify that the IDPH “Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form” is completely filled out.

- Completely fill out the IDPH Communicable Diseases Laboratory Requisition, including health department authorization code written at the top.

- Package the specimens in a biohazard specimen bag with absorbent material and include the IDPH Communicable Diseases Laboratory Requisition, the Health Department Authorization, and the Interim 2019 novel coronavirus (2019-nCoV) patient under investigation (PIU) form.

- Send the specimen and forms with cold packs to the Main Laboratory for shipment to IDPH Laboratories.

- Keep the specimen refrigerated until the courier arrives for pickup.

Results Processing:

- When final results are received, remove requisition from Miscellaneous Laboratory Tests file box.

- Place the extra LIS label on the results and place the results in the “Reports to be Scanned” basket in the LCC area.

- Page Infection Prevention and inform them of the results.

- LCC phlebotomy staff will scan the results into the LIS.

- If all testing is complete, discard the requisition in the appropriate PHI receptacle.

## REFERENCE: U.S. Department of Health and Human Services

Centers for Disease Control and Prevention