# Copley Memorial Hospital

Policy & Procedure

| Title: POCT ID NOW™ COVID-19 Assay | Author: BreAnn Greer |
| --- | --- |
| Date Initiated: March 30, 2020 Next Review Date: March 30, 2022 |  |

Purpose

This policy and procedure establishes guidelines for staff at Copley Memorial Hospital that perform point of care testing (POCT) COVID-19 test using the ID NOW™ instrument.

Policy

### PRINCIPLE:

The ID NOW™ COVID-19 assay performed on the ID NOW™ Instrument is a rapid molecular in vitro diagnostic test utilizing an isothermal nucleic acid amplification technology intended for the qualitative detection of nucleic acid from the SARS-CoV-2 viral RNA in direct nasal, nasopharyngeal or throat swabs and nasal, nasopharyngeal or throat swabs eluted in viral transport media from individuals who are suspected of COVID-19 by their healthcare provider.

ID NOW™ COVID-19 is an automated assay that utilizes isothermal nucleic acid amplification technology for the qualitative detection of SARS-CoV-2 viral nucleic acids. It is comprised of a Sample Receiver, containing elution/lysis buffer, a Test Base, comprising two sealed reaction tubes, each containing a lyophilized pellet, a Transfer Cartridge for transfer of the eluted sample to the Test Base, and the ID NOW Instrument.

The reaction tubes in the Test Base contain the reagents required for amplification of SARS-CoV-2, as well as an internal control. The templates (similar to primers) designed to target SARS-CoV-2 RNA amplify a unique region of the RdRp segment. Fluorescently-labeled molecular beacons are used to specifically identify each of the amplified RNA targets.

To perform the assay, the Sample Receiver and Test Base are inserted into the ID NOW™ Instrument. The sample is added to the Sample Receiver and transferred via the Transfer Cartridge to the Test Base, initiating target amplification. Heating, mixing and detection are provided by the instrument.

# CLINICAL SIGNIFICANCE:

# Results are for the identification of SRS-CoV-2 RNA. The SARS-CoV-2 RNA is generally detectable in respiratory samples during the acute phase of infection. Positive results are indicative of the presence of SARS-CoV-2 RNA; clinical correlation with patient history and other diagnostic information is necessary to determine patient infection status. Positive results do not rule our bacterial infection or co-infection with other viruses. Testing facilities within the United States and its territories are required to report all positive results to the appropriate public health authorities. Negative results do not preclude SARS-CoV-2 infection and should not be used as the sole basis for patient management decisions. Negative results must be combined with clinical observations, patient history, and epidemiological information.

# Procedure

Specimen COLLECTION:

A. 	Positive Patient Identification – Refer to hospital policy and procedure Patient

Identification & Banding for Inpatient and Outpatient.

B.	Specimen Identification – Label the swab, at the time of sample collection, with the

patient chart label which includes patient’s full name and financial number.

C.  	Specimen Collection:

Use freshly collected specimens for optimal test performance. Inadequate specimen collection or improper sample handling/storage/ transport may yield erroneous results.

For optimal performance, use the swabs provided in the test kit. Alternatively, foam, rayon, HydraFlock® Flocked swab (standard tip) and HydraFlock® Flocked  swab (mini-tip), Copan Mini Tip Flocked Swab, or Copan Standard Flocked swabs can be sued to collect nasal swab samples.

Puritan PurFlock Standard Tip Ultra Flocked Swabs, Puritan PurFlock Mini Tip Ultra Flocked Swabs and Copan Standard Rayon Tip Swabs are not suitable for use in this assay.

To collect a nasal swab sample, carefully insert the swab into the nostril exhibiting most visible drainage, or the nostril that is most congested if drainage is not visible. Using gentle rotation, push the swab until resistance is met at the level of the turbinates (less than one inch into the nostril). Rotate the swab several times against the nasal wall then slowly remove from the nostril. Using the same swab, repeat sample collection in the other nostril.

Specimen Transport and Storage

Direct nasal, throat, or nasopharyngeal swabs should be tested as soon as possible after collection. If immediate testing is not possible, the nasal, throat, or nasopharyngeal swab can be held in its original package at room temperature (15-30°C) for up to two hours prior to testing. If a direct throat or nasopharyngeal swab specimen will be held longer than two hours it must be refrigerated at 2-8°C and tested within 24 hours from the time of collection.

### REAGENTS AND EQUIPMENT:

- Materials

- ID NOW™ COVID-19  assay test kit:

Test Base	      Sample 		Transfer

Receiver		Cartridge

2. ID NOW™ Instrument

3. ID NOW™ universal printer

4. Thermal printer 59MM Labels (manufacturer part number 26333)

5. Barcode Scanner

6. Gloves / Personal Protective Equipment (PPE)

7. Patient Swabs: Sterile swabs (foam) for use with the ID NOW COVID-19 Test.

- Storage and Stability:

Store kit at 2-30°C. The ID NOW Strep 2 kit is stable until the expiration date marked on the outer packaging and containers. Ensure all test components are at room temperature before use.

- Warnings and Precautions:

- Cleaning and Maintenance

- The ID NOW™ instrument is maintenance-free and has no serviceable parts.  In the case of instrument failure or damage, contact Point of Care Department.

- Clean using 70% ethanol or hospital approved disinfectant wipe (10% bleach solution).  Do not spray or pour solution directly onto instrument when cleaning.  Ensure no excess liquid is used when cleaning as it may damage the instrument.

- Do not disassemble the instrument for cleaning.

- Do not immerse in water or cleaning solutions.

- Do not clean with soap or other solutions.

CALIBRATION:

The ID NOW™ Instrument is factory calibrated and does not require any further calibration and verification.  However, if the instrument was transported or moved, a performance check using ID NOW™ positive and negative controls is recommended to ensure proper functionality.

TEST PROCEDURE

Before testing with ID NOW™ COVID-19:

- Allow all samples to reach room temperature

- Allow all test pieces to reach room temperature

- Check that a reagent pellet is visible at the bottom of each of the reaction tubes prior to inserting the Test Base in the ID NOW™ Instrument. Do not use the Test Base if a pellet is not visible at the bottom of each reaction tube.

To Perform Test:

1. Follow the step-by-step instructions shown on the instrument screen.

Note: If testing a swab transported in media, the collection swab is to be tested following the step-by-step instructions shown on the instrument screen.

Note: The optimum environmental operating conditions for ID NOW™ COVID-19  are: 15-30°C and 10-80% relative humidity.

### STEPWISE PROCEDURE:

### Step 1

Turn on the ID NOW™ Instrument - press the power button on the side of the instrument.

Note: If the unit is unattended for one hour, the instrument will go to

a black screen power save mode. Touch the screen to return the unit

to active display operation.

Enter User ID

Press ‘ ’ after entry.

Touch ‘Run Test’

Touch ‘COVID-19’

Enter Patient ID using on screen keyboard or barcode scanner

Touch ‘ ’.

Verify that the ID was entered correctly, then touch ‘ ’ to confirm entry.

Step 2

Open the Lid and Insert Orange Test Base into Orange Test Base holder

Caution: Do not apply excessive force. Excessive force could damage the instrument.

Confirm that the correct test is displayed on the screen.

Touch ‘OK’ to proceed.

Caution: Once the Test Base has been placed in the holder, the user will have 10 minutes to confirm the test. If the test is not confirmed within 10 minutes, the instrument will time out and the Test Base must be removed and discarded.

Note: If the incorrect Test Base has been inserted, remove and dispose of the incorrect Test Base. Close the lid. The instrument will then run a self-test before proceeding to the Home screen. Press Run Test and restart the test using the correct Test Base.

Step 3

Insert Blue Sample Receiver into the Blue Sample Receiver holder

Caution: Do not apply excessive force. Excessive force could damage the instrument.

Caution: Confirm that the Sample Receiver foil pouch indicates that it is for use with ID NOW™ COVID-19 (not another ID NOW™ assay). Confirm that the foil seal on the Sample Receiver is for the COVID-19 assay. If not, then remove the Sample Receiver and replace it with a new Sample Receiver for ID NOW™ COVID-19.

Caution: Once the Sample Receiver has been placed in the holder, the user will have 10 minutes to start the test (Steps 3 through 5). If the test is not started within 10 minutes, the instrument will time out and all test pieces (Test Base and Sample Receiver) must be removed and discarded. The instrument will proceed to the Home screen. Press Run Test and restart the test using a new Test Base and Sample Receiver.

Wait for the Sample Receiver to Warm Up. Do not remove the Sample Receiver from the instrument once Warm Up begins.

Step 4

Direct Nasal, Throat, or Nasopharyngeal Swab Test Procedure

When prompted, remove the foil seal and place the patient swab to be tested into the Sample Receiver.

Caution: To ensure the Sample Receiver remains in the instrument while removing the foil seal, place two fingers along the outer edge of the Sample Receiver to hold it in place. If the Sample Receiver spills after warm up, cancel the test by pressing the Home button. Remove and discard the test pieces (Sample Receiver and Test Base) and clean the instrument. Press Run Test to start a new test using a new Test Base and Sample Receiver.

Discard the swab.

Skip to Step 5a

Nasal, Throat or Nasopharyngeal Swab Eluted in Transport Media Test Procedure

When prompted, remove the foil seal and add 0.2ml of sample to the Sample Receiver using the disposable pipettes provided in the kit.

Step 5a

Press the White Transfer Cartridge into the Blue Sample Receiver

Listen for a click.

When the Transfer Cartridge is properly attached to the Sample Receiver, the orange indicator on the Transfer Cartridge will rise. If the orange indicator does not rise, continue pushing onto the Sample Receiver until it does.

Caution: The orange indicator should be observed closely. If the orange indicator does not fully rise, the Transfer Cartridge may not collect enough sample and may lead to false negative or invalid results.

Step 5b

Lift and then connect the Transfer Cartridge to the Test Base

When the Transfer Cartridge is properly attached to the Test Base, the orange indicator on the Transfer Cartridge will descend. If the orange indicator does not descend, continue pushing onto the Test Base until it does.

Caution: If the orange indicator does not fully descend, not enough sample will be dispensed. This may potentially result in invalid or false negative results.

Step 6

Close the Lid.

DO NOT OPEN THE LID until the Test Complete message appears on the screen.

Note: The test will be cancelled if the lid is opened.

Caution: This screen will be displayed for up to 30 seconds once the Transfer Cartridge is detected. If the instrument does not detect that the lid has been closed by then, it will time out and all test pieces (Sample Receiver, Test Base, and Transfer Cartridge) must be removed and discarded. The instrument will proceed to the Home screen. Collect a new sample from the patient. Press Run Test and restart the test using a new Test Base and Sample Receiver.

Caution: DO NOT OPEN THE LID. The test will be cancelled and all test pieces (Sample Receiver, Test Base, and Transfer Cartridge) must be removed and discarded. A test result will not be reported or saved in the instrument

When amplification and detection is complete, the instrument will automatically save the data before advancing to the results screen.

Caution: The test is not saved until the completed result is displayed. Do not open the lid until the results are displayed.

The Test Results screen displays either a Negative or Positive result for a successfully completed test. If a test error occurs, the display will read ‘Invalid’. Refer to the Result Interpretation Section for Interpretation of Results.

Press Print to print test results, press New Test to run another test, Press Home to return to the Home screen

After printing, or if New Test or Home are selected, the instrument will prompt to open the lid and discard the used test pieces.

Remove test pieces by lifting the Transfer Cartridge attached to the Test Base, and clicking it into the Sample Receiver, by pressing into the Sample Receiver.

Caution: Do not try to remove the Sample Receiver by any other method as there is a risk of spilling the patient sample.

All test pieces will be connected and can now be removed from the instrument and disposed of according to federal, state and local regulations.

Caution: DO NOT disassemble the Transfer Cartridge and the Test Base before disposal.

Close the lid. The instrument will then run a Self-Test before showing the Home screen or Enter Patient ID screen, depending on the previous selection.

### QUALITY CONTROL:

ID NOW COVID-19 has built-in procedural controls. The result of the Procedural Control is displayed on the screen and is automatically stored in the instrument with each test result. This can be reviewed later by selecting Review Memory on the instrument.

- Internal Control:

- ID NOW COVID-19 contains an internal control that has been designed to control for sample inhibition, and assay reagent function. In positive samples where target amplification is strong, the internal control is ignored and the target amplification serves as the ‘control’ to confirm that the clinical sample was not inhibitory and that assay reagent performance was robust. At a very low frequency, clinical samples can contain inhibitors that may generate invalid results.

- Procedural Control Valid displayed on the instrument screen indicates that the assay reagents maintained their functional integrity and the sample did not significantly inhibit assay performance.

- External Positive and Negative Controls:

- Good laboratory practice suggests the use of positive and negative controls to ensure that test reagents are working and that the test is correctly performed.

- ID NOW COVID-19 kits contain Positive and Negative Control Swabs. These swabs will monitor the entire assay.

- Test Positive and Negative Control Swabs once with each new shipment received.

- Test Positive and Negative Control Swabs with every new lot number of reagent.

- Test Positive and Negative Control Swabs once for each untrained operator.

- Further controls may be tested in order to conform to local, state and/or federal regulations, accrediting groups, or your lab’s standard Quality Control procedures.

Note: Run both Positive and Negative Controls on the same ID NOW instrument.

- To Perform a QC Test: Touch ‘Run QC Test’

| - Touch ‘COVID-19’ E. F. Confirm Test Confirm the test type to match the QC sample intended for testing by touching ‘OK’ and following the on screen prompts to complete testing. The user has the option to enter an ID for the QC Sample being run.. |  |
| --- | --- |

RESULT INTERPRETATION:

When the test is complete, the results are clearly displayed on the instrument screen.

NOTE:

If an Invalid result is received, one additional test may be run using the same Sample Receiver. The instructions below should be followed:

- Remove the connected Test Base and Transfer Cartridge from the instrument and connect the Test Base portion to an UNUSED Sample Receiver. The connected Test Base and Transfer Cartridge MUST be attached to a Sample Receiver prior to disposal. The Sample Receiver from a new Transfer Cartridge package may be used for this.

- Remove the blue Sample Receiver separately and carefully from the instrument. The Sample Receiver should be retained and kept upright, to avoid spilling the liquid contents.

- From the Home Screen, start a new test. Follow the screen prompts, however when asked to insert the Sample Receiver, reuse the Sample Receiver and DO NOT re-elute the swab.

LIMITATIONS:

- The performance of the ID NOW COVID-19 was evaluated using the procedures provided in this package insert only. Modifications to these procedures may alter the performance of the test.

- False negative results may occur if a specimen is improperly collected, transported, or handled. False negative results may also occur if amplification inhibitors are present in the specimen or if inadequate levels of viruses are present in the specimen.

- As with any molecular test, if the virus mutates in the target region, COVID-19 may not be detected or may he detected less predictably.

- The test cannot rule out diseases caused by other bacterial or viral pathogens.

ANALYTICAL STUDIES (LIMIT OF DETECTION):

REFERENCES:

4.ID NOWTM  COVID-19 Package Insert, Abbott Diagnostics Scarborough, Inc., 10 Southgate Road, Scarborough, Maine 04074, USA, IN425024 Rev. 1 2018/04

5.Technical Support 1-855-731-2288 or ts.scr@abbott.com