## TITLE:  Accumulation of Normal Population Statistics Protocol

PURPOSE

Because test results can be different for different populations, it is important to establish normal population studies on the population from which most of the test samples are

collected  to ensure correct and adequate normal ranges.  This is accomplished annually with the rollover of new lot reagents and QC materials. Verification of our Reference Range as well as comparisons of “current vs new lots” and instrument comparisons are performed at this time.

### PERSONNEL

Medical Technologists and Technicians

### EQUIPMENT & REAGENT PREPARATION

CS-2500 and CA-660 Coagulation Analyzers

### SPECIMEN

1.	Blood is collected using a nontraumatic venipuncture technique.  It is

placed into a blue top (sodium citrate 3.2) evacuated test tube and mixed well

with the anticoagulation.  The final proportion of blood to anticoagulant is

9:1 so tubes cannot be over or under filled.

2.	Visually inspect the test tube for clots before centrifugation.

3.	Spin the sample as quickly as possible after collection.

Spin at 4500 rcf for 3 minutes in the coagulation centrifuge or at the speed and time to

render platelet poor plasma.

4.	The centrifuged plasmas should be tested immediately.  If delays are

expected decant separated plasmas into clean, properly identifiable

plastic tubes with a plastic pipette making sure no red cells are picked up.

If delays are expected, cover the test specimen with parafilm or caps and store

at 2-8C for no more than 2 hours.

5.	A sample must be tested within 2 hours.

### QUALITY CONTROL

See Coagulation Procedure 4840-CO-0100, Coagulation Quality Control.

### STEPWISE PROCEDURE

1.	Draw a citrated (blue top) tube from all employees who come to the lab

for employee health physicals, current employees and/or outpatients or ER patients not on any medications.  A minimum of 20 samples are used.

2.	Run a PT and PTT according to procedures in Coagulation Procedure manual.

3.	Record results on Normal Population Log (see example).

4.	When enough samples have been run, enter the data into the EP Evaluator using the

Verification of Reference Interval choice, (located on the Chem computer at the desk area

at the end of the bench).

5.         Type in the proposed reference interval (our current range).  Enter data.

6.	The data entered will result in a “PASS” or “FAIL” for that reference interval.

If not, contact the Sr. Tech. for readjustment of normal ranges. The new ranges will be submitted to the IS department for testing and implementation.  The mean calculated is the arithmetic mean.  You can calculate the geometric mean on line by typing in “Geometric Math Calculator”, enter the data and calculate.  Make screen shots to save as documentation.  Perform this task for both analyzers.

### REPORTING RESULTS

None indicated

### REFERENCES

None available