###### TITLE:  Prothrombin Time – Sysmex CA-660

###### PRINCIPLE

The Prothrombin Time test (PT) is a global screening test with three major applications. These include 1) detection of single or combined deficiencies of the extrinsic coagulation pathway indicative of hereditary and acquired coagulation disorders, liver disease or vitamin K deficiency; 2) monitoring test for oral anticoagulant therapy; and 3) assay for specific extrinsic coagulation factors. Thromboplastin and calcium ions are combined with test plasma at 37C; the PT is the time, in seconds, required for a fibrin clot to form.

SPECIMEN COLLECTION

Type:

Blue top (3.2%) sodium citrate vacutainer tube. Both the 2mL or 3mL tubes are acceptable as long they are filled properly (nine parts of freshly collected blood with one part of 0.11 mol/L (3.2% ) sodium citrate).

For proper collection, see policy 4840-CO-0120 Proper Collection & Preparation of Patient

Samples.

Handling Conditions:

The specimen should be transported at room temperature.  Prior to centrifugation, the whole blood specimen is checked for clot formation by gentle inversion and an applicator stick. Centrifuge the capped specimen for a minimum of 3 minutes at 4500 RPMs or at a speed and time that will produce platelet poor plasma. Although is it recommended to test samples as soon as possible the specimen is stable for 24 hours when stored at 15-25C. Do not store on ice or refrigerated. If testing is not complete within 24 hours, specimens can be frozen at -20°C or below for short term storage (up to 2 weeks) or -70°C or below for long term storage. Frozen plasma samples must be rapidly thawed at 37C while gently mixing and tested immediately.  Do not store at 37C for more than 5 minutes. Mixing is critical before testing.

EQUIPMENT AND MATERIALS

Equipment:

Sysmex CA-660 Instrument

4.0 mL Sample Cups

Reaction Tubes

Waste and Rinse Bottles

Materials:

Dade Innovin

Control material:

Dade Ci-Trol Coagulation Controls, Level 1, Level 3

Materials (cont.):

CA Clean™ I and Clean II

Preservative -free distilled or deionized water

Preparation

- Dade Innovin—lyophilized recombinant human tissue factor and synthetic phospholipids (thromboplastin); calcium ions; a heparin neutralizing compound; buffer; and stabilizers (Bovine Serum Albumin).

Reconstitute each vial of lyophilized Innovin with sterile deionized water using the volume stated on the vial label.

(according to label on vial).

Immediately after addition of water, mix contents of vial well before use to ensure a       homogenous solution. If allowed to stand, mix contents of vial before use. Continuous mixing is not necessary. Store at 2-8C. Do not freeze.

Stability after reconstitution:

On board	24 hours

2-8C 		10 days at (closed vial)

15-25C 	5 days in (closed vial)

37C 		24 hours at (closed vial)

This product is for in vitro diagnostic use.

2.	Dade Ci-Trol Coagulation Control, Level 1 (Normal) -Lyophilized preparation of human   plasma, stabilizers and buffers. Used for Quality Control.

3.	Dade Ci-Trol Coagulation Control Levels 3 (Abnormal)-Lyophilized preparations of

Human plasma, stabilizers and buffers. Used for Quality Control.

Stability after reconstitution:

| Condition | Ci-Trol 1, 3 |
| --- | --- |
| +2 to +8°C | 16 hours (closed vial) |
| +15 to +25°C | 8 hours (closed vial) |
| -20°C | -------------- |
| *Frozen and thawed once | NO |

  - Note:  Do not expose thawed control to 15 to 25°C for longer than 2 hours.

4.	CA Clean I is liquid and ready for use.
Store at 2-8 °C. Do Not Freeze.

Stable unopened at 2-8 °C until expiration date on bottle.

Opened bottles are stable for 30 days at 2-8 °C.

5. CA Clean II is liquid and ready for use.

Store at 5-35°C

Opened stablility is 2 months.

These products are for in vitro diagnostic use.

Performance Parameters:

Indication of deterioration: Controls will show deviations in results from the established laboratory range.

###### QUALITY CONTROL

Dade Ci-Trol Control Level 1

Dade Ci-Trol Control Level 3

(1)   Controls should be tested at the initiation of testing each day, upon reagent changes, and at least once each 8 hour shift of instrument use.

(2)   Controls should be run in the same manner as the test samples.

- Control tolerance limits--the range is calculated based on +/-2.0 to +/-2.5 standard deviations (SD) from the mean control value.

(4)   Corrective action when tolerance limits are exceeded:

- Rerun out-of-range control material.

- Verify reagent performance.

(c)  Check instrument performance.

- Document actions taken to identify and correct the problem before reporting any patient data.

PROCEDURES FOR SETUP

Loading Reagents: 
(when the instrument is not operating)

- 1.   Open Light Shield.

- 2.   Place Innovin® Reagent in Reagent holder position 1.

- 3.   Place CA Clean I in rinse position I ,CA Clean II in position II.

TO REPORT INR RESULTS

In order to report INR results enter the Innovin reagent lot specific ISI value and mean of the reference interval.

- [MAIN MENU]
[STANDARD CURVE]
[SELECT TESTS], (Select PT)
[MANUAL ENTRY]

- Arrow down to right field.

- Enter the lot specific ISI value and press [ENTER].

- Enter the determined PT mean value and press [ENTER], then [QUIT].

- Choose [SET].

NOTE: An INR will automatically be calculated if the ISI and geometric mean of the normal reference interval have been entered under the PT Standard Curve.

Loading Consumables and Discarding Waste Material: 
(do the following as necessary)

- 1.  Replenish Reaction tubes. 
Reaction tubes should be replenished as needed. The instrument will keep track of the last reaction tube position used.

- 2.  Replenish Rinse Solution.

  - Confirm that the instrument is not operating or has been interrupted due to insufficient rinse (error message “Replenish Rinse fluid”).

  - Remove the cap with the float switch from the Rinse solution container by turning the cap counter-clockwise. 
Caution: Do not touch the float switch. Rinse the float switch thoroughly with deionized or distilled water if contact is made.

  - Replace the Rinse solution with distilled or deionized water.

  - Replace the float switch by tightening the cap clockwise to close.

  - Verify that the tubing is securely connected and is not kinked.

  - If the analysis was interrupted due to insufficient Rinse, press [Resume] key.

- 3.  Dispose of Liquid Waste.

  - Confirm that the instrument is not operating or has been interrupted due to waste is full.

  - Remove the cap with the float switch from the Waste container by turning the cap counter-clockwise.
Caution: Biohazard

  - Dispose of the Biohazardous liquid waste.

  - Replace the float switch by tight the cap clockwise to close.

  - Verify that the tubing is securely connected and is not kinked.

- 4.  Dispose of Used Reaction Tubes from the Trash Box.

  - Confirm that the instrument is not operating.

  - On the right side of the instrument, remove the reaction tube trash box.

  - Discard the used reaction tubes.
Caution: Biohazard

  - Clean the reaction tube trash box with tap water, and thoroughly wipe off moisture from the reaction tube trash box.  Restore the reaction tube trash box.

- QUALITY CONTROL PROCESSING

- Controls are run from the sample rack.

- 1.  From the Main Menu, press [ID NO. ENTRY].  Use the assigned QC numbers.

- Press the [PT] key.

- Press [ENTER].

- Place the control material in a sample rack.

- If necessary, continue to program routine patient or other levels of quality control material. When completed, press [Start].

- Enter results from the QC in to the LIS.

## SAMPLE PROCESSING

Select test group to be used. Go to [Main Menu], [Settings], [Analysis Settings], [Test group]. Arrow to appropriate group. Select [Return], [Fix]. Return to Main Menu.

Samples can be programmed and loaded on the analyzer using one of four options depending whether the sample is barcoded and/or the host is connected:

- Manually enter sample ID (non-barcoded) / 
No automatic inquiry of tests (no host connection or the host is down)

- Manually enter sample ID (non-barcoded) / 
Automatic inquiry of tests (host connection operational)

- Sample ID number (barcoded sample) read by barcode reader / 
No automatic inquiry of tests (no host connection or the host is down)

- Sample ID number (barcoded sample) read by barcode reader / 
Automatic inquiry of tests (host connection operational)

I.	Manually enter sample ID (non-barcoded) / 
No automatic inquiry of tests (no host connection or the host is down)

- 1.  From the Main Menu, press [ID NO. Entry]. If ID No. Entry does not appear, press [Next]       for next available rack/position.

- 2.  Use the up and down arrow keys to select the desired sample position.

- 3.  The numeric keypad is displayed. Enter the sample ID number.
      Press the [PT] key.

- 4.   Continue to select test to be performed on the sample if necessary.

- 5.   Press [Enter]. Each test selected should show the symbol "".

- Load the sample in the respective sample rack position.

- Repeat until all sample ID numbers and tests are entered for the respective sample rack positions.

- When all samples are programmed and loaded in the respective sample rack position, press [Quit] on the numeric keypad.

- With all reagents and consumables on board, place the sample rack into position.

- Press [Start].

- Select [First Tube] or [Continue]: 
[Continue] will pick up cuvettes from the last position of cuvettes used.
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- The message “Analyzing” will be displayed. When all samples have been aspirated, the “Replace Rack? Yes” message will appear.

II.	Manually enter sample ID (non-barcoded) / 
Automatic inquiry of tests (host connection operational)

- 1. From the Main Menu, press [ID NO. Entry]. If ID No. Entry does not appear, press [Next] for next available rack/position.

- 2. Use the up and down arrow keys to select the desired sample position.

- 3. The numeric keypad is displayed. Enter the sample ID number.

- 4. Press [Enter].

- 5. Load samples in the respective rack position.

- 6. Repeat until all sample ID numbers are entered for the respective sample rack positions.

- 7. When all samples are programmed and loaded in the respective sample rack positions, press [Quit] on the numeric keypad

- 8. With all reagents and consumables on board, place the sample rack into position.

- 9. Press [Start].

- 10. Select [First Tube] or [Continue]. 
[Continue] will pick up cuvettes from the last position of cuvettes used.
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- 11. The message “Analyzing” will be displayed. When all samples have been aspirated, the “Replace Rack? Yes” message will appear.

III. Sample ID number (barcoded sample) read by barcode reader / 
No automatic inquiry of tests (no host connection or the host is down)

- 1.  From the Main Menu, press [ID NO. Entry]. If ID No. Entry does not appear, press [Next] for next available rack/position.

- 2.  Use the up and down arrow keys to select each tube position on the rack. Press the [PT] key.

- 3.  Continue to select test to be performed on the sample if necessary.

- 4.  Press [Enter]. Each test selected should show the symbol "".

- 5.  Load the sample in the respective sample rack position.

- 6.  With all reagents and consumables on board, place the sample rack into position. The bar code labels will face the user.

- 7. Press [Start].

- 8. Select [First Tube] or [Continue]. [Continue] will pick up cuvettes from the last position of cuvettes used. [First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- 9.The message "Analyzing" will be displayed. When all samples have been aspirated, the "Replace Rack? Yes" message will appear.

IV. Sample ID number (barcoded sample) read by barcode reader / 
Automatic inquiry of tests (host connection operational)

- 1.  When the host computer is connected using bi-directional communication, host inquiry         takes place when the sample ID is read and the analysis parameters are automatically registered

- 2.  Load the sample in the respective sample rack position.

- 3.  Press [Start] to begin processing.

- 4.  Select [First Tube] or [Continue]. 
[Continue] will pick up cuvettes from the last position of cuvettes used. 
[First] will begin picking up cuvettes from the top, right hand position of reaction tube tray.

- 5.  The message "Analyzing" will be displayed. When all samples have been aspirated, the "Replace Rack? Yes" message will appear.

NOTE: The cause of any printed error code must be investigated and the appropriate corrective action taken prior to reporting results.

### REPORTING RESULTS

PT results are reported in seconds. These results should be related to the reference interval. All PT results are reported with the INR.  Results of patients on oral anticoagulation should be reported as an INR. The INR is calculated from the lot-specific and instrument specific ISI of Innovin (see package insert) and the geometric mean of the normal reference interval.  Review values on the CA-660 printout, monitor with any error codes or flagging, and in the computer system with any previous history.  Review and verify quality control values in the LIS. Results will autoverify if QC has been run, is within limits and there are no instrument error codes, flags or critical results.

Hemolyzed, lipemic, or icteric samples must be noted as a comment with the result. In cases of extreme lipemia, icteric or hemolysis a result may not be possible if a valid curve is not obtained. Hemolyzed specimens should be re-collected. Lipemic or icteric specimens may have to be sent out to a reference lab for testing.

Refer to instrument Quick Guides for examples of curve errors.

Reference interval values will be based on the ISI value for the current lot of Innovin.

Reference Ranges:

Reference interval values will be based on the ISI value for the current lot of Innovin.

Analytical Measuring Range (AMR): 8.7 – 130.0. Results >130 seconds or an INR of >10.0 should be rerun and the specimen integrity of the sample should be evaluated.  If the PT is <9.3 on analyzer the INR will populate <0.9 into LIS, this is correct and results should be released.  If the PT is close to 9.3 then the INR sometimes calculates to a number less than 0.9 and <0.9 will populate into result field which is correct and results should be released.

Reference Range: PT	9.3-11.7 sec

Critical Values: INR: 	Outpatient: INR > 4.0

Inpatient: INR >5.0 or no clot detected

NICU:  >16.0 sec. or INR >5.0

All Critical Values must be phoned to the patient’s nurse or physician as a Critical Value. Document in the LIS with the (RBTO) comment:  RN’s full name, time/date and your initials.

As part of a routine check, the Sr. Tech or designee checks random printed patient reports as they are faxed to verify the following: the patient value, INR and that the correct Reference Range has printed.

PROCEDURE NOTES

Overall performance of PT testing is dependent on reagent and instrument performance. Acceptable variability (imprecision) should be such that the total day-to-day coefficient of variation (CV) of the analytic system is less than 5% on the same lot of control plasma.

Instrument Operations

Turn on the Power

- Turn on the power on the left side of the instrument.   The system will automatically performs a roughly 10-second self check, and the Main Menu will appear.

- When the detector and cooler reach an analysis-permitting temperature, the Main Menu screen displays “Ready”.  (This can take 5-30 minutes) (Make sure to replenish reaction tubes before running)

Turn OFF the Power

- Confirm that the instrument status is “Ready” before turning off the power switch.

Emergency Stop

- By pressing the Mechanical Stop switch, analysis operation can be stopped immediately.

- Press the Mechanical Stop Switch.  Analysis operation will stop immediately.

- When the Mechanical Stop switch is pressed during photo detection, the screen for confirming whether to stop photo detection comes on.

- To stop analysis Press [YES] key.

- To start discard/rinse operation, press [YES] key.

- When rinse operation cannot be performed, press [No] key.

- When an unanalyzed order remains, the screen for choosing whether to continue analysis comes on.

- Press [Start] key or [Cancel] key.

[Start] key: Starts analysis

[Cancel] key: Interrupts analysis and returns the screen to Main Menu.

LIMITATIONS OF THE PROCEDURE

Oral anticoagulants depress the production of factors II, VII, IX and X in the liver by inhibiting the action of Vitamin K. The PT is sensitive to the levels of factor II, VII and X and is used to monitor patient therapy with oral anticoagulants.Many commonly administered drugs may affect the PT results. These should be considered when unusual or unexpected abnormal results are obtained. Unexpected abnormal results should be followed by further coagulation studies to determine the source of the abnormality.

INTERFERENCES:

No interference up to:

| Triglycerides | 373 mg/dL |
| --- | --- |
| Hemoglobin | 1000 mg/dL |
| Bilirubin | 48 mg/dL |

REFERENCES

- Dade Innovin package insert, Siemens Healthcare Diagnostics Inc., Newark, DE, May   2008

- Dade Ci-Trol Coagulation Controls, Levels 1, 2 and 3 package inserts, Siemens

Healthcare Diagnostics Inc., Newark, DE, May 2008.

3.   Clinical Laboratory Standards Institute. Collection, Transport and Processing of Blood

Specimens for Testing Plasma-Based Coagulation Assays:  Approved Guideline-Fifth Edition. CLSI Publication H21-A5. Wayne, PA, January, 2008

4.   Clinical Laboratory Standards Institute, One-stage Prothrombin Time Test (PT) and Activated   Partial Thromboplastin Time (APTT) Test Approved Guideline: H47-A. Wayne, PA, June 1996.

5. Application sheet for PT Innovin® on CA-660

6. Sysmex CA-660 series Operator’s Manual

## S:\Laboratory Policies and Procedures\ Coagulation  7/4/2020 TML