## TITLE:   ARCHITECT CALIBRATION CONFIGURATION

## PRINCIPLE:

The purpose of this procedure is to give detailed instructions for calibration configuration on the Architect ci4100.

CLINICAL SIGNIFICANCE:

Perform this procedure to add a new lot number to a configured calibrator set and change the default lot.

REAGENTS:

See individual assay specific procedures

EQUIPMENT:

Architect ci4100

QUALITY CONTROL:

See individual procedures for more information and 7180-CH-310 Chemistry Quality Control.

### STEPWISE PROCEDURE:

i System Reagents

BHCG, BNP, Myoglobin, Troponin, and TSH

These analytes have the calibrator factors pre built into the Architect.  To change to a new (red dotted) calibrator follow these steps.

- Go to Orders -- Calibration Orders

- Enter cassette and position

- Select assay that needs new calibrator lot entered

- Select Assay options…(F5) at bottom of screen

- Enter lot number and expriation date and Done

- Add order to save the lot number

c System Reagents

Perform this procedure to enter a new calibrator lot number and/or change the default lot.

- All the modules need to be Stopped or Ready

- Log in under System Administrator.  Login: ADMIN  Password: ADM

- Go to System, then Configuration, then QC – Cal settings

- Select Calibrator Set from the QC – Cal categories list on the Configuration screen.

- Select the desired calibrator set, and then select F6 – Configure.  The Configure calibrator set window displays.

- Select the Lot number list button, and then select one of the following:

- The desired lot number if changing the default lot # (yellow sticker)

- New Lot if entering a lot that has not been previous built (red sticker)

- Select the Default check box to use this lot number as the default for calibration orders for assays in this calibrator set.  Need to select Default if changing to a lot that has a yellow sticker for furture calibration orders.  This will ensure the correct cal factors are used when calibrating.  Place green sticker on box when done.

- If selected New Lot enter the box Lot number in the Lot number data entry box.

- Enter a value in the Expiration date data entry box.

- Highlight the assay(s) from the table, and then select Define data.  The Define calibrator data window displays.

- Enter a value in the Concentration data entry box for each calibrator.

- Use the previous/next buttons to display each assay if you selected more than one, and then repeat step 8 for each.

- Select Done to return to the Configure calibrator set window.  The assay-specific data displays in the table.

- Select Done to save your changes.

- Return Modules to Running Status.

Import Calibrator Data

(MCC Calibrator)

Perform this procedure to import calibrator set data from a USB flash drive to the SCC.

- Select Calibrator set from the QC – Cal categories list on the Configuration screen.

- Select the desired calibrator set from the Calibrator sets list an then select F6 – Configure.  The configure calibrator set window displays.

- Select the Lot number list button and then select Import.

- Insert USB flash drive (in the MCC calibrator box) and select OK.  USB port is located on the CPU in the lower right hand side of Architect.

- Select the desired data file and then select Done.

- Select Done to return to the Configure calibrator set window.

- Select Done to save the imported data.

- Select OK if the data was imported from a USB flash drive then remove the drive from the USB port.

ASSAYS:

ICT Serum Calibrator

- Prebuilt into Architect so update calibrator lot number like i side reagents.

- Na, K, Cl

ACET

- Acetaminophen (ug/mL)

- 1 calibrator level

- Calibrator in the reagent box.  Value on the box.

AmmCA

- Ammonia (umol/L)

- 1 calibrator level

- Calibrator in the reagent box.  Value on reagent box and bottle.

Bil

- BiliD and BiliT (mg/dL)

- 2 calibrator levels

- Calibrator values on paper in cal box

C02C

- C02 (mmol/L)

- 2 calibrator levels

- Calibrator values on the bottles

CRP32

- CRP (mg/L)

- 6 calibrator levels

- Cal 1 5.00 (White)

- Cal 2 10.00 (Lt. Yellow)

- Cal 3 20.00 (Lt. Green)

- Cal 4 40.00 (Lt. Blue)

- Cal 5 160.00 (Magenta)

- Cal 6 320.00 (Brown)

DOAQ

- AmpQ, BarbQ, BenzQ, CocQ, and PCPQ

- 2 calibrator levels: DOA Neg Cal and DOA MC1 Cal 2

- Use DOA MC1 Cal 2 calibrator lot number/expiration for the calibrator set information

- DOA Neg Calibrator enter value of 0 for Cal 1

- Enter DOA MC1 Cal 2 values in Cal 2

- Values on the outside of Calibrator boxes and package insert

OPI3Q

- Qpi3Q

- 2 calibrator levels: DOA Neg Cal and Opiates 300 Calibrator

- Use Opiates 300 Calibrator lot number/expiration for the calibrator set information

- DOA Neg Calibrator enter value of 0 for Cal 1

- Enter Opiates 300 calibrator value in Cal 2 (300 ng/mL)

- Values on the outside of Calibrator boxes and package insert

THCQ

- THC (ng/mL)

- 2 calibrator levels: DOA Neg Cal and THC50 Cal

- Use THC50 Cal calibrator lot number/expiration for the calibrator set information

- DOA Neg Calibrator enter value of 0 for Cal 1

- Enter THC50 Cal value in Cal 2 (50 ng/mL)

- Values on the outside of Calibrator boxes and package insert

ETH

- Ethanol (mg/dL)

- 2 calibrator levels:  Ethanol Neg Cal and Ethanol 100 Cal

- Use Ethanol 100 calibrator lot number/expiration for the calibrator set information

- Ethanol Neg Calibrator value 0.0 mg/dL for Cal 1

- Ethanol 100 calibrator value 100.00 mg/dL for Cal 2

- Values on the outside of Calibrator boxes and package insert

Lipase

- Lipase (U/L)

- 1 calibrator

- Calibrator value is on the bottle

LIPIDMCC

- DLDL and UHDL (mg/dL)

- Lipid Multiconstituent Calibrator

- 1 calibrator

- Calibrator values on paper in cal box

MCC

- AlbG, BUN, CaC, Chol, CreaC, GluC, LactA, MAG, TP, TRIG

- Multiconstituent Calibrator

- 2 calibrators

- Use USB card to import calibrator data

SALI

- Salicylate (mg/dL)

- 1 calibrator

- Calibrator value on the box and bottle

UPRO

- Urine/CSF Protein (mg/dL)

- 5 calibrators

- Calibrator values on package insert and bottles

| Ref # | Cal Set Name | Calibrator (levels) | Analytes | Open Expiration |
| --- | --- | --- | --- | --- |
| 506-30 | ACET | Acetaminophen L3K Calibrator (1) in kit, value on box | Acet (ug/mL) | Until kit Expiration @ 2-8oC |
| 6K89-30 | AmmCA | Ammonia Ultra (1) in kit, value on reagent box and bottle | Amm (umol/L) | 120 days @ 2-8oC |
| 1E66-04 | Bil | Bilirubin Calibrator (2) | BiliD, BiliT (mg/dL) | 7 days @ 2-8oC |
| 8K28-02 | - | BNP Calibrators (6) Prebuilt factors, enter lot on cal order screen | BNPStat | Frozen 90 days @2-8oC |
| 1E64-02 | C02C | Carbon Dioxide Calibrator (2) | CO2C (mmol/L) | 30 days @ 2-8oC |
| 6K26-10 | 32CRP | CRP Calibrator Set (use CRP05, CRP10, CRP20, CRP40, CRP160, CRP320) see cheat sheet | CRP32 (mg/L) | 3 months @ 2-8oC |
| 3L43-05 3L43-02 | DOAQ | DOA Neg Calibrator DOA MC1 Cal 2 calibrator Enter lot number/exp from Cal 2 | AmpQ, BarbQ, BenzQ, CocQ, OpiQ, PCPQ | Until kit Expiration @ 2-8oC |
| 3L43-05 3L34-02 | OPQ | DOA Neg Calibrator Opiates 300 calibrator Enter lot number/exp from Opiates 300 | OpiQ | Until kit Expiration @ 2-8oC |
| 3L43-05 3L41-02 | THCQ | THC50 Cal calibrator DOA Neg Calibrator Enter lot number/exp from THC50 | THC5Q | Until kit Expiration @ 2-8oC |
| 3L36-01 3L36-02 | ETH | Ethanol Neg Calibrator Ethanol 100 Calibrator Enter lot number/exp from ETOH100 | EtohM (mg/dL) | Until kit Expiration @ 2-8oC |
| 1E46-03 | - | ICT Serum Calibrator (2) Enter new lot with iside directions | Cl-C, K-C, Na-C | 7 days @ 2-8oC |
| 3E16-02 | Lipase | Lipase Calibrator (1) | Lip (U/L) | 28 days @ 2-8oC |
| 5P56-01 | LIPIDMCC | Lipid Multiconstituent Calibrator (1) | DLDL, UHDL (mg/dL) | 7 days @ 2-8oC |
| 1E65-05 | MCC | MCC Multiconstituent Calibrator (2) | AlbG, BUN(Urea), CaC, Chol, CreaC, GluC, LactA, MAG, TP, Trig | 7 Days @ 2-8oC |
|  | - | None (Water Blank only) | ALT, AST, AlkP, Amy, CK | on board analyzer |
| 3K01-01 | SALI | Salicylate Calibrator (1) | Sali (mg/dL) | Until kit Expiration @ 2-8oC |
| 2K43-01 | - | STAT Myoglobin Calibrators (6) Prebuilt factors, enter lot on cal order screen | MYOGLOBIN | Until kit Expiration @ 2-8oC |
| 2K41-01 | - | Tnl II (Troponin) Calibrators (6) Prebuilt factors, enter lot on cal order screen | Tnl II (Troponin) | Until kit Expiration @ 2-8oC |
| 7K78-01 | - | Total B-hCG Calibrator set (6) Prebuilt factors, enter lot on cal order screen | STAT-B-hCG | Until kit Expiration @ 2-8oC |
| 7K62-01 | - | TSH Calibrators (2) Prebuilt factors, enter lot on cal order screen | TSH | Until kit Expiration @ 2-8oC |
| 1E71-02 | UPro | Urine/CSF Protein Calibrator (5) | Upro/CSF (mg/dL) | Until kit Expiration @ 2-8oC |

REPORTING AND INTERPRETING OF RESULTS:

Calibration records are reviewed for acceptability and QC is verified before running patient specimens.

REFERENCES:

Architect System Operations Manual,  September 2012