TITLE: Pharmacy Sterility Environmental Culture

PRINCIPLE:  Provide instruction for the receipt, ordering, culturing and result interpretation of gloved fingertip sampling, aseptic media fill, and surface sampling specimens collected by pharmacy staff.

PERSONNEL:  Medical Technologist and Technicians, LCC staff

# EQUIPMENT AND SUPPLIES:

35oC CO2 incubator

Chocolate Agar Plate (CHOC)

TSB Growth Media Pouch (THIO)

Surface Sampling Paddle

STEPWISE PROCEDURE:

Pharmacy will submit inoculated media for testing. There are 4 types of testing.

- Initial x3 Finger Tip test – 3 sets of 2 CHOC plates. 3 separate orders will be placed.

- Ongoing Finger Tip test – 1 set of 2 CHOC plates. 1 order will be placed

- Monthly Surface samples – 1 paddle for each surface samples. Separate order for each surface sample paddle.

- Ongoing Media Fill w/ Finger Tip test and Hood Surface sample – 2 CHOC, 1 pouch of TSB Growth Media and 1 surface sample paddle “hood”.

The Chocolate plates (fingertip sampling) shall be labeled with the collector’s name, date, time and an R or L to designate which hand fingertips were sampled. The TSB Broth (media fill) will be labeled with the collector’s name, date, and time. Sampling paddles will have location sampled.

Ordering in Soft system:

Patient Name: Pharmacy, Sterility

Order: Culture Environmental

Source: Pharmacy Sterility Check
Body Site: Free text the Name of the Pharmacy staff that collected the samples and types of samples collected. Example:.(John Doe R,L) or (John Doe R,L Thio, Hood Paddle). If surface sample paddle received, free text Paddle and the sampling location written on the container.

Plating:

When specimen is plated in microbiology 2 CHOC, 1 THIO and 1 PADDL media labels will print. Label media without covering sample source information.

- All CHOC (fingertip) and Paddle (Surface Samples) media will be incubated in 37C CO2 incubator for 2 days, then room temperature (20-25C) for additional 5 days.

- All TSB Growth Media will be incubated at room temp (20-25C) for 7 days then an additional 7 days in 37C CO2 incubator.

Reading and Reporting Culture:

- Any colony growth shall be quantified and have a basic identification based on Gram Stain and spot biochemical testing such as Staphylococcus Coagulase Negative, Diptheroids, Bacillus sp., GNR, Pseudomonas Sp. Etc.

- Examine media daily for growth and update culture accordingly

- Save media with growth for Pharmacy

- Print Final reports and place in Pharmacy folder

1. Initial x3 Finger Tip test – Incubate 2 CHOC plates in CO2 incubator for 48 hours then room temperature for 5 days.

At 24 hours incubation inspect for growth. Issue Prelim

~“Right Hand: No Growth to date”

~“Left Hand: No Growth to date”

After 48 Hours incubate 2 CHOC plates at room temperature for 5 additional days.

After 7 days issue Final

“Right Hand: No Growth at 7 Days”

“Left Hand: No Growth at 7 Days”

2. Ongoing Finger Tip Test - Incubate 2 CHOC plates in CO2 incubator for 48 hours then room temperature for 5 days.

At 24 hours incubation inspect for growth. Issue Prelim

~“Right Hand: No Growth to date”

~“Left Hand: No Growth to date”

After 48 Hours incubate 2 CHOC plates at room temperature for 5 additional days.

After 7 days issue Final

“Right Hand: No Growth at 7 Days”

“Left Hand: No Growth at 7 Days”

3. Monthly Surface samples -Incubate Paddle in CO2 incubator for 48 hours then room temperature for 5 additional days.

At 24 hours incubation inspect for growth. Issue Prelim

~“Surface Sample: No Growth to date”

After 7 days issue Final

“Surface Sample: No Growth at 7 days.”

4. Ongoing Media Fill w/ Finger Tip test and Hood Surface sample-– Incubate 2 CHOC plates and Paddle in Co2 incubator for 48 hours then room temperature for 5 days. Incubate TSB Growth Media at room temperature for 7 days and then 7 additional days in Co2 incubator.

At 24 hours incubation inspect for growth. Issue Prelim

~“Right Hand: No Growth to date”

~“Left Hand: No Growth to date”

~“Surface Sample: No Growth to date”

~“Media Fill: No Growth to date”

After 7 days update CHOC and Paddle results. Issue Prelim

“Right Hand: No Growth at 7 Days”

“Left Hand: No Growth at 7 Days”

“Surface Sample: No Growth at 7 days.”

~“Media Fill: No Growth to date”

After 14 days issue Final

“Right Hand: No Growth at 7 Days”

“Left Hand: No Growth at 7 Days”

“Surface Sample: No Growth at 7 days.”

“Media Fill: No Growth at 14 Days.”

If TSB becomes turbid issue Final

“Media Fill: Turbid- Growth Detected”

REFERENCES:

2016 The United States Pharmacopeial Convention. Nov. 2016 pg.59

© 2021 USPC - 797 PHARMACEUTICAL COMPOUNDING〈797〉—STERILE PREPARATIONS, 〈797〉 Pharmaceutical Compounding—Sterile Preparations. This proposal is based on the version of the chapter official as of May 1, 2020.

Current DocID: GUID-A4CAAA8B-6F02-4AB8-8628-09E102CBD703_6_