| Owner Title / Approval: | John R Samuel, MT(ASCP) Signature on file |
| --- | --- |
| Administrative Director Approval: | Not required |
| Medical Director Approval: | I have reviewed this procedure and approve it for use. |
|  | Cindy Sturtz, MD Signature on file |

# PURPOSE:

This document establishes a procedure for performing ionized calcium testing using the Chem 8+ cartridge on the Abbott iSTAT handheld analyzer.

REAGENTS/SUPPLIES:

- Abbott iSTAT handheld analyzer

- Abbott iSTAT Chem8+ Cartridge [Cat#09P31-25]

  - Store cartridges at 2-80C until expiration date listed on packaging.

  - Allow cartidge[s] to warm at room temperature in the sealed package for 5-10 minutes prior to use.

  - Once warmed to room temperature, sealed cartridges are stable for 14 days at room temperature.  Date all room temp cartridges with +14-day expiration date.

  - Once warmed to room temperature, DO NOT return cartridge[s] to refrigeration.

- Disposable micropipettes

SPECIMEN/SPECIMEN COLLECTION:

- Whole blood sample collected in Lithium Heparin.

- Tube must be filled to the fill volume stated on the tube label [ie “3.0 mL”]

- If not testing within 30 minutes, refrigerate at 2-80 C until testing can be performed.

- Allow refrigerated samples to equilibrate to room temperature prior to testing.

- Testing should be performed as soon as possible after collection.  Samples greater than 24 hours old will be rejected.

- Sample must remain closed until testing is performed.

QUALITY CONTROL:

- i-STAT TriControls, Level 1  [Cat#05P71-01]

- i-STAT TriControls, Level 3  [Cat#05P73-01]

- Routine quality control is performed monthly – the analyzer will inform user when QC is due.

- Additionally, each shipment of reagent cartridges - whether a new lot or not - will be QC’d same day as receipt to ensure that cartridges have not been compromised in shipping.

PROCEDURE:  Patient Testing

| Step | Action |
| --- | --- |
| 1 | If refrigerated, remove cartridge and allow to stand at room temperature for at least 5 minutes in the sealed envelope before using. |
| 2 | Turn the i-STAT handheld On. |
| 3 | Press 2 for i-STAT cartridge. |
| 4 | At the prompt “Scan or Enter Operator ID,” scan your HH Employee ID number barcode. - You may manually enter your employee ID number using the keypad and pressing the Enter key. You will be prompted to manually reenter your employee ID for confirmation. |
| 5 | At the prompt “Scan or Enter Patient ID,” scan the patient tube label barcode. |
| 6 | The analyzer screen will display the Patient ID number with the comment “ID Not In Valid ID List.” Select option 2 – Continue to proceed. |
| 7 | At the screen prompt “Scan Catridge Lot Number,” scan the barcode number on the cartridge envelope. |
| 8 | Remove cartridge from the cartridge envelope. |
| 9 | Ensure that the patient sample is well mixed. |
| 10 | Using a disposable micropipette, transfer patient sample to the sample well on the test cartridge. Observe as the sample migrates up the channel from the sample well. |
| 11 | Close the cartridge. |
| 12 | Insert the cartridge into the analyzer. |
| 13 | Results are available in 3 to 4 minutes. Do not move analyzer during testing. |

RESULT REPORTING:

| Step | Action |
| --- | --- |
| 1 | Upon completion, test results are displayed on the the analyzer screen. - If the analyzer has turned off, press the power key and select option 1 “Last Results” to view test results. - Use the arrow keys to navigate through test results if required/desired. |
| 2 | Referring to the analyzer display, manually enter the iCa [ionized calcium] result into Meditech. |
| 3 | Document the patient information and result on the paper “Manual Chemistry Result Log.” |

REFERENCE RANGE:

Normal range:  1.12 – 1.32 mmol/L

Critical  value:  < 0.82 mmol/L or  >1.55 mmol/L

REFERENCES:

- Ionized Calcium/ICA, Package insert revision 7/15/16.  Art# 714179-00P, Abbott Point of Care.

Document History

| Date of Origination and Document Control Number | CHEM 4300.0 3/16/2018 |  |
| --- | --- | --- |
|  |  | Prepared by: John R Samuel, MT(ASCP) |
| Revision History/ Biannual Review: |  |  |
| Revision History/ Biannual Review: |  |  |
| Revision History/ Biannual Review: |  |  |
| Revision History/ Biannual Review: |  |  |
| Revision History/ Biannual Review: |  |  |

| Medical Director or Designee Approval: | I have reviewed this document and approve it for use ___ pending approval of Medical Director of record. ___ change in Medical Director of record. ______________________________ Date:___________ Signature on file |
| --- | --- |